## 문헌특성_IOL
1. 선택문헌의 특성
no. | 1저자 (연도) | 연구설계 (RCT, NRCT, Obs) | Trial명(NCT no.) (있는 경우) 예. | 연구국가 | 연구대상자 |  |  | 환자특성 | [9 empty] | 중재군(I) |  | 2군 |  | 3군 |  | 4군 |  | 결과변수 (1. VA, 2.CS, 3. Satisfaction, 4. Inpendence, 5. Complicaiton | 추적관찰기간 (개월) | 저자결론
[5 empty] | 대상자 모집기간 | 대상자 | 선택/배제기준 | 연구대상자 수: 명(안) | [6 empty] | 기저특성 |  |  | 중재군 |  | 대조군(C) |  | 대조/중재군명 |  | 대조/중재군명
[8 empty] | 총 | 비교군 수 (2군, 3군 등) | 중재군 (N명(안)) | 대조군 (N명(안)) | 3군 (N명(안)) | 4군 (N명(안)) | 탈락률(%) 전체(중재 vs. 대조) | 평균연령 | 남성 (%) | 양안시술여부(Y/N) |  | 렌즈명 |  | 렌즈명 |  | 대조군렌즈명 |  | 대조군렌즈명
10_국내 | Kretz (2016) | Prospective study | NA | 독일 | NA | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | 배제기준: patients had a history of glaucoma or retinal detachment, corneal disease, irregular corneal astigmatism, abnormal iris, macular degeneration or retinopathy, neurophthalmic disease, history of ocular inflammation, or previous ocular surgery or corneal astigmatism of 1.25 diopter (D) or greater | 50명(100안) | 1 | 50명(100안) |  |  |  |  | 57.5 | NA | Yes | Diffractive Trifocal IOL | AT LISA tri 839MP | [6 empty] | 1, 3, 4, 5 | 3개월 | 백내장 수술 후 AT LISA tri 839MP IOL 이식은 최소 수준의 광현상(photic phenomena), 시력관련 일상활동, 수술 후 환자만족데 긍정적인 영향을 미치는 효과적인 시력회복을 제공함
28_국내 | 곽준영 (2012) | Prospective study | NA | 한국 | 2010.9-2010.12 | 백내장수술 | 포함기준: 40-80세 사이의 연령으로 노인성 백내장 진단을 받고 술후 돋보기 착용을 원치 않는 환자들을 대상으로 함 배제기준: 각막 난시는 -1.0디옵터 미만으로 하였고, 녹내장이나 망막 박리, 각막 혼탁, 각막 또는 안내 수술의 기왕력, 홍채 이상, 황반 변성을 포함한 시력에 영향을 줄 수 있는 망막질환이나 각막 굴절수술을 시행 받은 사람은 제외함. 또한 직업적인 야간운전이나 금속 세공과 같은 정밀한 근거리 작업을 요하는 사람이나 중간거리와 근거리를 포함한 모든 다양한 거리를 원거리 시력처럼 선명하게 보기를 원하는 시력에 대한 기대치가 지나치게 높은 사람은 제외함 | 20명(40안) | 1 | 20명(40안) |  |  |  |  | 57.2 | 45% | Yes | Diffractive Bifocal IOL | Acri.LISA 366D | [6 empty] | 1, 2 | 6개월 | Acri.LISA 366D다초점 인공수정체는 우수한 원거리, 중간거리, 근거리 시력교정 및 대비감도를 보이고 백내장 술 후 환자의 만족도를 높여 노안교정에도 유용하리라 생각함
46_국내 | 최문정 (2020) | Prospective study | NCT03997890 | 한국 | 2016.6-2018.6 | 백내장수술 | 배제기준: the presence of amblyopia, keratoconus, previous corneal or refractive surgery, chronic or recurrent uveitis, acute ocular diseases, previous ocular surgery, glaucoma, and any ocular diseases that could affect postoperative visual acuity. | 96명(192안) | 1 | 96명(192안) |  |  |  | 4% | 61.19 | 35.4% | Yes | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric | [6 empty] | 1, 3, 4, 5 | 4-6개월 | Symfony EDOF IOL은 우수한 원거리, 중간거리, 기능적 근거리 시력을 제공함. 시각적 결과는 눈에 띄는 수준의 안경 독립성 및 환자 만족도와 관련이 있음
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study |  | 독일 | NA | Cataract surgery | 배제기준: history of glaucoma or retinal detachment, corneal disease, regular corneal astigmatism greater 0.75 D, irregular corneal astigmatism, abnormal iris, macular degeneration or retinopathy, neurophthalmic disease, history of ocular inflammation or previous ocular surgery, and expected postoperative CDVA of ≤0.3 logMAR. | 30명(60안) | 1 | 30명(60안) |  |  |  |  | 68.3 | 46% | Yes | Refractive Bifocal IOL | Lentis-313MF15 | [6 empty] | 1, 3, 4, 5 | 3개월 | 새로운 비구면, 비대칭, + 1.5D 근거리 추가 MIOL은 중간거리에 대한 우수한 성능 및 근거리에 대한 기능적 결과와 함께 원거리 시각 기능에 대한 좋은 결과를 제공함
78_국내 | 권영기 (2015) | Retrospective study |  | 한국 | 2014.7-2014.12 | 백내장수술 | 포함기준: 전면각막난시가 1.5D 이하이며 안경을 벗고자 하는 환자를 대상으로 함 배제기준: 안외상의 기왕력, 안과 수술력, 불규칙 난시가 있거나 술 후 시력에 영향을 줄 수 있는 각막혼탁, 망막질환, 녹내장, 약시 등이 있는 환자 | 22명(44안) | 1 | 22명(44안) |  |  |  | 9.091% | 52.6 | 45.45% | Yes | Diffractive Trifocal IOL | AT Lisa Tri 839MP | [6 empty] | 1, 3, 4 | 6개월 | 회절성 삼중초점 인공수정체는 우수한 원거리, 중간거리, 근거리 시력을 보여주었고 환자의 만족도가 높았음
198_국내 | 박율리 (2018) | 전향, 무작위배정 |  | 한국 | 2012.2-2012.8 | 백내장수술 | 포함기준: 50-80세 백내장, 난시는 -1.0디옵터 미만으로 제한하였고, 당뇨가 없고 망막 소견이 정상인 환자를 대상으로 함. 또한 환자의 주관적인 측면에서 돋보기를 착용하지 않길 원하는 환자를 대상으로 함 배제기준: 약시, 사시, 녹내장이나 망막박리, 각막 혼탁, 각막 또는 안내 수술의 기왕력, 홍채 이상, 황반변성을 포함한 시력에 영향을 줄 수 있는 망막질환이나 각막 굴절수술을 시행 받은 사람, 수술 상담 후 수술 결과에 대한 기대치가 지나치게 높은 환자는 제외함 | 21명(40안) | 2 | 11명(20안) | 10명(20안) |  |  |  | 중재: 59.7, 대조: 63.7 | 38.1% | Yes | Hybrid Bifocal IOL | OptiVis | Diffractive Bifocal IOL | AcrySof ReSTOR D1 |  |  |  |  | 1, 2, 3, 4, 5 | 2개월 | OptiVis 인공수정체는 원거리 및 근거리 시력저하없이 우수한 중간거리 시력교정이 가능하여 높은 만족도를 보임
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) |  | 스페인 | 2005.1-2006.11 | Emmetropic presbyopia | 선택기준: emmetropic eyes (sphere and cylinder ±0.25 diopters [D]), aged between 50 and 60 years, and the desire for correction of distance and near vision without the use of spectacles or contact lenses. 배제기준: cataract, CDVA worse than 20/20 in either eye, history of glaucoma or retinal detachment, corneal disease, previous corneal or intraocular surgery, abnormal iris, pupil deformation, macular degeneration or retinopathy, neuro-ophthalmic diseases, and history of ocular inflammation. | 23명(46안) | 1 | 23명(46안) |  |  |  | 0% | 53.8 | 30.43% | Yes | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL | [6 empty] | 1, 3, 4, 5 | 6개월 | 정시 눈에 ReSTOR MIOL을 이식하여 노안 렌즈를 교체하는 것은 노안 교정을위한 효과적이고 안전한 방법임
272 | Pedrotti (2017) | Prospective single-blind comparative study |  | 이탈리아 |  | Cataract surgery | 포함기준: bilateral senile cataract and age between 50 and 70 years. 배제기준: corneal astigmatism greater than 1.00 D, myopia greater than 6.00 D, amblyopia, previous anterior or posterior segment surgery, and history of other ocular pathologies impairing visual function. | 42명 | 2 | 21명 | 21명 |  |  | 0% | 중재: 64.9, 대조: 66.1 | 45.24% | Yes | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF | Monofocal IOL | Tecnis 1-piece |  |  |  |  | 1, 2, 5 | 12개월 | 이 연구결과는 Comfort IOL +1.5D가 단초점 IOL과 비슷한 대비감도, 우수한 중간거기 안경독립성, 낮은 주관적 증상과 함께 좋은 시력, 비슷한 대비감도를 보여줌
323 | Gundersen (2016) | Randomized prospective study |  | 노르웨이 | NA | Cataract surgery | 포함기준: >50 years old, had presented for uncomplicated cataract surgery, and were interested in reducing their dependence on glasses for daily life. They had to have regular corneal astigmatism and have a calculated IOL power that was within the available range for each IOL type. 배제기준: ocular pathology (besides cataract) and previous refractive surgery. If the surgeon felt there were factors that would be likely to affect the subject’s postoperative vision (eg, amblyopia and history of uveitis), the surgeon could exclude the subject. | 22명 | 2 | 11명 | 11명 |  |  |  | 중재: 62.1, 대조 70.2 | 45.45% | Yes | Diffractive Trifocal toric IOL | FineVision IOL (PhysIOL ® ; FineVision Toric, Liège, Belgium) | Diffractive bifocal toric IOL | ReSTOR SND1T |  |  |  |  | 1, 4, 5 | 3개월 | trifocal toric IOL은 bifocal toric IOL과 비교할 때 시각 기능 및 근거리, 원거리, low contrast VA에 부정적인 영향을 없이 중간거리 시력을 개선함. 삼중 초점 렌즈의 원환체 구성요소는 난시를 효과적으로 감소시키고 우수한 회전 안정성을 제공함
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study |  | 이탈리아 | 2015.9-2016.1 | Cataract surgery | 포함기준: lens opacity causing a reduction in visual quality and the motivation for spectacle independence. 배제기준: professional night drivers, pilots, and those with other occupations for whom induced dysphotopsia could put their career at risk; inability to cooperate; difficulties comprehending written hamper or spoken language; ocular comorbidity that might postoperative visual acuity; previous refractive surgery; pseudoexfoliation and/or zonular fiber weakness; optical biometry impractical because of dense cataract; axial lengths of 22.00 mm or shorter 41.00 diopters or 26.00 mm or longer; a mean central corneal power of (D) or lower or 46.00 D or higher; corneal astigmatism of 0.75 D or higher; an irregular astigmatism index of 0.54 or higher; root mean square of corneal higher-order aberrations of 0.30mm or higher at a 4.0 mm pupil diameter; and angle k or 0.29 mm or higher. | 60명(120안) | 3 | 20명(40안) | 20명(40안) | 20명(40안) |  | 21.05% | 중재: 66, 대조1: 67, 대조2: 68 | 53.33% | Yes | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  |  | 1, 4, 5 | 4개월 | 두개의 다초점 IOL은 중간시력이 필요한 환자에게 좋은 옵션이지만 삼중초점 IOL은 근거리 시력이 필요한 환자에게 더 좋을 수 있음. 시각적 부작용에 대한 유의한 인식은 다초점 IOL을 시술하기 전에 환자에게 이러한 영향에 대해 상담을 받아야 함을 보여줌
391 | Alio (2011) | Prospective randomized controlled clinical study |  | 오스트리아, 스페인 | NA | Presbyopia | 포함기준: age older than 45 years and the need for bilateral cataract refractive surgery for presbyopia in the presence of significant nuclear sclerosis. 배제기준: active ocular disease and illiteracy | 152명(304안) | 4 | 78안 | 84안 | 70안 | 72안 | 0% | 71.32 | NA | No | Apodized diffractive multifocal IOL | AcrySof Restor SN6AD3 (Alcon) | Diffractive bifocal IOL | Acri.Lisa 366D (Carl Zeiss Meditec AG) | Refractive Multifocal | Rezoom (Abbott Medical Optics, Inc.) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) | 1 | 6개월 | Diffractive MIOL은 Refractive MIOL 및 단초점 IOL보다 더 좋은 reading perfromance 를 보임
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | NCT02529488 | 호주, 칠레, 독일, 이탈리아, 스페인, 벨기에, 영국, 프랑스, 네덜란드 | 2015.9-2017.6 | Cataract surgery | 포함기준: aged 22 years or older with bilateral cataracts with planned cataract extraction by phacoemulsification. Additional inclusion criteria were preoperative corrected distance visual acuity (CDVA) worse than 0.20 logarithm of the minimum angle of resolution (logMAR) in both eyes, potential postoperative CDVA equal to or better than 0.20 logMAR in both eyes, and preoperative regular corneal astigmatism less than 1.00 D in both eyes. 배제기준: any clinically significant ocular abnormalities or diseases; degenerative eye disorders; glaucoma or ocular hypertension; rubella, congenital, traumatic, or complicated cataracts, or any conditions expected to cause intraocular inflammation; known color vision deficiencies; anterior chamber depth of 2.5 mm or less; history of or planned refractive surgeries; or previous corneal transplant. If any of the mentioned criteria were discovered during the first-eye surgery before the IOL came into contact with the eye, implantation was not performed, and the patient was discontinued from the study: any additional procedures during phacoemulsification and IOL implantation because of intraoperative complications or incision site or haptic placement not in accordance with the protocol. If these criteria were discovered during the second-eye surgery before the IOL came into contact with the eye, implantation of the second eye was not performed. | 149명(298안) | 1 | 149명(298안) |  |  |  | 2.685% | 68.9 | 38.26% | Yes | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) | [6 empty] | 1, 5 | 12개월 | 연구 IOL은 좋은 VA 결과를 제공함. Defocus 곡선은 근거리에서 중간 거리까지 20/25 Snellen 이상의 VA를 보여주었으며 AE의 비율은 낮았음
469 | Pedrotti (2020) | Prospective observational comparative clinical study |  | 이탈리아 | NA | Presbyopia | 선택기준: patients with intraoperative complications, previous ocular surgery including corneal, retinal or refractive surgery, previous ocular trauma, and any ocular disease such as amblyopia, chronic or recurrent uveitis, acute ocular disease, diabetes mellitus with retinal changes, glaucoma, pseudoexfoliation syndrome or zonular laxity, optic nerve atrophy, keratoconus, alteration in colour perception and corneal endothelial dystrophy. 배제기준: patients with intraoperative complications, previous ocular surgery including corneal, retinal or refractive surgery, previous ocular trauma, and any ocular disease such as amblyopia, chronic or recurrent uveitis, acute ocular disease, diabetes mellitus with retinal changes, glaucoma, pseudoexfoliation syndrome or zonular laxity, optic nerve atrophy, keratoconus, alteration in colour perception and corneal endothelial dystrophy. | 50명(100안) | 2 | 25명(50안) | 25명(50안) |  |  | 0% | 중재: 70.16 대조: 72.12 | NA | Yes | Diffractive EDOF | Tecnis Symfony ZXR00 | Diffractive trifocal IOL | Acrysof IQ PanOptix TNFT00 |  |  |  |  | 1, 2, 5 | 3개월 | 두 안구 내 수정체는 서로 다른 특징을 나타내어도 가성 노안을 피할 수있는 유효한 옵션입니다. Symfony는 환자가 더 나은 객관적이고 주관적인 시력 및 대비 감도를 얻을 수 있도록합니다. PanOptix는 근거리 시력 및 안경 독립 요건에서 더 나은 결과를 제공합니다.
471 | Kohnen (2020) | Prospective single-arm study |  | 독일 | 2017.4-2019.9 | Cataract surgery, Presbyopia | 포함기준: bilateral cataract surgery or refractive lens exchange with implantation of the toric trifocal PanOptix IOL, total corneal refractive power (anterior and posterior astigmatism) greater than 0.75 D measured with the Pentacam (Oculus Optikgeräte GmbH). 배제기준: Patients were excluded from the study in case of previous ocular surgeries, ocular pathologies, and a reduced endothelial cell count. | 25명(50안) | 1 | 25명(50안) |  |  |  | 0% | 60 | 32% | Yes | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 | [6 empty] | 1, 2, 3, 4, 5 | 3개월 | Toric PanOptix IOL의 시각적 성능은 모든 거리에서 우수한 시력을 보여줌. 90% 이상이 refractive cylinder이 0.75D 이하로 감소하고 일부 광학 현상에도 불구하고 환자 만족도가 높으며 수술 3 개월 후 높은 안경 독립성을 달성함
488 | Alio (2020) | Prospective consecutive case series |  | 스페인 | NA | Cataract surgery, Presbyopia | 포함기준: candidates with early cataract and presbyopia looking for spectacle independence (cataract or clear lens extraction patients) who qualify for bilateral implantation and patients wishing to be spectacle independent. 배제기준: presence of any other ocular pathology which could compromise visual outcomes (amblyopia, anterior segment comorbidities, glaucoma, retinal disorders or neuro-ophthalmic disease) and preoperative astigmatism higher than 1.0 D measured with optical biometer or corneal topography. | 10명(20안) | 1 | 10명(20안) |  |  |  | 0% | 63.8 | NA | Yes | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) | [6 empty] | 1, 2, 3, 4 | 12개월 | Precizon Presbyopic NVA IOL (OPHTEC BV)은 좋은 시각적 결과를 제공함. 다초점 IOL은 원거리, 중간거리 및 근거리에서 우수한 VA와 만족스러운 대비 감도로 인해 높은 비율의 안경 독립성을 제공함. 시력의 질과 만족도 설문지에서 환자 만족도가 높았으며 광 현상을 나타내는 환자 비율이 낮았음
513 | Power (2018) | Prospective cohort analysis |  | 아일랜드 | 2016.9-2017.9 | Cataract surgery, Presbyopia | 포함기준: range of preoperative astigmatism that was included was 0-1.25D. Cases with pre-existing ocular pathology were excluded 배제기준: targets for mini-monovision, incomplete data, other ocular pathology. Data were prospectively collected on pre- and postoperative refraction, keratometry, distance vision, near vision, surgical wound site and Surgically Induced Astigmatism (SIA). | 53명(94안) | 1 | 53명(94안) |  |  |  | 0% | NA | NA | No | Diffractive EDOF IOL | Symfony | [6 empty] | 1 | 중앙값 4.1개월 | Symfony IOL은 노안을 해결하고 수술 후 안경 의존도를 줄이는 효과적인 수술 수단임. 수술 전 각막 난시가 1D 이상인 환자가 통계적으로 유의하게 열등하고 예측하기 어려운 시각적 결과를 보이기 때문에 잠재적인 안경 독립성을 제공 할 때 주의를 해야 함
536 | Escandon-Garcia (2018) | Prospective, nonrandomized, examiner-masked case series |  | 포르투갈 | NA | Cataract surgery | 포함기준: no active ocular disease except cataract, nonsevere dry eye, uneventful cataract surgery and postoperative healing process, clear posterior capsule and lens implant, no pupillary abnormality, postop refractive error within ±1.00 diopters (D), and an unaided postoperative visual acuity of 0.10 logMAR or better. 배제기준: IOL dislocation, posterior capsule opacification, or any vitreous or retinal disease. | 45명(90안) | 3 | 7명(14안) | 15명(30안) | 23명(46안) |  | 0% | 63 | 22.22% | Yes | Diffractive Trifocal IOL | AcrySof® IQ PanOptix (TFNT00) | Diffractive EDOF | TECNIS® Symfony model ZXR00 (Abbott Medical Optics, Santa Ana, USA) | Diffractive Trifocal IOL | FineVision Pod F (PhysIOL, Liège, BE) |  |  | 1, 2 | 3개월 | Trifocal IOL은 근거리에서 더 나은 성능을 제공하는 반면 EDOF IOL은 중간 거리에서 더 나은 성능을 제공함. Light Distortion Analyzer로 측정한 Objective dysphotopsia는 trifocal IOL에 비해 EDOF IOL에서 감소하지 않음
621 | Friedrich (2012) | Prospective consecutive study |  | 독일 | NA | Cataract surgery, Presbyopia | 포함기준: Patients 40 years or older with a visually significant cataract or presbyopia demanding surgical correction, corneal astigmatism of less than 1.50 D 배제기준: history of glaucoma or retinal detachment, corneal disease, irregular corneal astigmatism, abnormal iris, macular degeneration or retinopathy, neurophthalmic disease, or history of ocular inflammation이 있으면 제외함 | 19명(31안) | 1 | 19명(31안) |  |  |  | 0% | 65 | 42.1% | No | Diffractive Bifocal IOL | TECNIS ZMB00 | [6 empty] | 1, 5 | 6개월 | 비구면 회절형 다초점 IOL Tecnis ZMB00은 백내장 제거 또는 노안 교정을 위한 수정체 유화술 후 원거리 및 근거리 시각 기능의 복원을 제공하며, 이는 이식된 IOL의 광학적 power와 무관하고 예측가능함
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | NCT03268746 | 한국 | NA | Cataract surgery | 포함기준: Eligible subjects were those without ocular disease that could confound study outcomes who wanted an IOL that provided near, intermediate, and distance vision. Inclusion criteria were clear intraocular media other than cataract in both eyes, calculated lens power between + 16.0 D and + 24.0 D, and preoperative or expected postoperative regular corneal astigmatism of < 1.00 D. 배제기준: clinically significant corneal abnormalities; previous corneal transplantation; ocular trauma; previous refractive surgery or refractive procedures throughout the study duration; history of concurrent retinal conditions; anterior chamber ≤2.5 mm not caused by swollen cataract; concurrent anterior or posterior segment inflammation; and expectation of ocular surgical treatment, large capsulotomy, or retinal laser treatment during the study (excluding neodymium-doped yttrium aluminum garnet [Nd:YAG] capsulotomy). | 44명(88안) | 1 | 44명(88안) |  |  |  | 15.38% | 60 | 25% | Yes | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) | [6 empty] | 1, 2, 3, 4, 5 | 3개월 | 한국인에서 수술 후 3 개월 후 삼중초점 TFNT00 IOL의 시각 성능은 모든 거리에서 양안 UCVA에 대해 0.1 logMAR 미만으로 높은 피험자 만족도와 안경 독립성을 나타냄
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | ChiCTR-ONC-17011119 | 중국 | 2016.9-2018.2 | Cataract surgery | 포함기준: aged from 50 to 85 years, no active ocular disease except cataracts, nonsevere dry eye, angle κ within 0.5 mm, no pupillary abnormality, and no surgery within 3 months. 배제기준: reduced zonular/ capsular stability, corneal astigmatism higher than 1.25 D, previous refractive surgery, poor mobility, IOL dislocation, posterior capsule opacification, or any ocular comorbidity (amblyopia, retina, or optic nerve pathology) that may hamper postoperative acuity. | 104명(141안) | 3 | 34명(47안) | 34명(47안) | 36명(47안) |  | 7.965% | 중재: 69.57, 대조1: 69.72, 대조2: 68.56 | 49.04% | No | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  |  | 1, 2, 4 | 3개월 | Tecnis Symfony ZXR00과 Lentis Comfort LS-313 MF15는 모두 뛰어난 시각 복원력과 안정된 거리 및 중간 시력, 우수한 주관적 시각 기능 및 우수한 대비 감도를 제공함. EDOF IOL은 다초점 IOL보다 더 나은 객관적인 광학 품질과 더 두드러진 dysphotopsia 증상을 보임
691 | Bohm (2019) | Prospective case series |  | 독일 | 2013-2017 | Presbyopia | 배제기준: previous ocular surgery, corneal astigmatism tomography greater than 1.50 diopters (D) measured with Scheimpflug (Pentacam, Oculus Optikger€ ate GmbH), ocular pathology or corneal abnormalities, and an endothelial cell count below 2000 cells/mm2 . | 105명(210안) | 4 | 27명(54안) | 27명(54안) | 26명(52안) | 20명(50안) | 0% | 중재: 63.41, 대조1: 63.51, 대조2: 69.23, 대조3: 66.72 | 49.52% | Yes | Diffractive Trifocal IOL | AcrySof PanOptix | Diffractive Trifocal IOL | AT LISA tri 839MP | Diffractive EDOF toric IOL | TECNIS Symfony ZXR00 | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF30 | 1, 4, 5 | 3개월 | 4 개의 IOL은 비슷하게 우수한 CDVA를 보임. EDOF IOL은 다른 IOL보다 약간 더 나은 DCIVA를 보였지만 DCNVA는 더 나빴음. panfocal IOL만이 50cm에서 더 나은 DCIVA를 보임
702 | Levinger (2019) | Interventional case series |  | 이스라엘 | 2014.1-2015.12 | Presbyopia | 포함기준: function without distance correction and had a near addition (add) of at least +1.25D 배제기준: a history of glaucoma, retinal detachment, corneal disease, corneal surgery, ocular inflammation, neuro-ophthalmic disease, macular degeneration, and retinopathy. | 26명(26안) | 1 | 26명(26안) |  |  |  | 0% | 53.8 | 57.69% | No | Diffractive Trifocal IOL | FineVision Micro F | [6 empty] | 1, 2, 3, 5 | 6개월 | 정시 노안 환자에서 FLACS 및 diffractive 삼중초점 IOL (FineVision Micro F)을 시술한 비우성안의 unilateral RLE는 이 예비 증례연구에서 합병증 없이 만족스러운 근거리, 원거리, 중간거리 시력을 제공함
732 | Vounotrypidis (2017) | Prospective longitudinal case series |  | 독일 | NA | Cataract surgery | 포함기준: age 55 years and 85 years scheduled for bilateral cataract surgery for visually significant cataract were included in the study if they desired spectacle independence and had corneal astigmatism less than 1.00 diopter (D) against the rule (ATR) and less than 1.25 D with the rule (WTR). 배제기준: patients with a history of zonular instability, previous ocular surgery, corneal detachment, disease, irregular corneal astigmatism, glaucoma or retinal maculopathies, neuro-ophthalmic disease, ocular inflammation, or pupil anomalies were excluded. Patients with a mesopic pupil size smaller than 3.0 mm were also excluded to prevent low functionality of the IOL. | 22명(44안) | 1 | 22명(44안) |  |  |  | 0% | 69.2 | 45.45% | Yes | Refractive EDOF IOL | Lentis comfort LS-313 MF15 | [6 empty] | 1, 4, 5 | 3개월 | 1.5D의 근거리 추가로 향상된 EDOF IOL은 중간거리 및 원거리의 우수한 시각 성능과 허용 가능한 근거리 복원을 보여줌. 환자 만족도가 매우 높았으며, 광현상으로 불편한 환자는 거의 없었음
782 | Chang (2019) | Prospective, observational case series |  | 홍콩 | 2012.8-2014.4 | Cataract surgery, Presbyopia | 포함기준: patients aged 40 years or older with preoperative corneal astigmatism of 1.75 diopters (D) or less. 배제기준: preexisting ocular conditions including age-related macular degeneration and glaucoma, systemic diseases such as uncontrolled diabetes mellitus that might affect the postoperative vision, Fuchs dystrophy, and a history of corneal refractive surgery. | 36명(72안) | 1 | 36명(72안) |  |  |  | 0% | 55.9 | 11.11% | Yes | Diffractive Bifocal IOL | TECNIS ZMB00 | [6 empty] | 1, 2, 3, 5 | 6개월 | 백내장 유무에 관계없이 노안환자에서 이중초점 IOL의 양측시술은 만족스러운 시력과 높은 안격독립률을 보였으며, Halos와 glare는 일반적임
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study |  | 오스트리아, 프랑스, 독일, 포르투갈, 스페인 | NA | Cataract surgery | 포함기준: significant cataract, estimated postoperative corneal astigmatism of 0.75 D or higher, and presbyopia with the wish for postoperative spectacle or contact lens independence at all distances. 배제기준: previous intraocular surgery, glaucoma, history of uveitis or retinal detachment, peripheral retinal lesions not treated prophylactically, iris atrophy, corneal disease, macular degeneration, unreal expectations and any neuro-ophthalmological disease. | 38명(57안) | 1 | 38명(57안) |  |  |  | 0% | 58.6 | NA | No | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) | [6 empty] | 1, 3, 4 | 2-4개월 | 다초점 원환제 IOL의 시술은 각막 난시가 있는 눈에서 매우 우수한 시각적 재활을 제공하는 안전한 시술임
1162 | Venter (2014) | Prospective consecutive case series |  | 남아프리카 | NA | Presbyopia | 포함기준: age of 45 years or older, corrected distance visual acuity (CDVA) of 6/6 or better in each eye, ametropia combined with presbyopia, and corneal astigmatism less than 1.50 diopter (D). 제외기준: a history of glaucoma or retinal detachment, corneal disease, corneal surgery, ocular inflammation, neuro-ophthalmic disease, and macular degeneration or retinopathy. | 53명(106안) | 1 | 53명(106안) |  |  |  | 0% | 58.2 | 39.62% | Yes | Refractive bifocal IOL | SBL-3 | [6 empty] | 1, 3, 5 | 3개월 | 새로운 refractive rotationally asymmetric IOL는 근거리, 중간거리, 원거리에 대해 좋은 시력을 제공함. 이 IOL 성능을 평가하려면 장기간 추적 관찰이 필요함
1220 | Mojzis (2014) | Case series |  | 스페인 | NA | Presbyopia | 포함기준: refractive lens exchange for presbyopia, bilateral surgery, astigmatism of less than 1.0 diopter (D), uneventful surgery, and IOL power between 0.0 D and +32.0D 배제기준: previous ocular surgery, ocular disease, complications during surgery, and corneal astigmatism of more than 1.0 D. | 30명(60안) | 1 | 30명(60안) |  |  |  | NA | 57.9 | NA | Yes | Diffractive trifocal IOL | AT Lisa tri 839MP | [6 empty] | 1, 3, 4, 5 | 6개월 | 노안환자에서 삼중초점 IOL은 근거리, 중간거리, 원거리 시력을 개선하고 3 foci 사용은 근거리, 원거리 시력 손실없이 유의한 중간거리 시력개선을 제공함
1288 | Tsaousis (2013) | Observational case study |  | 그리스 | NA | Cataract surgery | 포함기준: age between 55 and 80 years, uncomplicated phacoemulsification, independence from spectacle use following cataract surgery, implantation of the same IOL bilaterally, and at least 15 months following the operation in the second eye. 제외기준: any type of fundus pathology, corneal astigmatism of more than 1.50 diopters (D), and deviation from the target refraction of more than 1.00 D. | 20명(40안) | 1 | 20명(40안) |  |  |  | 0% | 70 | 45% | Yes | Diffractive multifocal IOL | AcrySof IQ ReSTOR IOL SN60D3 | [6 empty] | 1 | 26개월 | 결과는 시력 개선 측면에서 양측 다초점 IOL 이식을 한 환자에서 양안 시력의 상당한 이점이 있음을 확인하였으며 이 효과는 모든 거리에서 나타남
1350 | Ferrer-Blasco (2012) | Prospective study |  | 스페인 | NA | Presbyopia | 포함기준: age between 45 and 70 years and the motivation to reduce the dependency of any form of spectacle or contact lens correction for distance and near. 제외기준: ≥1 diopters (D) of corneal astigmatism, history of glaucoma or retinal detachment (RD), corneal disease, and previous corneal or intraocular surgery. | 30명(60안) | 1 | 30명(60안) |  |  |  | 0% | Myopic group: 50.83, Hyperopic group: 52.14 | 65% | Yes | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 | [6 empty] | 1, 5 | 6개월 | RLE시 ReSTOR ® SN6AD3 IOL는 높은 ametropia 및 노안 교정에 있어 우수한 안전성, 효능 및 예측 가능성을 보여줌
1409 | Altaie (2012) | Prospective, observational study |  | 뉴질랜드 | 2004.12-2008.3 | Cataract or CLE | 포함기준: (i) motivation, that is, the desire to no longer require (or minimize) the use of spectacle or contact lens correction for distance and near vision; (ii) the ability to return for follow-up assessments up to 6 months post surgery; and (iii) the ability to provide informed consent. 배제기준: amblyopia, history of glaucoma, corneal disease, previous corneal or intraocular surgery, abnormalities of the iris or pupil distortion, large ambient pupil diameter (>5.00 mm), macular degeneration, other retinal disease, neuro-ophthalmic diseases and a history of ocular inflammation. | 203명(363안) | 1 | 203명(363안) |  |  |  | 0% | Cataract group: 66.53, CLE group: 55.86 | 42.7% | No | Diffractive bifocal IOL | AcrySof Natural ReSTOR IOL (SN60D3) | [6 empty] | 1, 3, 4, 5 | 6개월 | apodized, diffractive 다초점 IOL는 예상대로 우수한 near uncorrected visual acuity과 distance uncorrected visual acuity, spectacle independence을 제공함. 소수의 피험자들은 IOL 설계와 관련이 있을 수 있는 halos, other visual phenomena을 보고함
1485 | Alfonso (2011) | Prospective study |  | 스페인 | NA | Presbyopia | 포함기준: age between 35 and 70 years and the desire not to wear any form of spectacle or contact lens correction for distance and near vision. 배제기준: included cataract, BCVA <20/25 in either eye, preoperative astigmatism >1.50 D, history of glaucoma or retinal detachment (RD), corneal disease, previous corneal or intraocular surgery, abnormal iris, pupil deformation, macular degeneration or retinopathy, neuroophthalmic diseases, and history of ocular inflammation. | 33명(66안) | 1 | 33명(66안) |  |  |  | 0% | 53.26 | 18.18% | Yes | Diffractive bifocal IOL | Acri.LISA 366D | [6 empty] | 1, 5 | 6개월 | RLE 후 Acri.LISA 366D 이중초점 IOL의 양측이식은 ametropia 및 노안교정을 위한 효과적인 시술임
1544 | Alfonso (2010) | Prospective study |  | 스페인 | NA | Presbyopia | 포함기준: age between 35 and 70 years and the desire not to wear any form of spectacle or contact lens correction for distance and near vision. 배제기준: cataract, best-corrected distance visual acuity (BCVA) < 20/25 in either eye, preoperative astigmatism > 1.50 diopters (D), history of glaucoma or retinal detachment (RD), corneal disease, previous corneal or intraocular surgery, abnormal iris, pupil deformation, macular degeneration or retinopathy, neuro-ophthalmic diseases, and history of ocular inflammation. | 17명(34안) | 1 | 17명(34안) |  |  |  | 0% | 52.9 | 17.65% | Yes | Diffractive bifocal IOL | 447D IOL | [6 empty] | 1, 5 | 6개월 | RLE 후 distance-dominant diffractive bifocal IOL 시술 6개월 결과는 ametropia 및 노안 교정에서 효능, 안전성 및 예측 가능성을 입증함
1556 | Alfonso (2010) | Prospective study |  | 스페인 | NA | Presbyopia | 포함기준: age between 40 and 75 years and motivation (the desire to no longer wear any form of spectacle or contact lens correction for distance and near vision). 배제기준: included cataract, best-corrected distance VA (BCVA) <20/25 in either eye, preoperative astigmatism >1.50 diopters (D), history of glaucoma or retinal detachment (RD), corneal disease, previous corneal or intraocular surgery, abnormal iris, pupil deformation, macular degeneration or retinopathy, neuro-ophthalmic diseases, and history of ocular inflammation. | 65명(130안) | 1 | 65명(130안) |  |  |  | 0% | 52.3 | 23.08% | Yes | Diffractive bifocal IOL | Acri.Twin | [6 empty] | 1, 5 | 6개월 | Acri.Twin bifocal IOL의 asymmetric bilateral implantation을 통한 RLE는 ametropia 및 노안 교정을 위한 효과적인 시술임. 본 연구의 결과는 근시 및 원시 환자 모두에게 원거리 및 근거리에서 안전하고 효과적이었음
3093 | Tan (2019) | Prospective study |  | 중국 | 2016.7~ | Cataract surgery, Presbyopia | 포함기준: diagnosis of age-related cataract and/or presbyopia, correctable ametropia with or without cataract, demand for spectacle independenc, and reasonable expectations regaring the postoperative outcome 배제기준: astigmatism of more than 1.5D, keratoconus or other keratopathies affecting visual acuity, fundus lesions, pupil diameter is more than 6mm or less than 2 mm under natural light, professional drivers and those who often drive at night, history of previous eye surgery, patient suffering anxiety or paranoia, and patients with systemic disease | 40명(55안) | 1 | 40명(55안) |  |  |  | 0% | 70.82 | 55% | No | Diffractive EDOF IOL | Tecnis Symfony IOL | [6 empty] | 1, 2, 3, 4, 5 | 3개월 | TECNIS Symforny는 원거리, 중간거리 시력, 및 functional range near 시력개선을 보여주고, 안경독립성, 환자만족도가 높음
3220 | Almulhim (2018) | Prospective nonrandomized study |  | 사우디아라비아 | 2016.12-2017.6 | Cataract surgery, Presbyopia | 포함기준: The included patients had cataract or presbyopia/pre- presbyopia suitable for refractive lens exchange and were seeking spectacle independence. The excluded patients had the defects such as history of glaucoma, retinal detach- ment, corneal disease, irregular corneal astigmatism, abnor- mal iris, macular degeneration, advanced retinopathy, neuro ophthalmic disease, the history of ocular inflammation or previous ocular surgery. Eight eyes with regular astigma- tism of >1.25 D were included and underwent a separate analysis of refractive outcome as they have extreme cylinder readings. | 41명(82안) | 1 | 41명(82안) |  |  |  | 9% | NA | 48.78% | Yes | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP | [6 empty] | 1, 3, 4, 5 | 3개월 | Diffractive trifocal IOL은 시력과 관련된 일상 활동의 수행에 긍정적인 영향과 함께 환자에게 방해가 되지 않는 광현상을 최소화하면서 만족스러운 시각 및 굴절 결과를 제공함
3241 | Zamora-De-La-Cruz (2018) | Observational longitudinal and prosective study |  | 멕시코 | 2016.3-2016.8 | Cataract surgery, Presbyopia | 포함기준: without ocular comorbidities, with astigmatism <1 dipoter, with no history of previous eye surgery and that could comply with long-term follow up 배제기준: patients who could not be implanted bilaterally with the same IOL of who had alterations that conditioned bias of the results were excluded | 12명(24안) | 2 | 7명(14안) | 5명(10안) |  |  | 0% | NA | NA | Yes | Diffractive bifocal IOL | ReSTOR +2.50 | Diffractive trifocal IOL | Acrysof IQ PanOptix |  |  |  |  | 1, 3, 4 | NA | 두 렌즈 모두 환자들사이에 좋은 수용성을 보임. 현재 두 렌즈 중 하나를 결정하는 수용가능한 하나의 기준은 환자의 나이, 활동에 따른 시력요구도임
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series |  | 스페인 | NA | Cataract surgery | 포함기준: 45-87세, patients with incipient or moderate cataract reporting a significant reduction in visual quality or presbyopic patients suitable for refractive lens exchange and refractive astigmastism <-2.00 D. 배제기준: history of glaucoma or retinal detachment, corneal disease, previous corneal or intraocular surgery, an abnormal iris, pupil deformation, macular degeneration, or retinopathy, neuroopthalmic disease, and history of ocular inflammation. | 60명(120안) | 4 | 15명(30안) | 15명(30안) | 15명(30안) | 15명(30안) | 0% | 중재: 63.07, 대조1: 66.78, 대조2: 62.15, 대조3: 67.14 | NA | Yes | Diffractive trifocal IOL | AT LISA tri 839MP | Diffractive trifocal IOL | FineVision | Diffractive bifocal IOL | AcrySof ReSTOR SN6AD1 | monofocal IOL | Acri.Smart 48S | 1 | 3개월 | 다초점 IOL은 좋은 원거리, 근거리 및 중간거리 시력을 제공함. AT LISA tri MP839는 FineVision 및 AcrySof ReSTOR SN6AD1 IOL 보다 중간거리 시력에서 더 나은 결과를 보여주었음. 각 렌즈의 디포커스 곡선 프로파일을 알면 의사는 최상의 IOL을 선택할 수 있음
4009 | Venter (2013) | Prospective Case series study |  | 영국 | 2010.6-2011.1 | ametropic Presbyopia | 포함기준: The study cohort comprised presbyopic ametropic patients who had cataract or were not candidates for laser vision correction. All eyes had naturally occurring keratometric cylinder greater than 1.50 D. 배제기준: history of glaucoma or retinal detachment, corneal disease, previous corneal surgery, history of ocular inflammation, neuro-ophthalmic disease, macular degeneration, and retinopathy | 58명(89안) | 1 | 58명(89안) |  |  |  | 0% | 54.8 | 60% | No | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT | [6 empty] | 1, 3, 4, 5 | 3개월 | surface-embedded near section이 있는 multifocal toric IOL은 기존의 각막 난시를 효과적으로 교정함, IOL은 수정체 낭에서 안정적이었고 시각적으로 유의한 합병증은 없었으며 수술 후 3 개월까지 유의한 rotation out이 없었음
4240 | Van Der Linden (2012) | Comparative case series |  | 네덜란드 | NA | Presbyopia | 포함기준: multifocal IOL implantation were ocular pathology other than cataract (eg, dry-eye syndrome, macular disease) diabetes with or without retinopathy, corneal astigmatism more than 1.75 diopters (D), and an American Society of Anesthesiologists classification of III or higher | 117명(233안) | 2 | 45명(90안) | 72명(143안) |  |  | 0% | 중재: 61.8, 대조: 59.7 | 42.49% | Yes | Refractive bifocal IOL | Lentis Mplus LS-312 | Diffractive bifocal IOL | ReSTOR SN6AD1 |  |  |  |  | 1, 4, 5 | 3개월 | 새로운 부문 별 추가 다초점 IOL은 UDVA 및 dysphotopsia 측면에서 diffractive apodized 다초점 IOL과 비교하였음. diffractive apodized 다초점 IOL은 30cm 및 40cm 거리에서의 시력이 더 좋았고 환자 만족도가 더 높았음
2 | Kim (2019) | Retrospective review | NA | 한국 | 2009.11-2017.4 | Cataract surgery | Cataract patients who underwent uneventful phacoemulsification and multifocal IOL implantation during the study period were included in the sample. Those with histories of trauma, corneal opacity, and severe dry eye were excluded. Cases with intraoperative complications were also excluded. | 63명(88안) | 3 | 53안 | 18안 |  | 대조(17안) | NA | 57.93 | 34.09% |  | Trifocal IOL | AT LISA tri 839MP | Bifocal IOL | AcrySof IQ ReSTOR SN6AD1 | Bifocal IOL | Tecnis MF ZLB00 |  |  |  | 1개월 | Trifocal IOL과 Bifocal IOL 모두 시력향상에 효과적이지만 Trifocal IOL 환자의 굴절이상은 근시경향을 보여줌
10 | Kretz (2016) | Prospective study | NA | 독일 | NA | cataract or presbyopia/ pre-presbyopia | Exclusion: patients had a history of glaucoma or retinal detachment, corneal disease, irregular corneal astigmatism, abnormal iris, macular degeneration or retinopathy, neurophthalmic disease, history of ocular inflammation, or previous ocular surgery or corneal astigmatism of 1.25 diopter (D) or greater | 50명 | 1 | 50명 | NA |  | NA | NA | 57.5 | NA |  | TrifocalIOL | T LISA tri 839MP | NA | [6 empty] | 3개월
28 | 곽준영 (2012) | Prospective study | NA | 한국 | 2010.9-2010.12 | Cataract surgery | 40-80세 사이의 연령으로 노인성 백내장 진단을 받고 술후 돋보기 착용을 원치 않는 환자들을 대상. 각막 난시는 -1.0디옵터 미만, 녹내장이나 망막 박리, 각막 혼탁, 각막 또는 안내 수술의 기왕력, 홍채 이상, 황반 변성을 포함한 시력에 영향을 줄 수 있는 망막질환이나 각막 굴절수술을 시행 받은 사람은 제외 직업적인 야간운전이나 금속 세공과 같은 정밀한 근거리 작업을 요하는 사람이나 중간거리와 근거리를 포함한 모든 다양한 거리를 원거리 시력처럼 선명하게 보기를 원하는 시력에 대한 기대치가 지나치게 높은 사람은 제외 | 20명(40안) | 1 | 20명(40안) | NA | NA | NA | NA | 57.2 | 45% | Y | Multifocal IOL | Acri.LISA 366D | [7 empty] | 3개월

## 결과지표_연속형_IOL
2. 연속형 결과변수
no. | 1저자 (연도) | 연구설계 | 대상자 | 중재군명 | 렌즈명 | 대조군명 | 렌즈명 | 하위그룹 | 결과그룹 | 결과지표명 | 시력구분(Uncorrected, corrected) | 시력구분(Distant, Intermediate, Near) | 단위(LogMAR, Snellen, et al.,) | 안구(Biocular, monocular) | 비고 | 측정시점 (개월) | 분석단위(명/안) | 중재군 |  |  | 대조군 |  |  | 최종-기저평균차(변화량) |  |  |  | 두 군간 차이 | 비고
[18 empty] | N | mean | SD | N | mean | SD | 중재군 | 변화량 SD | 대조군 | 변화량 SD | P value
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 수술전 |  | 100 | 0.72 | NA
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 3개월 |  | 100 | 0.06 | NA | [7 empty] | <0.001 | p-value로 SD 추정
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular |  | 수술전 |  | 50 | -
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular |  | 3개월 |  | 50 | 0.04 | NA
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | 수술전 |  | 100 | 0.16 | NA
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | 3개월 |  | 100 | 0.04 | NA | [7 empty] | <0.001 | p-value로 SD 추정
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UNVA | uncorrected | Near | LogMAR | Monocular |  | 수술전 |  | 100 | -
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UNVA | uncorrected | Near | LogMAR | Monocular |  | 3개월 |  | 100 | 0.06 | NA
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UNVA | uncorrected | Near | LogMAR | Binocular |  | 수술전 |  | 50 | -
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UNVA | uncorrected | Near | LogMAR | Binocular |  | 3개월 |  | 50 | 0.01 | NA
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UIVA | uncorrected | Intermediate | LogMAR | Monocular |  | 수술전 |  | 100 | -
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UIVA | uncorrected | Intermediate | LogMAR | Monocular |  | 3개월 |  | 100 | 0.09 | NA
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UIVA | uncorrected | Intermediate | LogMAR | Binocular |  | 수술전 | 명 | 50 | -
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UIVA | uncorrected | Intermediate | LogMAR | Binocular |  | 3개월 | 명 | 50 | 0.04 | NA
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | Spectacle independece | Reading |  |  | 0-40점(0점 best) |  |  | 3개월 | 명 | 50 | 10.33 | 12.47
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | Spectacle independece | Computer work |  |  | 0-40점(0점 best) |  |  | 3개월 | 명 | 50 | 5.71 | 11.9
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | Spectacle independece | Distant |  |  | 0-40점(0점 best) |  |  | 3개월 | 명 | 50 | 3.92 | 9.77
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  |  | VA | BCVA | best-corrected | Distant(4m) | LogMAR | Binocular |  | 수술전 | 명 | 20 | 0.31 | 0.23
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular |  | 1개월 | 명 | 20 | 0.09 | 0.13
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular |  | 6개월 | 명 | 20 | 0.1 | 0.11
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Binocular |  | 1개월 | 명 | 20 | 0.08 | 0.11
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Binocular |  | 6개월 | 명 | 20 | 0.06 | 0.14
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  |  | VA | UIVA | uncorrected | Intermediate(70cm) | LogMAR | Binocular |  | 1개월 | 명 | 20 | 0.27 | 0.15
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  |  | VA | UIVA | uncorrected | Intermediate(70cm) | LogMAR | Binocular |  | 6개월 | 명 | 20 | 0.26 | 0.14
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular |  | 1개월 | 명 | 20 | 0.07 | 0.12
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular |  | 6개월 | 명 | 20 | 0.08 | 0.14
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  |  | VA | DCNVA | distance corrected | Near(33cm) | LogMAR | Binocular |  | 1개월 | 명 | 20 | 0.07 | 0.11
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  |  | VA | DCNVA | distance corrected | Near(33cm) | LogMAR | Binocular |  | 6개월 | 명 | 20 | 0.06 | 0.1
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  | photopic | CS |  |  |  | log | Monocular | 1.5CPD | 1개월 | 안 | 40 | 1.38 | 0.32
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  | photopic | CS |  |  |  | log | Monocular | 3CPD | 1개월 | 안 | 40 | 1.73 | 0.24
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  | photopic | CS |  |  |  | log | Monocular | 6CPD | 1개월 | 안 | 40 | 1.75 | 0.21
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  | photopic | CS |  |  |  | log | Monocular | 12CPD | 1개월 | 안 | 40 | 1.33 | 0.18
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  | photopic | CS |  |  |  | log | Monocular | 18CPD | 1개월 | 안 | 40 | 0.77 | 0.27
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  | mesopic | CS |  |  |  | log | Monocular | 1.5CPD | 1개월 | 안 | 40 | 1.37 | 0.27
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  | mesopic | CS |  |  |  | log | Monocular | 3CPD | 1개월 | 안 | 40 | 1.72 | 0.16
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  | mesopic | CS |  |  |  | log | Monocular | 6CPD | 1개월 | 안 | 40 | 1.63 | 0.14
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  | mesopic | CS |  |  |  | log | Monocular | 12CPD | 1개월 | 안 | 40 | 1.01 | 0.19
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  | mesopic | CS |  |  |  | log | Monocular | 18CPD | 1개월 | 안 | 40 | 0.5 | 0.22
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  |  | Satisfaction | 근거리시력 |  |  | 0-300점(300점 best) |  |  | 1개월 | 명 | 20 | 244 | 36
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  |  | Satisfaction | 원거리시력 |  |  | 0-300점(300점 best) |  |  | 1개월 | 명 | 20 | 256 | 21
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | Diffractive Bifocal IOL | Acri.LISA 366D |  |  |  | Satisfaction | 운전 |  |  | 0-300점(300점 best) |  |  | 1개월 | 명 | 20 | 221 | 42
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Binocular |  | 4-6개월 | 명 | 96 | -0.04 | 0.07
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular |  | 4-6개월 | 명 | 96 | -0.02 | 0.07
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  |  | VA | UIVA | uncorrected | Intermediate(70cm) | LogMAR | Binocular |  | 4-6개월 | 명 | 96 | 0.04 | 0.07
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | 4-6개월 | 명 | 96 | 0.17 | 0.14
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  | Toric in a least one eye | VA | CDVA | corrected | Distant(4m) | LogMAR | Binocular |  | 4-6개월 | 명 | 96 | -0.04 | 0.07
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  | Toric in a least one eye | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular |  | 4-6개월 | 명 | 96 | 0 | 0.09
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  | Toric in a least one eye | VA | UIVA | uncorrected | Intermediate(70cm) | LogMAR | Binocular |  | 4-6개월 | 명 | 96 | 0.04 | 0.06
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  | Toric in a least one eye | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | 4-6개월 | 명 | 96 | 0.13 | 0.12
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  |  | Satisfaction |  |  | Distant(4m) | 0-10점(10점 best) | Binocular |  | 4-6개월 | 명 | 96 | 9 | NA | [8 empty] | median
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  |  | Satisfaction |  |  | Intermediate(70cm) | 0-10점(10점 best) | Binocular |  | 4-6개월 | 명 | 96 | 9.5 | NA | [8 empty] | median
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  |  | Satisfaction |  |  | near(40cm) | 0-10점(10점 best) | Binocular |  | 4-6개월 | 명 | 96 | 8 | NA | [8 empty] | median
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  | Toric | Satisfaction |  |  | Distant(4m) | 0-10점(10점 best) | Binocular |  | 4-6개월 | 명 | 11 | 9 | NA | [8 empty] | median
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  | Toric | Satisfaction |  |  | Intermediate(70cm) | 0-10점(10점 best) | Binocular |  | 4-6개월 | 명 | 11 | 9 | NA | [8 empty] | median
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  | Toric | Satisfaction |  |  | near(40cm) | 0-10점(10점 best) | Binocular |  | 4-6개월 | 명 | 11 | 7 | NA | [8 empty] | median
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular |  | 수술전 | 명 | 30 | 0.46 | 0.29
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular |  | 3개월 | 명 | 30 | 0 | 0.09
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Binocular |  | 수술전 | 명 | 30 | 0.29 | 0.14
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Binocular |  | 3개월 | 명 | 30 | -0.08 | 0.09
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  | monocular SRD | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 30 | 0.437 | 0.12
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  | monocular SRD | VA | UNVA | uncorrected | Near(subjective) | LogMAR | Monocular |  | 3개월 | 안 | 30 | 0.336 | 0.165
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  | monocular SRD | VA | DCNVA | distance corrected | near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 30 | 0.385 | 0.195
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  | monocular SRD | VA | DCNVA | distance corrected | Near(subjective) | LogMAR | Monocular |  | 3개월 | 안 | 30 | 0.316 | 0.21
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  | monocular SRD | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 30 | 0 | 0.095
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  | monocular SRD | VA | UIVA | uncorrected | Intermediate(subjective) | LogMAR | Monocular |  | 3개월 | 안 | 30 | -0.02 | 0.088
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  | monocular SRD | VA | DCIVA | distance corrected | intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 30 | -0.02 | 0.107
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  | monocular SRD | VA | DCIVA | distance corrected | intermediate(subjective) | LogMAR | Monocular |  | 3개월 | 안 | 30 | -0.1 | 0.113
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular |  | 수술전 | 안 | 44 | 0.47 | 0.38
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 1개월 | 안 | 44 | 0.23 | NA
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 44 | 0.21 | NA
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 6개월 | 안 | 44 | 0.2 | NA
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Monocular |  | 1개월 | 안 | 44 | 0.22 | NA
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 44 | 0.2 | NA
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Monocular |  | 6개월 | 안 | 44 | 0.22 | NA
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 1개월 | 안 | 44 | 0.02 | NA
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 44 | 0.01 | NA
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 6개월 | 안 | 44 | 0.01 | NA
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | 1개월 | 안 | 44 | 0.16 | NA
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | 3개월 | 안 | 44 | 0.15 | NA
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | 6개월 | 안 | 44 | 0.15 | NA
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Binocular |  | 1개월 | 안 | 44 | 0.12 | NA
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Binocular |  | 3개월 | 안 | 44 | 0.11 | NA
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Binocular |  | 6개월 | 안 | 44 | 0.13 | NA
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular |  | 1개월 | 안 | 44 | 0 | NA
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular |  | 3개월 | 안 | 44 | 0 | NA
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular |  | 6개월 | 안 | 44 | 0 | NA
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis | Diffractive Bifocal IOL | AcrySof ReSTOR D1 |  | VA | UCVA | uncorrected |  | LogMAR | Binocular |  | 수술전 | 안 | 20 | 0.41 | 0.09 | 20 | 0.45 | 0.18
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis | Diffractive Bifocal IOL | AcrySof ReSTOR D1 |  | VA | BCVA | best-corrected |  | LogMAR | Binocular |  | 수술전 | 안 | 20 | 0.35 | 0.08 | 20 | 0.33 | 0.15
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis | Diffractive Bifocal IOL | AcrySof ReSTOR D1 |  | VA | UCVA | uncorrected | Distant(5m) | LogMAR | Binocular |  | 2개월 | 안 | 20 | 0.08 | 0.07 | 20 | 0.09 | 0.11
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis | Diffractive Bifocal IOL | AcrySof ReSTOR D1 |  | VA | BCVA | best-corrected | Distant(5m) | LogMAR | Binocular |  | 2개월 | 안 | 20 | 0.03 | 0.08 | 20 | 0.01 | 0.02
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis | Diffractive Bifocal IOL | AcrySof ReSTOR D1 |  | VA | UCVA | uncorrected | Intermediate(70cm) | LogMAR | Binocular |  | 2개월 | 안 | 20 | 0.11 | 0.08 | 20 | 0.32 | 0.17
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis | Diffractive Bifocal IOL | AcrySof ReSTOR D1 |  | VA | DCVA | distance corrected | Intermediate(70cm) | LogMAR | Binocular |  | 2개월 | 안 | 20 | 0.11 | 0.08 | 20 | 0.37 | 0.21
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis | Diffractive Bifocal IOL | AcrySof ReSTOR D1 |  | VA | BCVA | best-corrected | Intermediate(70cm) | LogMAR | Binocular |  | 2개월 | 안 | 20 | 0.04 | 0.11 | 20 | 0.15 | 0.13
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis | Diffractive Bifocal IOL | AcrySof ReSTOR D1 |  | VA | UCVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | 2개월 | 안 | 20 | 0.09 | 0.1 | 20 | 0.17 | 0.12
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis | Diffractive Bifocal IOL | AcrySof ReSTOR D1 |  | VA | DCVA | distance corrected | near(40cm) | LogMAR | Binocular |  | 2개월 | 안 | 20 | 0.06 | 0.07 | 20 | 0.16 | 0.13
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis | Diffractive Bifocal IOL | AcrySof ReSTOR D1 |  | VA | BCVA | best-corrected | near(40cm) | LogMAR | Binocular |  | 2개월 | 안 | 20 | 0.05 | 0.07 | 20 | 0.06 | 0.09
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis | Diffractive Bifocal IOL | AcrySof ReSTOR D1 |  | Satisfaction |  | uncorrected | Near(40cm) | 1-5점(5점 best) | Binocular |  | 2개월 | 명 | 11 | 4.57 | NA | 10 | NA | NA
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis | Diffractive Bifocal IOL | AcrySof ReSTOR D1 |  | Satisfaction |  | uncorrected | Intermediate(70cm) | 1-5점(5점 best) | Binocular |  | 2개월 | 명 | 11 | 4.79 | NA | 10 | NA | NA
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis | Diffractive Bifocal IOL | AcrySof ReSTOR D1 |  | Satisfaction |  | uncorrected | Distant(5m) | 1-5점(5점 best) | Binocular |  | 2개월 | 명 | 11 | 4.61 | NA | 10 | NA | NA
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis | Diffractive Bifocal IOL | AcrySof ReSTOR D1 |  | Satisfaction |  | uncorrected |  | 1-5점(5점 best) | Binocular |  | 2개월 | 명 | 11 | 4.67 | NA | 10 | 4.28 | NA
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | UDVA | uncorrected | Distant(6m) | Snellen decimal | Binocular |  | 6개월 | 안 | 46 | 0.95 | 0.09
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | CDVA | corrected | Distant(6m) | Snellen decimal | Binocular |  | 수술전 | 안 | 46 | 0.95 | 0.13
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | CDVA | corrected | Distant(6m) | Snellen decimal | Binocular |  | 6개월 | 안 | 46 | 0.95 | 0.12
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | UNVA | uncorrected | Near(33cm) | Snellen decimal | Binocular |  | 6개월 | 안 | 46 | 0.95 | 0.07
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | DCNVA | distance corrected | Near(33cm) | Snellen decimal | Binocular |  | 수술전 | 안 | 46 | 0.92 | 0.09
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | DCNVA | distance corrected | Near(33cm) | Snellen decimal | Binocular |  | 6개월 | 안 | 46 | 0.99 | 0.09
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | Satisfaction | Far vision |  |  | 0-9점(9점 best) |  |  | 6개월 | 명 | 23 | 8.3 | 0.6
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | Satisfaction | intemediate vision |  |  | 0-9점(9점 best) |  |  | 6개월 | 명 | 23 | 7.1 | 1.4
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | Satisfaction | near vision |  |  | 0-9점(9점 best) |  |  | 6개월 | 명 | 23 | 8.1 | 1.5
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | Satisfaction | night vision |  |  | 0-9점(9점 best) |  |  | 6개월 | 명 | 23 | 7.7 | 1.3
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | Satisfaction | global satisfcation |  |  | 0-9점(9점 best) |  |  | 6개월 | 명 | 23 | 8.5 | 1.2
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | Satisfaction | glare |  |  | 1-4점(1점 best) |  |  | 6개월 | 명 | 23 | 1.2 | 0.6
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | Satisfaction | halos |  |  | 1-4점(1점 best) |  |  | 6개월 | 명 | 23 | 2.5 | 0.5
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | first implanted eye | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 93 | 0.371 | 0.204 | 89 | 0.376 | 0.26
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | second implanted eye | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 93 | 0.238 | 0.17 | 89 | 0.233 | 0.224
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | 6개월 | 명 | 86 | 0.082 | 0.064, 0.101 | 81 | 0.136 | 0.113, 0.159 | [5 empty] | 90% CI로 SD 계산하기
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular |  | 6개월 | 명 | 86 | 0.049 | 0.026, 0.072 | 81 | 0.116 | 0.092, 0.139 | [5 empty] | 90% CI로 SD 계산하기
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular |  | 6개월 | 명 | 86 | 0.014 | -0.004, 0.031 | 81 | 0.003 | -0.018, 0.024 | [5 empty] | 90% CI로 SD 계산하기
272 | Pedrotti (2017) | Prospective single-blind comparative study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF | Monofocal IOL | Tecnis 1-piece |  | VA | UNVA | uncorrected | Near(30cm) | LogMAR | Binocular |  | 12개월 | 명 | 21 | 0.54 | 0.09 | 21 | 0.58 | 0.08
272 | Pedrotti (2017) | Prospective single-blind comparative study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF | Monofocal IOL | Tecnis 1-piece |  | VA | UIVA | uncorrected | Intermediate(50cm) | LogMAR | Binocular |  | 12개월 | 명 | 21 | 0.24 | 0.07 | 21 | 0.33 | 0.09
272 | Pedrotti (2017) | Prospective single-blind comparative study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF | Monofocal IOL | Tecnis 1-piece |  | VA | UIVA | uncorrected | Intermediate(70cm) | LogMAR | Binocular |  | 12개월 | 명 | 21 | 0.05 | 0.05 | 21 | 0.2 | 0.08
272 | Pedrotti (2017) | Prospective single-blind comparative study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF | Monofocal IOL | Tecnis 1-piece |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular |  | 12개월 | 명 | 21 | -0.01 | 0.06 | 21 | 0.02 | 0.05
272 | Pedrotti (2017) | Prospective single-blind comparative study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF | Monofocal IOL | Tecnis 1-piece |  | VA | CNVA | corrected | near(30cm) | LogMAR | Binocular |  | 12개월 | 명 | 21 | 0.01 | 0.03 | 21 | 0 | 0.02
272 | Pedrotti (2017) | Prospective single-blind comparative study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF | Monofocal IOL | Tecnis 1-piece |  | VA | CIVA | corrected | intermediate(50cm) | LogMAR | Binocular |  | 12개월 | 명 | 21 | 0.02 | 0.03 | 21 | 0.01 | 0.03
272 | Pedrotti (2017) | Prospective single-blind comparative study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF | Monofocal IOL | Tecnis 1-piece |  | VA | CIVA | corrected | Intermediate(70cm) | LogMAR | Binocular |  | 12개월 | 명 | 21 | -0.01 | 0.02 | 21 | 0 | 0.02
272 | Pedrotti (2017) | Prospective single-blind comparative study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF | Monofocal IOL | Tecnis 1-piece |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Binocular |  | 12개월 | 명 | 21 | -0.02 | 0.05 | 21 | -0.01 | 0.04
323 | Gundersen (2016) | Randomized prospective study | Cataract surgery | Diffractive Trifocal toric IOL | FineVision IOL (PhysIOL ® ; FineVision Toric, Liège, Belgium) | Diffractive bifocal toric IOL | ReSTOR SND1T |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 22 | 0.03 | -0.1 | 22 | 0.08 | 0.13
323 | Gundersen (2016) | Randomized prospective study | Cataract surgery | Diffractive Trifocal toric IOL | FineVision IOL (PhysIOL ® ; FineVision Toric, Liège, Belgium) | Diffractive bifocal toric IOL | ReSTOR SND1T |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 22 | -0.01 | -0.06 | 22 | 0.01 | 0.07
323 | Gundersen (2016) | Randomized prospective study | Cataract surgery | Diffractive Trifocal toric IOL | FineVision IOL (PhysIOL ® ; FineVision Toric, Liège, Belgium) | Diffractive bifocal toric IOL | ReSTOR SND1T |  | VA | LCVA |  |  | LogMAR | Monocular | low-contrast (10%) VA | 3개월 | 안 | 22 | 0.32 | 0.13 | 22 | 0.35 | 0.09
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | VA | CDVA | corrected | Distant(3m) | LogMAR | Monocular |  | 수술전 | 안 | 40 | 0.44 | 0.05 | 40 | 0.4 | 0.07 | [5 empty] | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | VA | CDVA | corrected | Distant(3m) | LogMAR | Monocular |  | 수술전 | 안 | 40 | 0.44 | 0.05 | 40 | 0.45 | 0.05 | [5 empty] | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | VA | UDVA | uncorrected | Distant(3m) | LogMAR | Monocular |  | 4개월 | 안 | 40 | 0 | 0.04 | 40 | 0.03 | 0.05 | [5 empty] | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | VA | CDVA | corrected | Distant(3m) | LogMAR | Monocular |  | 4개월 | 안 | 40 | -0.01 | 0.01 | 40 | -0.01 | 0.02 | [5 empty] | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | VA | UIVA | uncorrected | Intermediate(67cm) | LogMAR | Monocular |  | 4개월 | 안 | 40 | 0.23 | 0.07 | 40 | 0.27 | 0.08 | [5 empty] | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | VA | DCIVA | distance corrected | Intermediate(67cm) | LogMAR | Monocular |  | 4개월 | 안 | 40 | 0.13 | 0.07 | 40 | 0.16 | 0.07 | [5 empty] | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 4개월 | 안 | 40 | 0.02 | 0.06 | 40 | 0.07 | 0.08 | [5 empty] | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | VA | DCNVA | distance corrected | near(40cm) | LogMAR | Monocular |  | 4개월 | 안 | 40 | 0.01 | 0.04 | 40 | 0.07 | 0.07 | [5 empty] | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | VA | UDVA | uncorrected | Distant(3m) | LogMAR | Monocular |  | 4개월 | 안 | 40 | 0 | 0.04 | 40 | 0.02 | 0.06 | [5 empty] | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | VA | CDVA | corrected | Distant(3m) | LogMAR | Monocular |  | 4개월 | 안 | 40 | -0.01 | 0.01 | 40 | -0.01 | 0.02 | [5 empty] | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | VA | UIVA | uncorrected | Intermediate(67cm) | LogMAR | Monocular |  | 4개월 | 안 | 40 | 0.23 | 0.07 | 40 | 0.42 | 0.09 | [5 empty] | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | VA | DCIVA | distance corrected | Intermediate(67cm) | LogMAR | Monocular |  | 4개월 | 안 | 40 | 0.13 | 0.07 | 40 | 0.29 | 0.11 | [5 empty] | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 4개월 | 안 | 40 | 0.02 | 0.06 | 40 | 0.38 | 0.1 | [5 empty] | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | VA | DCNVA | distance corrected | near(40cm) | LogMAR | Monocular |  | 4개월 | 안 | 40 | 0.01 | 0.04 | 40 | 0.32 | 0.09 | [5 empty] | 3군으로 중재군 n수 50% 입력필요
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | Apodized diffractive multifocal IOL | AcrySof Restor SN6AD3 (Alcon) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) |  | VA | UDVA | uncorrected | Distant(6m) | LogMAR | Monocular |  | 수술전 | 안 | 78 | 0.62 | 0.33 | 72 | 0.8 | 0.48
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | Diffractive bifocal IOL | Acri.Lisa 366D (Carl Zeiss Meditec AG) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) |  | VA | UDVA | uncorrected | Distant(6m) | LogMAR | Monocular |  | 수술전 | 안 | 84 | 0.61 | 0.32 | 72 | 0.8 | 0.48
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | Refractive Multifocal | Rezoom (Abbott Medical Optics, Inc.) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) |  | VA | UDVA | uncorrected | Distant(6m) | LogMAR | Monocular |  | 수술전 | 안 | 70 | 0.71 | 0.46 | 72 | 0.8 | 0.48
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | Apodized diffractive multifocal IOL | AcrySof Restor SN6AD3 (Alcon) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) |  | VA | CDVA | corrected | Distant(6m) | LogMAR | Monocular |  | 수술전 | 안 | 78 | 0.23 | 0.22 | 72 | 0.26 | 0.22
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | Diffractive bifocal IOL | Acri.Lisa 366D (Carl Zeiss Meditec AG) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) |  | VA | CDVA | corrected | Distant(6m) | LogMAR | Monocular |  | 수술전 | 안 | 84 | 0.14 | 0.15 | 72 | 0.26 | 0.22
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | Refractive Multifocal | Rezoom (Abbott Medical Optics, Inc.) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) |  | VA | CDVA | corrected | Distant(6m) | LogMAR | Monocular |  | 수술전 | 안 | 70 | 0.29 | 0.22 | 72 | 0.26 | 0.22
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | Apodized diffractive multifocal IOL | AcrySof Restor SN6AD3 (Alcon) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) |  | VA | UDVA | uncorrected | Distant(6m) | LogMAR | Monocular |  | 1개월 | 안 | 78 | 0.16 | 0.14 | 72 | 0.12 | 0.15
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | Diffractive bifocal IOL | Acri.Lisa 366D (Carl Zeiss Meditec AG) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) |  | VA | UDVA | uncorrected | Distant(6m) | LogMAR | Monocular |  | 1개월 | 안 | 84 | 0.13 | 0.13 | 72 | 0.12 | 0.15
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | Refractive Multifocal | Rezoom (Abbott Medical Optics, Inc.) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) |  | VA | UDVA | uncorrected | Distant(6m) | LogMAR | Monocular |  | 1개월 | 안 | 70 | 0.13 | 0.13 | 72 | 0.12 | 0.15
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | Apodized diffractive multifocal IOL | AcrySof Restor SN6AD3 (Alcon) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) |  | VA | CDVA | corrected | Distant(6m) | LogMAR | Monocular |  | 1개월 | 안 | 78 | 0.08 | 0.12 | 72 | 0.02 | 0.1
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | Diffractive bifocal IOL | Acri.Lisa 366D (Carl Zeiss Meditec AG) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) |  | VA | CDVA | corrected | Distant(6m) | LogMAR | Monocular |  | 1개월 | 안 | 84 | 0.05 | 0.1 | 72 | 0.02 | 0.1
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | Refractive Multifocal | Rezoom (Abbott Medical Optics, Inc.) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) |  | VA | CDVA | corrected | Distant(6m) | LogMAR | Monocular |  | 1개월 | 안 | 70 | 0.05 | 0.1 | 72 | 0.02 | 0.1
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | Apodized diffractive multifocal IOL | AcrySof Restor SN6AD3 (Alcon) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) |  | VA | UDVA | uncorrected | Distant(6m) | LogMAR | Monocular |  | 6개월 | 안 | 78 | 0.15 | 0.12 | 72 | 0.09 | 0.15
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | Diffractive bifocal IOL | Acri.Lisa 366D (Carl Zeiss Meditec AG) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) |  | VA | UDVA | uncorrected | Distant(6m) | LogMAR | Monocular |  | 6개월 | 안 | 84 | 0.12 | 0.11 | 72 | 0.09 | 0.15
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | Refractive Multifocal | Rezoom (Abbott Medical Optics, Inc.) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) |  | VA | UDVA | uncorrected | Distant(6m) | LogMAR | Monocular |  | 6개월 | 안 | 70 | 0.12 | 0.13 | 72 | 0.09 | 0.15
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | Apodized diffractive multifocal IOL | AcrySof Restor SN6AD3 (Alcon) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) |  | VA | CDVA | corrected | Distant(6m) | LogMAR | Monocular |  | 6개월 | 안 | 78 | 0.06 | 0.08 | 72 | 0.04 | 0.08
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | Diffractive bifocal IOL | Acri.Lisa 366D (Carl Zeiss Meditec AG) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) |  | VA | CDVA | corrected | Distant(6m) | LogMAR | Monocular |  | 6개월 | 안 | 84 | 0.06 | 0.1 | 72 | 0.04 | 0.08
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | Refractive Multifocal | Rezoom (Abbott Medical Optics, Inc.) | Monofocal IOL | Acri.Smart 48S (Carl Zeiss Meditec AG) |  | VA | CDVA | corrected | Distant(6m) | LogMAR | Monocular |  | 6개월 | 안 | 70 | 0.06 | 0.011 | 72 | 0.04 | 0.08
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 146 | 0.39 | 0.19
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 146 | 0.33 | 0.13
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular |  | 12개월 | 명 | 145 | 0.02 | 0.11
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Binocular |  | 6개월 | 명 | 145 | -0.04 | 0.11
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 12개월 | 명 | 145 | 0.01 | 0.13
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 12개월 | 명 | 145 | 0.01 | 0.1
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | DCIVA | distance corrected | intermediate(60cm) | LogMAR | Binocular |  | 12개월 | 명 | 145 | 0.04 | 0.12
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | DCIVA | distance corrected | intermediate(80cm) | LogMAR | Binocular |  | 12개월 | 명 | 145 | 0.08 | 0.14
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | UNVA | uncorrected | near | LogMAR | Binocular |  | 12개월 | 명 | 145 | 0.07 | 0.11
469 | Pedrotti (2020) | Prospective observational comparative clinical study | Presbyopia | Diffractive EDOF | Tecnis Symfony ZXR00 | Diffractive trifocal IOL | Acrysof IQ PanOptix TNFT00 |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular |  | 3개월 | 명 | 25 | 0 | 0.09 | 25 | -0.02 | 0.09
469 | Pedrotti (2020) | Prospective observational comparative clinical study | Presbyopia | Diffractive EDOF | Tecnis Symfony ZXR00 | Diffractive trifocal IOL | Acrysof IQ PanOptix TNFT00 |  | VA | BCDVA | best-corrected | Distant(4m) | LogMAR | Binocular |  | 3개월 | 명 | 25 | -0.09 | 0.05 | 25 | -0.08 | 0.06
469 | Pedrotti (2020) | Prospective observational comparative clinical study | Presbyopia | Diffractive EDOF | Tecnis Symfony ZXR00 | Diffractive trifocal IOL | Acrysof IQ PanOptix TNFT00 |  | VA | UDIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular |  | 3개월 | 명 | 25 | 0.1 | 0.03 | 25 | 0.12 | 0.04
469 | Pedrotti (2020) | Prospective observational comparative clinical study | Presbyopia | Diffractive EDOF | Tecnis Symfony ZXR00 | Diffractive trifocal IOL | Acrysof IQ PanOptix TNFT00 |  | VA | BDCIVA | best distance corrected | intermediate(60cm) | LogMAR | Binocular |  | 3개월 | 명 | 25 | 0.07 | 0.03 | 25 | 0.1 | 0.03
469 | Pedrotti (2020) | Prospective observational comparative clinical study | Presbyopia | Diffractive EDOF | Tecnis Symfony ZXR00 | Diffractive trifocal IOL | Acrysof IQ PanOptix TNFT00 |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | 3개월 | 명 | 25 | 0.26 | 0.08 | 25 | 0.14 | 0.05
469 | Pedrotti (2020) | Prospective observational comparative clinical study | Presbyopia | Diffractive EDOF | Tecnis Symfony ZXR00 | Diffractive trifocal IOL | Acrysof IQ PanOptix TNFT00 |  | VA | BDCNVA | best distance corrected | near(40cm) | LogMAR | Binocular |  | 3개월 | 명 | 25 | 0.26 | 0.05 | 25 | 0.14 | 0.04
469 | Pedrotti (2020) | Prospective observational comparative clinical study | Presbyopia | Diffractive EDOF | Tecnis Symfony ZXR00 | Diffractive trifocal IOL | Acrysof IQ PanOptix TNFT00 |  | CS |  |  |  | c/deg |  | means of mean MTF cut-off evaluation | 3개월 | 명 | 25 | 34.25 | 11.8 | 25 | 27.41 | 13.5
1409 | Altaie (2012) | Prospective, observational study | Cataract or CLE | Diffractive bifocal IOL | AcrySof Natural ReSTOR IOL (SN60D3) |  |  | Cataract group | VA | BCVA | best-corrected |  | LogMAR | Monocular |  | 수술전 | 명 | 199 | 0.3 | 0.22
1409 | Altaie (2012) | Prospective, observational study | Cataract or CLE | Diffractive bifocal IOL | AcrySof Natural ReSTOR IOL (SN60D3) |  |  | CLE | VA | BCVA | best-corrected |  | LogMAR | Monocular |  | 수술전 | 명 | 164 | 0 | NA
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | UCVA | uncorrected | Distant(6m) | LogMAR | Monocular |  | 6개월 | 안 | 21 | 0.59 | 0.21
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | BCVA | best-corrected | Distant(6m) | LogMAR | Monocular |  | 수술전 | 안 | 21 | 0.81 | 0.19
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | BCVA | best-corrected | Distant(6m) | LogMAR | Monocular |  | 6개월 | 안 | 21 | 0.87 | 0.14
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | UCVA | uncorrected | Distant(6m) | LogMAR | Monocular |  | 6개월 | 안 | 45 | 0.76 | 0.23
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | BCVA | best-corrected | Distant(6m) | LogMAR | Monocular |  | 수술전 | 안 | 45 | 0.89 | 0.14
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | BCVA | best-corrected | Distant(6m) | LogMAR | Monocular |  | 6개월 | 안 | 45 | 0.92 | 0.15
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | LogMAR | Monocular |  | 수술전 | 안 | 21 | 0.93 | 0.11
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | LogMAR | Monocular |  | 6개월 | 안 | 21 | 0.96 | 0.11
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | BCNVA | best-corrected | Near(33cm) | LogMAR | Monocular |  | 수술전 | 안 | 45 | 0.97 | 0.06
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | BCNVA | best-corrected | Near(33cm) | LogMAR | Monocular |  | 6개월 | 안 | 45 | 0.97 | 0.14
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | UCVA | uncorrected | Distant(6m) | LogMAR | Monocular |  | 6개월 | 안 | 14 | 0.9 | 0.12
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | BCVA | best-corrected | Distant(6m) | LogMAR | Monocular |  | 수술전 | 안 | 14 | 0.82 | 0.16
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | BCVA | best-corrected | Distant(6m) | LogMAR | Monocular |  | 6개월 | 안 | 14 | 0.99 | 0.03
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | UCVA | uncorrected | Distant(6m) | LogMAR | Monocular |  | 6개월 | 안 | 20 | 0.88 | 0.18
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | BCVA | best-corrected | Distant(6m) | LogMAR | Monocular |  | 수술전 | 안 | 20 | 0.89 | 0.11
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | BCVA | best-corrected | Distant(6m) | LogMAR | Monocular |  | 6개월 | 안 | 20 | 0.96 | 0.09
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | LogMAR | Monocular |  | 수술전 | 안 | 14 | 0.98 | 0.04
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | LogMAR | Monocular |  | 6개월 | 안 | 14 | 1 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | UCNVA | uncorrected | Near(33cm) | LogMAR | Monocular |  | 6개월 | 안 | 20 | 0.98 | 0.07
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | UCVA | uncorrected | Distant(6m) | LogMAR | Monocular |  | 6개월 | 안 | 52 | 0.79 | 0.16
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | BCVA | best spectacle corrected | Distant(6m) | LogMAR | Monocular |  | 수술전 | 안 | 52 | 0.8 | 0.18
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | BCVA | best spectacle corrected | Distant(6m) | LogMAR | Monocular |  | 6개월 | 안 | 52 | 0.9 | 0.12
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | hyperopic | VA | UCVA | uncorrected | Distant(6m) | LogMAR | Monocular |  | 6개월 | 안 | 78 | 0.87 | 0.15
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | hyperopic | VA | BCVA | best spectacle corrected | Distant(6m) | LogMAR | Monocular |  | 수술전 | 안 | 78 | 0.94 | 0.11
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | hyperopic | VA | BCVA | best spectacle corrected | Distant(6m) | LogMAR | Monocular |  | 6개월 | 안 | 78 | 0.97 | 0.06
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | LogMAR | Monocular |  | 수술전 | 안 | 52 | 1 | 0
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | LogMAR | Monocular |  | 6개월 | 안 | 52 | 0.99 | 0.03
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | UCNVA | uncorrected | Near(33cm) | LogMAR | Monocular |  | 6개월 | 안 | 78 | 1 | 0
3093 | Tan (2019) | Prospective study | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Technis Symfony IOL |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 수술전 |  |  | 0.25 | NA
3093 | Tan (2019) | Prospective study | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Technis Symfony IOL |  |  | with glare | CS | 6CPD |  |  | log | [5 empty] | 1.43 | NA
3093 | Tan (2019) | Prospective study | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Technis Symfony IOL |  |  | with glare | CS | 12CPD |  |  | log | [5 empty] | 1.18 | NA
3093 | Tan (2019) | Prospective study | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Technis Symfony IOL |  |  | without glare | CS | 6CPD |  |  | log | [5 empty] | 1.6 | NA
3093 | Tan (2019) | Prospective study | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Technis Symfony IOL |  |  | without glare | CS | 12CPD |  |  | log | [5 empty] | 1.21 | NA
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UDVA | uncorrected | Distant | Decimal | Monocular |  | 수술전 | 안 | 82 | 0.35
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UDVA | uncorrected | Distant | Decimal | Monocular |  | 1개월 | 안 | 82 | 0.89
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UDVA | uncorrected | Distant | Decimal | Monocular |  | 3개월 | 안 | 82 | 0.9 | [8 empty] | <0.001 | p값으로 SD 산출
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | CDVA | corrected | Distant | Decimal | Monocular |  | 수술전 | 안 | 82 | 0.74
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | CDVA | corrected | Distant | Decimal | Monocular |  | 3개월 | 안 | 82 | 0.95
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | CNVA | corrected | near(40cm) | Decimal | Monocular |  | 수술전 | 안 | 82 | 0.78
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | Decimal | Monocular |  | 1개월 | 안 | 82 | 0.84
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | Decimal | Monocular |  | 3개월 | 안 | 82 | 0.87 | [8 empty] | 0.123 | p값으로 SD 산출
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UNVA | uncorrected | Near(40cm) | Decimal | Monocular |  | 1개월 | 안 | 82 | 0.88 | NA
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UNVA | uncorrected | Near(40cm) | Decimal | Monocular |  | 3개월 | 안 | 82 | 0.91 | NA | [7 empty] | 0.184 | p값으로 SD 산출
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series | Cataract surgery | Diffractive trifocal IOL | AT LISA tri 839MP | monofocal IOL | Acri.Smart 48S |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 30 | 0.63 | 0.58 | 30 | 1.12 | 0.7
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series | Cataract surgery | Diffractive trifocal IOL | AT LISA tri 839MP | monofocal IOL | Acri.Smart 48S |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 30 | 0.08 | 0.24 | 30 | 0.22 | 0.24
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series | Cataract surgery | Diffractive trifocal IOL | FineVision | monofocal IOL | Acri.Smart 48S |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 30 | 0.72 | 0.34 | 30 | 1.12 | 0.7
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series | Cataract surgery | Diffractive trifocal IOL | FineVision | monofocal IOL | Acri.Smart 48S |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 30 | 0.27 | 0.23 | 30 | 0.22 | 0.24
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series | Cataract surgery | Diffractive bifocal IOL | AcrySof ReSTOR SN6AD1 | monofocal IOL | Acri.Smart 48S |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 30 | 0.58 | 0.62 | 30 | 1.12 | 0.7
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series | Cataract surgery | Diffractive bifocal IOL | AcrySof ReSTOR SN6AD1 | monofocal IOL | Acri.Smart 48S |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 30 | 0.07 | 0.15 | 30 | 0.22 | 0.24
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series | Cataract surgery | Diffractive trifocal IOL | AT LISA tri 839MP | monofocal IOL | Acri.Smart 48S |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 30 | 0 | 0.07 | 30 | 0.14 | 0.16
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series | Cataract surgery | Diffractive trifocal IOL | AT LISA tri 839MP | monofocal IOL | Acri.Smart 48S |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 30 | 0.01 | 0.05 | 30 | 0.03 | 0.05
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series | Cataract surgery | Diffractive trifocal IOL | AT LISA tri 839MP | monofocal IOL | Acri.Smart 48S |  | VA | DCNVA | distance corrected | near | LogMAR | Monocular |  | 3개월 | 안 | 30 | 0.15 | 0.11 | 30 | 0.49 | 0.12
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series | Cataract surgery | Diffractive trifocal IOL | AT LISA tri 839MP | monofocal IOL | Acri.Smart 48S |  | VA | UNVA | uncorrected | near | LogMAR | Monocular |  | 3개월 | 안 | 30 | 0.17 | 0.1 | 30 | 0.51 | 0.17
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series | Cataract surgery | Diffractive trifocal IOL | FineVision | monofocal IOL | Acri.Smart 48S |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 30 | 0.1 | 0.09 | 30 | 0.14 | 0.16
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series | Cataract surgery | Diffractive trifocal IOL | FineVision | monofocal IOL | Acri.Smart 48S |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 30 | 0.05 | 0.07 | 30 | 0.03 | 0.05
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series | Cataract surgery | Diffractive trifocal IOL | FineVision | monofocal IOL | Acri.Smart 48S |  | VA | DCNVA | distance corrected | near | LogMAR | Monocular |  | 3개월 | 안 | 30 | 0.16 | 0.1 | 30 | 0.49 | 0.12
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series | Cataract surgery | Diffractive trifocal IOL | FineVision | monofocal IOL | Acri.Smart 48S |  | VA | UNVA | uncorrected | near | LogMAR | Monocular |  | 3개월 | 안 | 30 | 0.21 | 0.1 | 30 | 0.51 | 0.17
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series | Cataract surgery | Diffractive bifocal IOL | AcrySof ReSTOR SN6AD1 | monofocal IOL | Acri.Smart 48S |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 30 | 0.19 | 0.14 | 30 | 0.14 | 0.16
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series | Cataract surgery | Diffractive bifocal IOL | AcrySof ReSTOR SN6AD1 | monofocal IOL | Acri.Smart 48S |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 30 | 0.06 | 0.05 | 30 | 0.03 | 0.05
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series | Cataract surgery | Diffractive bifocal IOL | AcrySof ReSTOR SN6AD1 | monofocal IOL | Acri.Smart 48S |  | VA | DCNVA | distance corrected | near | LogMAR | Monocular |  | 3개월 | 안 | 30 | 0.15 | 0.1 | 30 | 0.49 | 0.12
3522 | Plaza-Puche (2016) | Prospective interventional comparative nonrandomized consecutive case series | Cataract surgery | Diffractive bifocal IOL | AcrySof ReSTOR SN6AD1 | monofocal IOL | Acri.Smart 48S |  | VA | UNVA | uncorrected | near | LogMAR | Monocular |  | 3개월 | 안 | 30 | 0.19 | 0.11 | 30 | 0.51 | 0.17
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 3개월 | 안 | 89 | 0.03 | 0.11
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular |  | 3개월 | 안 | 58 | -0.02 | 0.1
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 89 | 0.17 | 0.14
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | 3개월 | 안 | 58 | 0.13 | 0.12
4240 | Van Der Linden (2012) | Comparative case series | Presbyopia | Refractive bifocal IOL | Lentis Mplus LS-312 | Diffractive bifocal IOL | ReSTOR SN6AD1 |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 3개월 | 안 | 90 | 0.04 | 0.15 | 143 | 0.06 | 0.25
4240 | Van Der Linden (2012) | Comparative case series | Presbyopia | Refractive bifocal IOL | Lentis Mplus LS-312 | Diffractive bifocal IOL | ReSTOR SN6AD1 |  | VA | UNVA | uncorrected | Near(30cm) | LogMAR | Monocular |  | 3개월 | 안 | 90 | 0.05 | 0.14 | 143 | 0.15 | 0.08
4240 | Van Der Linden (2012) | Comparative case series | Presbyopia | Refractive bifocal IOL | Lentis Mplus LS-312 | Diffractive bifocal IOL | ReSTOR SN6AD1 |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 90 | 0.05 | 0.14 | 143 | 0.16 | 0.21
4240 | Van Der Linden (2012) | Comparative case series | Presbyopia | Refractive bifocal IOL | Lentis Mplus LS-312 | Diffractive bifocal IOL | ReSTOR SN6AD1 |  | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 90 | 0.57 | 0.58 | 143 | 0.63 | 0.61
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 50 | 0.2 | 0.117
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 50 | 0.08 | 0.142
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Monocular |  | 3개월 | 안 | 50 | 0.12 | 0.107
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 50 | 0.1 | 0.111
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCVA | distance corrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 50 | -0.05 | 0.095
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance corrected | intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 50 | 0.1 | 0.127
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance corrected | intermediate(60cm) | LogMAR | Monocular |  | 3개월 | 안 | 50 | 0.14 | 0.109
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCNVA | distance corrected | near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 50 | 0.05 | 0.129
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular |  | 3개월 | 명 | 25 | -0.06 | 0.1
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Binocular |  | 3개월 | 명 | 25 | 0.05 | 0.1
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular |  | 3개월 | 명 | 25 | 0.11 | 0.086
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | 3개월 | 명 | 25 | 0.02 | 0.097
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCVA | distance corrected | Distant(4m) | LogMAR | Binocular |  | 3개월 | 명 | 25 | -0.11 | 0.086
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance corrected | intermediate(80cm) | LogMAR | Binocular |  | 3개월 | 명 | 25 | 0.02 | 0.099
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance corrected | intermediate(60cm) | LogMAR | Binocular |  | 3개월 | 명 | 25 | 0.11 | 0.093
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCNVA | distance corrected | near(40cm) | LogMAR | Binocular |  | 3개월 | 명 | 25 | 0 | 0.093
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  | photopic | CS |  |  |  | logCSWebar |  |  |  | 명 | 25 | 1.4 | 0.4
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  | mesopic | CS |  |  |  | logCSWebar |  |  |  | 명 | 25 | 0.88 | 0.14
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  | mesopic with glare lighting condition | CS |  |  |  | logCSWebar |  |  |  | 명 | 25 | 0.79 | 0.13
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 20 | 0.51 | 0.3
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 20 | 0.13 | 0.22
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 20 | 0.13 | 0.09
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 20 | 0.02 | 0.06
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 20 | 0.22 | 0.11
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | DCNVA | distance corrected | near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 20 | 0.18 | 0.1
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | CNVA | corrected | near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 20 | 0.08 | 0.05
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 20 | 0.22 | 0.1
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 12개월 | 안 | 20 | 0.08 | 0.08
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 12개월 | 안 | 20 | -0.01 | 0.05
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 12개월 | 안 | 20 | 0.22 | 0.12
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | DCNVA | distance corrected | near(40cm) | LogMAR | Monocular |  | 12개월 | 안 | 20 | 0.16 | 0.13
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular |  | 수술전 | 명 | 10 | 0.34 | 0.26
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Binocular |  | 수술전 | 명 | 10 | 0.01 | 0.15
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular |  | 3개월 | 명 | 10 | 0.04 | 0.04
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Binocular |  | 3개월 | 명 | 10 | -0.03 | 0.04
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | 3개월 | 명 | 10 | 0.15 | 0.13
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | DCNVA | distance corrected | near(40cm) | LogMAR | Binocular |  | 3개월 | 명 | 10 | 0.15 | 0.14
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | CNVA | corrected | near(40cm) | LogMAR | Binocular |  | 3개월 | 명 | 10 | 0.08 | 0.05
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Binocular |  | 3개월 | 명 | 10 | 0.16 | 0.11
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular |  | 12개월 | 명 | 10 | -0.01 | 0.05
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Binocular |  | 12개월 | 명 | 10 | -0.11 | 0.13
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | 12개월 | 명 | 10 | 0.13 | 0.12
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | VA | DCNVA | distance corrected | near(40cm) | LogMAR | Binocular |  | 12개월 | 명 | 10 | 0.08 | 0.11
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  | photopic | CS | 3CPD |  |  | log |  |  | 12개월 | 명 | 10 | 1.68 | NA
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  | photopic | CS | 6CPD |  |  | log |  |  | 12개월 | 명 | 10 | 1.74 | NA
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  | photopic | CS | 12CPD |  |  | log |  |  | 12개월 | 명 | 10 | 1.53 | NA
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  | photopic | CS | 18CPD |  |  | log |  |  | 12개월 | 명 | 10 | 1.05 | NA
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  | photopic with glare | CS | 3CPD |  |  | log |  |  | 12개월 | 명 | 10 | 1.56 | NA
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  | photopic with glare | CS | 6CPD |  |  | log |  |  | 12개월 | 명 | 10 | 1.73 | NA
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  | photopic with glare | CS | 12CPD |  |  | log |  |  | 12개월 | 명 | 10 | 1.47 | NA
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  | photopic with glare | CS | 18CPD |  |  | log |  |  | 12개월 | 명 | 10 | 0.95 | NA
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  | mesopic | CS | 3CPD |  |  | log |  |  | 12개월 | 명 | 10 | 1.59 | NA
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  | mesopic | CS | 6CPD |  |  | log |  |  | 12개월 | 명 | 10 | 1.67 | NA
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  | mesopic | CS | 12CPD |  |  | log |  |  | 12개월 | 명 | 10 | 1.28 | NA
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  | mesopic | CS | 18CPD |  |  | log |  |  | 12개월 | 명 | 10 | 0.78 | NA
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  | mesopic with glare | CS | 3CPD |  |  | log |  |  | 12개월 | 명 | 10 | 1.5 | NA
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  | mesopic with glare | CS | 6CPD |  |  | log |  |  | 12개월 | 명 | 10 | 1.53 | NA
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  | mesopic with glare | CS | 12CPD |  |  | log |  |  | 12개월 | 명 | 10 | 1.06 | NA
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  | mesopic with glare | CS | 18CPD |  |  | log |  |  | 12개월 | 명 | 10 | 0.65 | NA
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 4개월 | 안 | 94 | 0.12 | 0.1
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular |  | 4개월 | 명 | 40 | 0.05 | 0.07
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 4개월 | 안 | 94 | 0.34 | 0.09
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | 4개월 | 명 | 40 | 0.29 | 0.06
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | CLE | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 45 | 0.04 | (0.03-0.06) | [8 empty] | 95%CI로 SD계산
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | CLE | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 4개월 | 안 | 45 | 0.11 | 0.09
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | CLE | VA | UDVA | uncorrected | Distant | LogMAR | Binocular |  | 4개월 | 명 | 22 | 0.05 | 0.07
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | CLE | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 4개월 | 안 | 45 | 0.33 | 0.07
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | CLE | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | 4개월 | 명 | 22 | 0.27 | 0.06
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | Cataract group | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 49 | 0.04 | (0.03-0.06) | [8 empty] | 95%CI로 SD계산
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | Cataract group | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 4개월 | 안 | 49 | 0.14 | 0.12
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | Cataract group | VA | UDVA | uncorrected | Distant | LogMAR | Binocular |  | 4개월 | 명 | 18 | 0.05 | 0.08
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | Cataract group | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 4개월 | 안 | 49 | 0.35 | 0.1
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | Cataract group | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | 4개월 | 명 | 18 | 0.3 | 0.07
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | 0-0.5D | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 4개월 | 안 | 60 | 0.11 | 0.09
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | 0-0.5D | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 4개월 | 안 | 60 | 0.33 | 0.09
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | 0.5-0.99D | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 4개월 | 안 | 20 | 0.1 | 0.08
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | 0.5-0.99D | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 4개월 | 안 | 20 | 0.35 | 0.1
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | >1D(high preopoerative corneal astigmatism) | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 4개월 | 안 | 14 | 0.21 | 0.14
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | >1D(high preopoerative corneal astigmatism) | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 4개월 | 안 | 14 | 0.39 | 0.07
536 | Escandon-Garcia (2018) | Prospective, nonrandomized, examiner-masked case series | Cataract surgery | Diffractive Trifocal IOL | AcrySof® IQ PanOptix (TFNT00) |  |  |  | VA | UCVA | uncorrected | Distant | LogMAR | Binocular |  |  | 명 | 7 | 0.07 | 0.1
536 | Escandon-Garcia (2018) | Prospective, nonrandomized, examiner-masked case series | Cataract surgery | Diffractive Trifocal IOL | AcrySof® IQ PanOptix (TFNT00) |  |  |  | VA | BSCVA | best distance spectacle-corrected | Distant | LogMAR | Binocular |  |  | 명 | 7 | -0.07 | 0.19
536 | Escandon-Garcia (2018) | Prospective, nonrandomized, examiner-masked case series | Cataract surgery | Diffractive EDOF | TECNIS® Symfony model ZXR00 (Abbott Medical Optics, Santa Ana, USA) |  |  |  | VA | UCVA | uncorrected | Distant | LogMAR | Binocular |  |  | 명 | 15 | 0.08 | 0.1
536 | Escandon-Garcia (2018) | Prospective, nonrandomized, examiner-masked case series | Cataract surgery | Diffractive EDOF | TECNIS® Symfony model ZXR00 (Abbott Medical Optics, Santa Ana, USA) |  |  |  | VA | BSCVA | best distance spectacle-corrected | Distant | LogMAR | Binocular |  |  | 명 | 15 | -0.1 | 0.19
536 | Escandon-Garcia (2018) | Prospective, nonrandomized, examiner-masked case series | Cataract surgery | Diffractive Trifocal IOL | FineVision Pod F (PhysIOL, Liège, BE) |  |  |  | VA | UCVA | uncorrected | Distant | LogMAR | Binocular |  |  | 명 | 23 | 0.08 | 0.09
536 | Escandon-Garcia (2018) | Prospective, nonrandomized, examiner-masked case series | Cataract surgery | Diffractive Trifocal IOL | FineVision Pod F (PhysIOL, Liège, BE) |  |  |  | VA | BSCVA | best distance spectacle-corrected | Distant | LogMAR | Binocular |  |  | 명 | 23 | -0.24 | 0.14
536 | Escandon-Garcia (2018) | Prospective, nonrandomized, examiner-masked case series | Cataract surgery |  |  |  |  | whole sample | VA | UCVA | uncorrected | Distant | LogMAR | Binocular |  |  | 명 | 45 | 0.08 | 0.12
536 | Escandon-Garcia (2018) | Prospective, nonrandomized, examiner-masked case series | Cataract surgery |  |  |  |  | whole sample | VA | BSCVA | best distance spectacle-corrected | Distant | LogMAR | Binocular |  |  | 명 | 45 | -0.16 | 0.27
621 | Friedrich (2012) | Prospective consecutive study | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UCVA | uncorrected | Distant | LogMAR | Monocular |  | 6개월 | 안 | 31 | 0.08 | 0.11
621 | Friedrich (2012) | Prospective consecutive study | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UCVA | uncorrected | Distant | LogMAR | Binocular |  | 6개월 | 명 | 19 | 0.01 | 0.08
621 | Friedrich (2012) | Prospective consecutive study | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | BCVA | best-corrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 31 | -0.41 | 0.32
621 | Friedrich (2012) | Prospective consecutive study | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | BCVA | best-corrected | Distant | LogMAR | Monocular |  | 6개월 | 안 | 31 | 0.03 | 0.05
621 | Friedrich (2012) | Prospective consecutive study | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UCVA | uncorrected | near | LogMAR | Monocular |  | 6개월 | 안 | 31 | 0.03 | 0.09
621 | Friedrich (2012) | Prospective consecutive study | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UCVA | uncorrected | near | LogMAR | Binocular |  | 6개월 | 명 | 19 | 0 | 0.08
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | BCDVA | best-corrected | Distant(4m) |  | Binocular |  | 1개월 | 명 | 44 | -0.015 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | UCDVA | uncorrected | Distant(4m) |  | Binocular |  | 1개월 | 명 | 44 | 0.049 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | UCIVA | uncorrected | Intermediate(60cm) |  | Binocular |  | 1개월 | 명 | 44 | 0.015 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | UCNVA | uncorrected | Near(40cm) |  | Binocular |  | 1개월 | 명 | 44 | 0.047 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | BCDVA | best-corrected | Distant(4m) |  | Binocular |  | 3개월 | 명 | 44 | -0.049 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | UCDVA | uncorrected | Distant(4m) |  | Binocular |  | 3개월 | 명 | 44 | 0.026 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | UCIVA | uncorrected | Intermediate(60cm) |  | Binocular |  | 3개월 | 명 | 44 | -0.025 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | UCNVA | uncorrected | Near(40cm) |  | Binocular |  | 3개월 | 명 | 44 | 0.029 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | VA | UCDVA | uncorrected | Distant(4m) |  | Monocular |  | 1개월 | 안 | 45 | 0.089 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | VA | UCIVA | uncorrected | Intermediate(60cm) |  | Monocular |  | 1개월 | 안 | 45 | 0.08 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | VA | UCNVA | uncorrected | Near(40cm) |  | Monocular |  | 1개월 | 안 | 45 | 0.116 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | VA | UCDVA | uncorrected | Distant(4m) |  | Monocular |  | 1개월 | 안 | 44 | 0.097 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | VA | UCIVA | uncorrected | Intermediate(60cm) |  | Monocular |  | 1개월 | 안 | 44 | 0.08 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | VA | UCNVA | uncorrected | Near(40cm) |  | Monocular |  | 1개월 | 안 | 44 | 0.104 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | VA | UCDVA | uncorrected | Distant(4m) |  | Monocular |  | 3개월 | 안 | 45 | 0.073 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | VA | UCIVA | uncorrected | Intermediate(60cm) |  | Monocular |  | 3개월 | 안 | 45 | 0.054 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | VA | UCNVA | uncorrected | Near(40cm) |  | Monocular |  | 3개월 | 안 | 45 | 0.089 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | VA | UCDVA | uncorrected | Distant(4m) |  | Monocular |  | 3개월 | 안 | 44 | 0.081 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | VA | UCIVA | uncorrected | Intermediate(60cm) |  | Monocular |  | 3개월 | 안 | 44 | 0.035 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | VA | UCNVA | uncorrected | Near(40cm) |  | Monocular |  | 3개월 | 안 | 44 | 0.091 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | photopic without glare | CS | 3CPD |  |  |  | Binocular |  | 3개월 | 명 | 44 | 1.7 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | photopic without glare | CS | 6CPD |  |  |  | Binocular |  | 3개월 | 명 | 44 | 1.9 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | photopic without glare | CS | 12CPD |  |  |  | Binocular |  | 3개월 | 명 | 44 | 1.6 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | photopic without glare | CS | 18CPD |  |  |  | Binocular |  | 3개월 | 명 | 44 | 1.1 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | photopic with glare | CS | 3CPD |  |  |  | Binocular |  | 3개월 | 명 | 44 | 1.7 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | photopic with glare | CS | 6CPD |  |  |  | Binocular |  | 3개월 | 명 | 44 | 1.9 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | photopic with glare | CS | 12CPD |  |  |  | Binocular |  | 3개월 | 명 | 44 | 1.5 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | photopic with glare | CS | 18CPD |  |  |  | Binocular |  | 3개월 | 명 | 44 | 1.1 | NA
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | VA | UDVA | uncorrected | Distant(5m) | LogMAR | Binocular |  | 수술전 | 명 | 34 | 0.63 | 0.34 | 36 | 0.66 | 0.4
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | VA | UDVA | uncorrected | Distant(5m) | LogMAR | Binocular |  | 수술전 | 명 | 34 | 0.64 | 0.38 | 36 | 0.66 | 0.4
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | VA | UDVA | uncorrected | Distant(5m) | LogMAR | Monocular |  | 1개월 | 안 | 47 | 0.12 | 0.14 | 47 | 0.17 | 0.2
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | VA | UDVA | uncorrected | Distant(5m) | LogMAR | Monocular |  | 1개월 | 안 | 47 | 0.18 | 0.16 | 47 | 0.17 | 0.2
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Monocular |  | 1개월 | 안 | 47 | 0.18 | 0.16 | 47 | 0.117 | 0.18
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Monocular |  | 1개월 | 안 | 47 | 0.18 | 0.2 | 47 | 0.117 | 0.18
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 1개월 | 안 | 47 | 0.34 | 0.19 | 47 | 0.41 | 0.2
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 1개월 | 안 | 47 | 0.3 | 0.21 | 47 | 0.41 | 0.2
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | VA | UDVA | uncorrected | Distant(5m) | LogMAR | Monocular |  | 3개월 | 안 | 47 | 0.1 | 0.13 | 47 | 0.15 | 0.17
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | VA | UDVA | uncorrected | Distant(5m) | LogMAR | Monocular |  | 3개월 | 안 | 47 | 0.12 | 0.13 | 47 | 0.15 | 0.17
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 47 | 0.14 | 0.12 | 47 | 0.2 | 0.18
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 47 | 0.16 | 0.17 | 47 | 0.2 | 0.18
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 47 | 0.34 | 0.19 | 47 | 0.38 | 0.16
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 47 | 0.32 | 0.18 | 47 | 0.38 | 0.16
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | VA | CDVA | corrected | Distant(5m) | LogMAR | Monocular |  | 3개월 | 안 | 47 | -0.01 | 0.07 | 47 | -0.01 | 0.08
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | VA | CDVA | corrected | Distant(5m) | LogMAR | Monocular |  | 3개월 | 안 | 47 | 0 | 0.05 | 47 | -0.01 | 0.08
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | VA | CIVA | corrected | intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 47 | 0.02 | 0.1 | 47 | 0.04 | 0.11
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | VA | CIVA | corrected | intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 47 | 0.03 | 0.07 | 47 | 0.04 | 0.11
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | VA | CNVA | corrected | near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 47 | 0.09 | 0.2 | 47 | 0.06 | 0.11
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | VA | CNVA | corrected | near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 47 | 0.08 | 0.11 | 47 | 0.06 | 0.11
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | VA | DCIVA | distance corrected | intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 47 | 0.11 | 0.13 | 47 | 0.21 | 0.14
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | VA | DCIVA | distance corrected | intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 47 | 0.14 | 0.11 | 47 | 0.21 | 0.14
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | VA | DCNVA | distance corrected | near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 47 | 0.37 | 0.17 | 47 | 0.44 | 0.16
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | VA | DCNVA | distance corrected | near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 47 | 0.36 | 0.15 | 47 | 0.44 | 0.16
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | day glares |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0.36 | 1.42 | 47 | 0.15 | 0.59
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | day glares |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0.38 | 1.19 | 47 | 0.15 | 0.59
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | night glares |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0.43 | 1.16 | 47 | 0.19 | 0.97
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | night glares |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0.57 | 1.54 | 47 | 0.19 | 0.97
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | halos |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 1.81 | 2.47 | 47 | 0.21 | 0.83
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | halos |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0.64 | 1.52 | 47 | 0.21 | 0.83
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | starbursts |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 1.3 | 2.07 | 47 | 0.43 | 1.3
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | starbursts |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0.83 | 1.61 | 47 | 0.43 | 1.3
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | hazy visions |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0 | 0 | 47 | 0.06 | 0.44
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | hazy visions |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0.32 | 0.94 | 47 | 0.06 | 0.44
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | blurred vision |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0.66 | 1.32 | 47 | 0.45 | 1.08
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | blurred vision |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0.6 | 1.17 | 47 | 0.45 | 1.08
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | distortion |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0 | 0 | 47 | 0 | 0
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | distortion |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0.9 | 0.41 | 47 | 0 | 0
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | double vision |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0.11 | 0.52 | 47 | 0.13 | 0.88
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | double vision |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0.3 | 0.86 | 47 | 0.13 | 0.88
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | fluctuation in vision |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0.38 | 1.11 | 47 | 0.13 | 0.61
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | fluctuation in vision |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0.13 | 0.61 | 47 | 0.13 | 0.61
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | focusing difficulty |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0.36 | 0.13 | 47 | 0.06 | 0.06
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | focusing difficulty |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0.19 | 0.14 | 47 | 0.06 | 0.06
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Diffractive EDOF toric IOL | Tecnis Symfony ZXR00 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | depth perception difficulty |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0.26 | 0.77 | 47 | 0.02 | 0.15
666 | Song (2020) | Single-center prospective, nonrandomized, participantand examiner-blinded cohort study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF15 | Refractive monofocal IOL | Lentis L-313 |  | Satisfaction | depth perception difficulty |  |  | 0-3점(0점 best) | Monocular |  | 3개월 | 안 | 47 | 0.09 | 0.58 | 47 | 0.02 | 0.15
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractivce Trifocal IOL | AcrySof PanOptix |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 54 | 0.65 | 0.4
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractivce Trifocal IOL | AcrySof PanOptix |  |  |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 54 | 0.21 | 0.26
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 54 | 0.69 | 0.5
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 54 | 0.18 | 0.21
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR00 |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 52 | 0.47 | 0.36
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR00 |  |  |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 52 | 0.33 | 0.23
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF30 |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 50 | 0.58 | 0.35
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF30 |  |  |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 50 | 0.22 | 0.16
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 54 | 0.04 | 0.113
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 54 | -0.05 | 0.072
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 54 | 0.06 | 0.107
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  |  | VA | DCNVA | distance corrected | near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 54 | 0.05 | 0.094
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Monocular |  | 3개월 | 안 | 54 | 0.08 | 0.12
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  |  | VA | DCIVA | distance corrected | intermediate(60cm) | LogMAR | Monocular |  | 3개월 | 안 | 54 | 0.07 | 0.119
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 54 | 0.15 | 0.136
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  |  | VA | DCIVA | distance corrected | intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 54 | 0.15 | 0.144
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 54 | 0.01 | 0.103
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 54 | -0.04 | 0.083
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 54 | 0.09 | 0.116
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | DCNVA | distance corrected | near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 54 | 0.06 | 0.131
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 54 | 0.07 | 0.1152
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | DCIVA | distance corrected | intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 54 | 0.04 | 0.109
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR00 |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 52 | 0.01 | 0.09
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR01 |  |  |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 52 | -0.05 | 0.016
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR02 |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 52 | -0.3 | 0.146
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR03 |  |  |  | VA | DCNVA | distance corrected | near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 52 | 0.3 | 0.144
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR04 |  |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Monocular |  | 3개월 | 안 | 52 | 0.08 | 0.114
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR05 |  |  |  | VA | DCIVA | distance corrected | intermediate(60cm) | LogMAR | Monocular |  | 3개월 | 안 | 52 | 0.08 | 0.15
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR06 |  |  |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 52 | 0.04 | 0.087
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR07 |  |  |  | VA | DCIVA | distance corrected | intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 52 | 0.03 | 0.083
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF30 |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 50 | 0.05 | 0.122
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF31 |  |  |  | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular |  | 3개월 | 안 | 50 | -0.03 | 0.9
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF32 |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 50 | 0.16 | 0.14
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF33 |  |  |  | VA | DCNVA | distance corrected | near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 50 | 0.12 | 0.13
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF36 |  |  |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 50 | 0.18 | 0.165
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF37 |  |  |  | VA | DCIVA | distance corrected | intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 50 | 0.16 | 0.167
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 26 | 0.13 | 0.04
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 6개월 | 안 | 26 | 0.18 | 0.32
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Monocular |  | 수술전 | 안 | 26 | 0.46 | 0.12
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Monocular |  | 6개월 | 안 | 26 | 0.17 | 0.21
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | VA | UNVA | uncorrected | Near(30cm) | LogMAR | Monocular |  | 수술전 | 안 | 26 | 0.66 | 0.17
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | VA | UNVA | uncorrected | Near(30cm) | LogMAR | Monocular |  | 6개월 | 안 | 26 | 0.02 | 0.1
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | Car driving, day |  | Distant | 1-6점(1점 best) |  |  | 6개월 | 명 | 26 | 1.3 | 0.57
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | Car driving, night |  | Distant | 1-6점(1점 best) |  |  | 6개월 | 명 | 26 | 2.5 | 1.76
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | TV |  | Distant | 1-6점(1점 best) |  |  | 6개월 | 명 | 26 | 1.55 | 0.99
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | Therater, concert |  | Distant | 1-6점(1점 best) |  |  | 6개월 | 명 | 26 | 1.5 | 1
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | At home |  | Distant | 1-6점(1점 best) |  |  | 6개월 | 명 | 26 | 1.6 | 1.9
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | cooking |  | intermediate(60cm) | 1-6점(1점 best) |  |  | 6개월 | 명 | 26 | 1.85 | 1.18
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | domestic work |  | intermediate(60cm) | 1-6점(1점 best) |  |  | 6개월 | 명 | 26 | 1.85 | 1.18
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | computer |  | intermediate(60cm) | 1-6점(1점 best) |  |  | 6개월 | 명 | 26 | 2.25 | 1.61
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | using tablet |  | near(30cm) | 1-6점(1점 best) |  |  | 6개월 | 명 | 26 | 2.35 | 1.59
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | using cell-phone |  | near(30cm) | 1-6점(1점 best) |  |  | 6개월 | 명 | 26 | 2.25 | 1.61
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | reading a book |  | near(30cm) | 1-6점(1점 best) |  |  | 6개월 | 명 | 26 | 2.65 | 1.89
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | reading newspaper |  | near(30cm) | 1-6점(1점 best) |  |  | 6개월 | 명 | 26 | 2.7 | 1.98
732 | Vounotrypidis (2017) | Prospective longitudinal case series | Cataract surgery | Refractive EDOF IOL | Lentis comfort LS-313 MF15 |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 44 | 0.64 | 0.4
732 | Vounotrypidis (2017) | Prospective longitudinal case series | Cataract surgery | Refractive EDOF IOL | Lentis comfort LS-313 MF15 |  |  |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Monocular |  | 수술전 | 안 | 44 | 0.67 | 0.24
732 | Vounotrypidis (2017) | Prospective longitudinal case series | Cataract surgery | Refractive EDOF IOL | Lentis comfort LS-313 MF15 |  |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Monocular |  | 수술전 | 안 | 44 | 0.78 | 0.42
732 | Vounotrypidis (2017) | Prospective longitudinal case series | Cataract surgery | Refractive EDOF IOL | Lentis comfort LS-313 MF15 |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 44 | 0.31 | 0.15
732 | Vounotrypidis (2017) | Prospective longitudinal case series | Cataract surgery | Refractive EDOF IOL | Lentis comfort LS-313 MF15 |  |  |  | VA | CIVA | corrected | intermediate(80cm) | LogMAR | Monocular |  | 수술전 | 안 | 44 | 0.48 | 0.26
732 | Vounotrypidis (2017) | Prospective longitudinal case series | Cataract surgery | Refractive EDOF IOL | Lentis comfort LS-313 MF15 |  |  |  | VA | CNVA | corrected | Near(33cm) | LogMAR | Monocular |  | 수술전 | 안 | 44 | 0.62 | 0.28
732 | Vounotrypidis (2017) | Prospective longitudinal case series | Cataract surgery | Refractive EDOF IOL | Lentis comfort LS-313 MF15 |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 3개월 | 안 | 44 | 0.07 | 0.1
732 | Vounotrypidis (2017) | Prospective longitudinal case series | Cataract surgery | Refractive EDOF IOL | Lentis comfort LS-313 MF15 |  |  |  | VA | UIVA | uncorrected | Intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 44 | 0.21 | 0.15
732 | Vounotrypidis (2017) | Prospective longitudinal case series | Cataract surgery | Refractive EDOF IOL | Lentis comfort LS-313 MF15 |  |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Monocular |  | 3개월 | 안 | 44 | 0.53 | 0.15
732 | Vounotrypidis (2017) | Prospective longitudinal case series | Cataract surgery | Refractive EDOF IOL | Lentis comfort LS-313 MF15 |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | 3개월 | 안 | 44 | 0.05 | 0.09
732 | Vounotrypidis (2017) | Prospective longitudinal case series | Cataract surgery | Refractive EDOF IOL | Lentis comfort LS-313 MF15 |  |  |  | VA | CIVA | corrected | intermediate(80cm) | LogMAR | Monocular |  | 3개월 | 안 | 44 | 0.17 | 0.15
732 | Vounotrypidis (2017) | Prospective longitudinal case series | Cataract surgery | Refractive EDOF IOL | Lentis comfort LS-313 MF15 |  |  |  | VA | CNVA | corrected | Near(33cm) | LogMAR | Monocular |  | 3개월 | 안 | 44 | 0.4 | 0.11
732 | Vounotrypidis (2017) | Prospective longitudinal case series | Cataract surgery | Refractive EDOF IOL | Lentis comfort LS-313 MF15 |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular |  | 3개월 | 명 | 22 | 0.01 | 0.08
732 | Vounotrypidis (2017) | Prospective longitudinal case series | Cataract surgery | Refractive EDOF IOL | Lentis comfort LS-313 MF15 |  |  |  | complication | Posterior capsule opacification |  |  | 0-4점(0점 best) |  |  | 3개월 | 명 | 22 | 0.23 | 0.16
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | LogMAR | Monocular |  | 수술전 | 안 | 70 | 0.05 | 0.15
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | UDVA | uncorrected | Distant(3m) | LogMAR | Monocular |  | 3개월 | 안 | 72 | 0.01 | 0.11
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | CDVA | corrected | Distant(3m) | LogMAR | Monocular |  | 3개월 | 안 | 72 | -0.09 | 0.06
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | UIVA | uncorrected | Intermediate(67cm) | LogMAR | Monocular |  | 3개월 | 안 | 72 | 0.24 | 0.14
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | DCIVA | distance corrected | Intermediate(67cm) | LogMAR | Monocular |  | 3개월 | 안 | 72 | 0.3 | 0.12
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Monocular |  | 3개월 | 안 | 72 | 0.08 | 0.07
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | DCNVA | distance corrected | Near(33cm) | LogMAR | Monocular |  | 3개월 | 안 | 72 | 0.06 | 0.05
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | UDVA | uncorrected | Distant(3m) | LogMAR | Monocular |  | 6개월 | 안 | 72 | 0.01 | 0.12
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | CDVA | corrected | Distant(3m) | LogMAR | Monocular |  | 6개월 | 안 | 72 | -0.1 | 0.05
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | UIVA | uncorrected | Intermediate(67cm) | LogMAR | Monocular |  | 6개월 | 안 | 72 | 0.26 | 0.17
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | DCIVA | distance corrected | Intermediate(67cm) | LogMAR | Monocular |  | 6개월 | 안 | 72 | 0.31 | 0.15
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Monocular |  | 6개월 | 안 | 72 | 0.09 | 0.08
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | DCNVA | distance corrected | Near(33cm) | LogMAR | Monocular |  | 6개월 | 안 | 72 | 0.06 | 0.07
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | UDVA | uncorrected | Distant(3m) | LogMAR | Binocular |  | 3개월 | 명 | 36 | -0.06 | 0.08
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | CDVA | corrected | Distant(3m) | LogMAR | Binocular |  | 3개월 | 명 | 36 | -0.11 | 0.04
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mespic | VA | CDVA | corrected | Distant(3m) | LogMAR | Binocular |  | 3개월 | 명 | 36 | -0.11 | 0.04
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | UIVA | uncorrected | Intermediate(67cm) | LogMAR | Binocular |  | 3개월 | 명 | 36 | 0.13 | 0.11
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | DCIVA | distance corrected | Intermediate(67cm) | LogMAR | Binocular |  | 3개월 | 명 | 36 | 0.18 | 0.09
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mespic | VA | DCIVA | distance corrected | Intermediate(67cm) | LogMAR | Binocular |  | 3개월 | 명 | 36 | 0.37 | 0.12
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular |  | 3개월 | 명 | 36 | 0.04 | 0.06
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | DCNVA | distance corrected | Near(33cm) | LogMAR | Binocular |  | 3개월 | 명 | 36 | 0.01 | 0.05
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mespic | VA | DCNVA | distance corrected | Near(33cm) | LogMAR | Binocular |  | 3개월 | 명 | 36 | 0.13 | 0.07
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | UDVA | uncorrected | Distant(3m) | LogMAR | Binocular |  | 6개월 | 명 | 36 | -0.05 | 0.09
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | CDVA | corrected | Distant(3m) | LogMAR | Binocular |  | 6개월 | 명 | 36 | -0.11 | 0.04
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mespic | VA | CDVA | corrected | Distant(3m) | LogMAR | Binocular |  | 6개월 | 명 | 36 | -0.11 | 0.04
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | UIVA | uncorrected | Intermediate(67cm) | LogMAR | Binocular |  | 6개월 | 명 | 36 | 0.12 | 0.13
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | DCIVA | distance corrected | Intermediate(67cm) | LogMAR | Binocular |  | 6개월 | 명 | 36 | 0.18 | 0.14
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mespic | VA | DCIVA | distance corrected | Intermediate(67cm) | LogMAR | Binocular |  | 6개월 | 명 | 36 | 0.4 | 0.11
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular |  | 6개월 | 명 | 36 | 0.04 | 0.06
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | photopic | VA | DCNVA | distance corrected | Near(33cm) | LogMAR | Binocular |  | 6개월 | 명 | 36 | 0.02 | 0.06
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mespic | VA | DCNVA | distance corrected | Near(33cm) | LogMAR | Binocular |  | 6개월 | 명 | 36 | 0.12 | 0.07
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 57 | 0.2 | 0.19
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 17 | 0.15 | 0.15
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 15 | 0.13 | 0.15
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 12 | 0.33 | 0.26
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 13 | 0.21 | 0.18
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 57 | 0.12 | 0.16
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 57 | 0.04 | 0.1
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 57 | 0.02 | 0.08
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 57 | 0 | 0.08
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 57 | 0.1 | 0.14
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 57 | 0.06 | 0.12
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | CNVA | corrected | near(40cm) | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 57 | 0.08 | 0.14
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | CNVA | corrected | near(40cm) | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 57 | 0.03 | 0.09
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 57 | 0.21 | 0.13
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 57 | 0.21 | 0.2
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 17 | 0.11 | 0.11
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | UDVA | uncorrected | Distant | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 17 | 0.02 | 0.09
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 17 | 0.01 | 0.08
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | CDVA | corrected | Distant | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 17 | -0.01 | 0.07
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 17 | 0.08 | 0.12
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 17 | 0.01 | 0.07
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | CNVA | corrected | near(40cm) | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 17 | 0.05 | 0.12
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | CNVA | corrected | near(40cm) | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 17 | -0.01 | 0.07
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 17 | 0.2 | 0.16
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 17 | 0.13 | 0.14
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 15 | 0.09 | 0.14
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | UDVA | uncorrected | Distant | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 15 | 0.03 | 0.05
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 15 | -0.01 | 0.08
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | CDVA | corrected | Distant | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 15 | 0.01 | 0.09
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 15 | 0.08 | 0.13
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 15 | 0.06 | 0.08
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | CNVA | corrected | near(40cm) | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 15 | 0.07 | 0.12
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | CNVA | corrected | near(40cm) | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 15 | 0.03 | 0.08
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 15 | 0.15 | 0.1
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 15 | 0.26 | 0.3
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 12 | 0.23 | 0.23
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | UDVA | uncorrected | Distant | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 12 | 0.02 | 0.09
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 12 | 0.08 | 0.11
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | CDVA | corrected | Distant | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 12 | 0.02 | 0.09
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 12 | 0.21 | 0.18
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 12 | 0.17 | 0.23
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | CNVA | corrected | near(40cm) | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 12 | 0.17 | 0.22
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | CNVA | corrected | near(40cm) | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 12 | 0 | 0
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 12 | 0.28 | 0.12
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 12 | 0.2 | 0.17
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 13 | 0.1 | 0.13
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | UDVA | uncorrected | Distant | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 13 | 0.07 | 0.14
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 13 | 0.02 | 0.06
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | CDVA | corrected | Distant | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 13 | -0.02 | 0.08
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 13 | 0.05 | 0.1
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 13 | 0.05 | 0.11
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | CNVA | corrected | near(40cm) | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 13 | 0.07 | 0.11
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | CNVA | corrected | near(40cm) | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 13 | 0.08 | 0.13
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular |  | final visit(2-4개월) | 안 | 13 | 0.22 | 0.12
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Binocular |  | final visit(2-4개월) | 안 | 13 | 0.23 | 0.14
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 106 | 0.12 | 0.12
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | VA | UNVA | uncorrected | Near(40cm) | LogMAR | Binocular |  | 3개월 | 안 | 53 | 0.08 | 0.1
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | VA | DCNVA | distance corrected | near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 106 | 0.11 | 0.1
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | VA | DCNVA | distance corrected | near(40cm) | LogMAR | Binocular |  | 3개월 | 안 | 53 | 0.08 | 0.09
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | VA | UIVA | uncorrected | Intermediate(70cm) | LogMAR | Monocular |  | 3개월 | 안 | 106 | 0.16 | 0.11
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | VA | UIVA | uncorrected | Intermediate(70cm) | LogMAR | Binocular |  | 3개월 | 안 | 53 | 0.13 | 0.1
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | VA | DCIVA | distance corrected | Intermediate(70cm) | LogMAR | Monocular |  | 3개월 | 안 | 106 | 0.15 | 0.1
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | VA | DCIVA | distance corrected | Intermediate(70cm) | LogMAR | Binocular |  | 3개월 | 안 | 53 | 0.1 | 0.09
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 3개월 | 안 | 106 | -0.03 | 0.09
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular |  | 3개월 | 안 | 53 | -0.08 | 0.08
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | 3개월 | 안 | 106 | -0.07 | 0.06
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Binocular |  | 3개월 | 안 | 53 | -0.1 | 0.06
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  | Night vision phenomena | Satisfaction | starbursts |  |  | 1-7점(1점 best) |  |  | 3개월 | 명 | 53 | 2.8 | 1.5
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  | Night vision phenomena | Satisfaction | glare |  |  | 1-7점(1점 best) |  |  | 3개월 | 명 | 53 | 3 | 1.6
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  | Night vision phenomena | Satisfaction | halos |  |  | 1-7점(1점 best) |  |  | 3개월 | 명 | 53 | 3.2 | 1.6
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  | Night vision phenomena | Satisfaction | double vision or ghosting |  |  | 1-7점(1점 best) |  |  | 3개월 | 명 | 53 | 2.5 | 1.6
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 60 | 0.53 | 0.47
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 60 | 0.02 | 0.21
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Monocular |  | 수술전 | 안 | 60 | 0.92 | 0.26
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | CNVA | corrected | Near(33cm) | LogMAR | Monocular |  | 수술전 | 안 | 60 | 0.17 | 0.19
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | DCNVA | distance corrected | Near(33cm) | LogMAR | Monocular |  | 수술전 | 안 | 60 | 0.68 | 0.19
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular |  | 수술전 | 안 | 60 | 0.76 | 0.27
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | CIVA | corrected | Intermediate(66cm) | LogMAR | Monocular |  | 수술전 | 안 | 60 | 0.13 | 0.23
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | DCIVA | distance corrected | Intermediate(66cm) | LogMAR | Monocular |  | 수술전 | 안 | 60 | 0.43 | 0.26
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 1개월 | 안 | 60 | -0.03 | 0.08
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | 1개월 | 안 | 60 | -0.05 | 0.07
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Monocular |  | 1개월 | 안 | 60 | 0.22 | 0.13
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | CNVA | corrected | Near(33cm) | LogMAR | Monocular |  | 1개월 | 안 | 60 | 0.2 | 0.11
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | DCNVA | distance corrected | Near(33cm) | LogMAR | Monocular |  | 1개월 | 안 | 60 | 0.2 | 0.11
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular |  | 1개월 | 안 | 60 | 0.08 | 0.11
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | CIVA | corrected | Intermediate(66cm) | LogMAR | Monocular |  | 1개월 | 안 | 60 | 0.07 | 0.1
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | DCIVA | distance corrected | Intermediate(66cm) | LogMAR | Monocular |  | 1개월 | 안 | 60 | 0.07 | 0.1
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 3개월 | 안 | 60 | -0.04 | 0.1
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | 3개월 | 안 | 60 | -0.06 | 0.09
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Monocular |  | 3개월 | 안 | 60 | 0.19 | 0.11
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | CNVA | corrected | Near(33cm) | LogMAR | Monocular |  | 3개월 | 안 | 60 | 0.14 | 0.1
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | DCNVA | distance corrected | Near(33cm) | LogMAR | Monocular |  | 3개월 | 안 | 60 | 0.17 | 0.1
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular |  | 3개월 | 안 | 60 | 0.11 | 0.1
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | CIVA | corrected | Intermediate(66cm) | LogMAR | Monocular |  | 3개월 | 안 | 60 | 0.08 | 0.1
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | DCIVA | distance corrected | Intermediate(66cm) | LogMAR | Monocular |  | 3개월 | 안 | 60 | 0.1 | 0.09
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular |  | 6개월 | 안 | 60 | -0.03 | 0.09
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Monocular |  | 6개월 | 안 | 60 | -0.05 | 0.08
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Monocular |  | 6개월 | 안 | 60 | 0.2 | 0.12
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | CNVA | corrected | Near(33cm) | LogMAR | Monocular |  | 6개월 | 안 | 60 | 0.13 | 0.1
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | DCNVA | distance corrected | Near(33cm) | LogMAR | Monocular |  | 6개월 | 안 | 60 | 0.17 | 0.11
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular |  | 6개월 | 안 | 60 | 0.08 | 0.1
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | CIVA | corrected | Intermediate(66cm) | LogMAR | Monocular |  | 6개월 | 안 | 60 | 0.06 | 0.11
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | VA | DCIVA | distance corrected | Intermediate(66cm) | LogMAR | Monocular |  | 6개월 | 안 | 60 | 0.08 | 0.1
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  | photopic | CS | 1.5CPD |  |  | log |  |  | 1개월 | 안 | 60 | 1.57 | 0.13
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  | photopic | CS | 1.5CPD |  |  | log |  |  | 6개월 | 안 | 60 | 1.62 | 0.15
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  | photopic | CS | 12CPD |  |  | log |  |  | 1개월 | 안 | 60 | 1.35 | 0.15
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  | photopic | CS | 12CPD |  |  | log |  |  | 6개월 | 안 | 60 | 1.4 | 0.16
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  | photopic | CS | 18CPD |  |  | log |  |  | 1개월 | 안 | 60 | 0.73 | 0.22
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  | photopic | CS | 18CPD |  |  | log |  |  | 6개월 | 안 | 60 | 0.85 | 0.25
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | Satisfaction | Television |  |  | 1-6점(1점 best) |  |  | 6개월 | 명 | 30 | 1.13 | 0.35
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | Satisfaction | Therater, concert |  |  | 1-6점(1점 best) |  |  | 6개월 | 명 | 30 | 1.23 | 0.43
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | Satisfaction | Driving at daytime |  |  | 1-6점(1점 best) |  |  | 6개월 | 명 | 30 | 1.33 | 0.48
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | Satisfaction | At home |  |  | 1-6점(1점 best) |  |  | 6개월 | 명 | 30 | 1.17 | 0.38
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | Satisfaction | Driving at night |  |  | 1-6점(1점 best) |  |  | 6개월 | 명 | 30 | 2.57 | 0.77
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | Satisfaction | cooking |  |  | 1-6점(1점 best) |  |  | 6개월 | 명 | 30 | 1.13 | 0.35
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | Satisfaction | Newspaper |  |  | 1-6점(1점 best) |  |  | 6개월 | 명 | 30 | 1.67 | 0.71
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | Satisfaction | computer |  |  | 1-6점(1점 best) |  |  | 6개월 | 명 | 30 | 1.67 | 0.8
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | Satisfaction | Homework |  |  | 1-6점(1점 best) |  |  | 6개월 | 명 | 30 | 1.1 | 0.31
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | Satisfaction | Overall |  |  | 1-6점(1점 best) |  |  | 6개월 | 명 | 30 | 1.43 | 0.57
1288 | Tsaousis (2013) | Observational case study | Cataract surgery | Diffractive multifocal IOL | AcrySof IQ ReSTOR IOL SN60D3 |  |  | worse eye | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 26개월 | 안 | 20 | 0.21 | 0.12
1288 | Tsaousis (2013) | Observational case study | Cataract surgery | Diffractive multifocal IOL | AcrySof IQ ReSTOR IOL SN60D3 |  |  | better eye | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 26개월 | 안 | 20 | 0.07 | 0.1
1288 | Tsaousis (2013) | Observational case study | Cataract surgery | Diffractive multifocal IOL | AcrySof IQ ReSTOR IOL SN60D3 |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular |  | 26개월 | 명 | 20 | 0 | 0.09
1288 | Tsaousis (2013) | Observational case study | Cataract surgery | Diffractive multifocal IOL | AcrySof IQ ReSTOR IOL SN60D3 |  |  | worse eye | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular |  | 26개월 | 안 | 20 | 0.32 | 0.15
1288 | Tsaousis (2013) | Observational case study | Cataract surgery | Diffractive multifocal IOL | AcrySof IQ ReSTOR IOL SN60D3 |  |  | better eye | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular |  | 26개월 | 안 | 20 | 0.18 | 0.14
1288 | Tsaousis (2013) | Observational case study | Cataract surgery | Diffractive multifocal IOL | AcrySof IQ ReSTOR IOL SN60D3 |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Binocular |  | 26개월 | 명 | 20 | 0.08 | 0.15
1288 | Tsaousis (2013) | Observational case study | Cataract surgery | Diffractive multifocal IOL | AcrySof IQ ReSTOR IOL SN60D3 |  |  | worse eye | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Monocular |  | 26개월 | 안 | 20 | 0.32 | 0.12
1288 | Tsaousis (2013) | Observational case study | Cataract surgery | Diffractive multifocal IOL | AcrySof IQ ReSTOR IOL SN60D3 |  |  | better eye | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Monocular |  | 26개월 | 안 | 20 | 0.2 | 0.09
1288 | Tsaousis (2013) | Observational case study | Cataract surgery | Diffractive multifocal IOL | AcrySof IQ ReSTOR IOL SN60D3 |  |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular |  | 26개월 | 명 | 20 | 0.11 | 0.1
1288 | Tsaousis (2013) | Observational case study | Cataract surgery | Diffractive multifocal IOL | AcrySof IQ ReSTOR IOL SN60D3 |  |  |  | VA | binocular summation |  | Distant(4m) | LogMAR |  | defined as the difference between the binocular and better eye VA | 26개월 | 명 | 20 | 0.07 | 0.05
1288 | Tsaousis (2013) | Observational case study | Cataract surgery | Diffractive multifocal IOL | AcrySof IQ ReSTOR IOL SN60D3 |  |  |  | VA | binocular summation |  | Intermediate(66cm) | LogMAR |  | defined as the difference between the binocular and better eye VA | 26개월 | 명 | 20 | 0.1 | 0.11
1288 | Tsaousis (2013) | Observational case study | Cataract surgery | Diffractive multifocal IOL | AcrySof IQ ReSTOR IOL SN60D3 |  |  |  | VA | binocular summation |  | Near(33cm) | LogMAR |  | defined as the difference between the binocular and better eye VA | 26개월 | 명 | 20 | 0.09 | 0.08
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | BCVA | best-corrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 30 | 0.03 | 0.06
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | UCVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 6개월 | 안 | 30 | 0.08 | 0.15
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | BCVA | best-corrected | Distant(4m) | LogMAR | Monocular |  | 6개월 | 안 | 30 | 0.01 | 0.02
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | BCVA | best-corrected | Distant(4m) | LogMAR | Monocular |  | 수술전 | 안 | 30 | 0.05 | 0.12
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | UCVA | uncorrected | Distant(4m) | LogMAR | Monocular |  | 6개월 | 안 | 30 | 0.11 | 0.14
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | BCVA | best-corrected | Distant(4m) | LogMAR | Monocular |  | 6개월 | 안 | 30 | 0.04 | 0.1
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | BCNVA | best-corrected | near(40cm) | LogMAR | Monocular |  | 수술전 | 안 | 30 | 0.1 | 0.01
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | UCVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 6개월 | 안 | 30 | 0.03 | 0.05
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | BCNVA | best-corrected | near(40cm) | LogMAR | Monocular |  | 6개월 | 안 | 30 | 0.01 | 0.02
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | BCNVA | best-corrected | near(40cm) | LogMAR | Monocular |  | 수술전 | 안 | 30 | 0.1 | 0.01
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | UCVA | uncorrected | Near(40cm) | LogMAR | Monocular |  | 6개월 | 안 | 30 | 0.08 | 0.01
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | BCNVA | best-corrected | near(40cm) | LogMAR | Monocular |  | 6개월 | 안 | 30 | 0.04 | 0.06

## 결과지표_범주형_IOL
3. 범주형 결과변수
no. | 1저자 (연도) | 연구설계 | 대상자 | 중재군명 | 렌즈명 | 대조군명 | 렌즈명 | 하위그룹 | 결과그룹 | 결과지표명 | 구분(나안, 교정( | 구분(거리) | 단위 | 안구기준 (Monocular, Binocular) | 정의 | 측정시점 (개월) | 분석단위(명/안) | 중재군 |  | 대조군 |  | 두 군간 차이 | OR, RR, HR |  | 95% CI | 비고
[18 empty] | Total N | event | Total N | event | P value | 단위 | risk rate
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UDVA |  | Distant |  | Monocular | 20/25 or better | 3개월 | 안 | 100 | 91
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UDVA |  | Distant |  | Monocular | 20/30 or better | 3개월 | 안 | 100 | 99
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UIVA |  | Intermediate(66cm) |  | Monocular | 20/25 or better | 3개월 | 안 | 100 | 79
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UIVA |  | Intermediate(66cm) |  | Monocular | 20/30 or better | 3개월 | 안 | 100 | 94
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UNVA |  | Near(40cm) |  | Monocular | 20/25 or better | 3개월 | 안 | 100 | 87
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | VA | UNVA |  | Near(40cm) |  | Monocular | 20/30 or better | 3개월 | 안 | 100 | 99
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | Complication | posterior capsule opacification | [5 empty] | 3개월 | 명 | 50 | 0
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | Satisfaction | completely happy | [5 empty] | 3개월 | 명 | 50 | 40
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | Satisfaction | moderate happy | [5 empty] | 3개월 | 명 | 50 | 8
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | Satisfaction | unhappy | [5 empty] | 3개월 | 명 | 50 | 2
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  | Satisfaction | Choose same IOL again | [5 empty] | 3개월 | 명 | 50 | 44
10_국내 | Kretz (2006) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  |  |  | recommend the surgery and IOL | [5 empty] | 3개월 | 명 | 50 | 43
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  |  | Spectacle independence | never or only occasionally use spectacle across all distance | [5 empty] | 4-6개월 | 명 | 96 | 72.96
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  | Toric | Spectacle independence | never or only occasionally use spectacle across all distance | [5 empty] | 4-6개월 | 명 | 11 | 8.25
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  |  | Spectacle independence | never or only occasionally use spectacle |  | Reading distance |  |  |  | 4-6개월 | 명 | 96 | 92.16
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  | Toric | Spectacle independence | never or only occasionally use spectacle |  | Reading distance |  |  |  | 4-6개월 | 명 | 11 | 10.12
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  |  | Spectacle independence | required spectacles most of time |  | Intermediate task |  |  |  | 4-6개월 | 명 | 96 | 29.76
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  | Toric | Spectacle independence |  |  | Intermediate task |  |  |  | 4-6개월 | 명 | 11 | 11
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  |  | Spectacle independence | frequently needed reading glass | [5 empty] | 4-6개월 | 명 | 96 | 12.96
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  | Toric | Spectacle independence | frequently needed reading glass | [5 empty] | 4-6개월 | 명 | 11 | 0.913
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  |  | Satisfaction | recommend the IOL | [5 empty] | 4-6개월 | 명 | 96 | 78
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  |  | Satisfaction | Choose same IOL again | [5 empty] | 4-6개월 | 명 | 96 | 77.0016
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  |  | Complication | macular edema | [6 empty] | 명 | 96 | 1
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  |  | Complication | refractive touch-up treatment | [6 empty] | 명 | 96 | 5
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | Diffractive EDOF | TECNIS Symfony(ZXR00), TECNIS Symfony Toric |  |  |  | Complication | IOL decentration or tilt | [6 empty] | 명 | 96 | 0
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | <0.0 | 3개월 | 안 | 60 | 36.6
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | <0.1 | 3개월 | 안 | 60 | 46.8
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | <0.2 | 3개월 | 안 | 60 | 60
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  |  | VA | UIVA | uncorrected | intermediate(80cm) | LogMAR | Binocular | <0.0 | 3개월 | 안 | 60 | 34.8
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  |  | VA | UIVA | uncorrected | intermediate(80cm) | LogMAR | Binocular | <0.1 | 3개월 | 안 | 60 | 40.002
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  |  | VA | UIVA | uncorrected | intermediate(80cm) | LogMAR | Binocular | <0.2 | 3개월 | 안 | 60 | 60
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  |  | Satisfaction | nighttime driving was possible without any complication | [5 empty] | 3개월 | 명 | 30 | 30
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  |  | Spectacle independence | intermdiate task |  | <60cm | Y/N |  |  | 3개월 | 명 | 30 | 26
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  |  | Spectacle independence | intermdiate task |  | >60cm | Y/N |  |  | 3개월 | 명 | 30 | 30
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  |  | Satisfaction | very satisfied |  |  | Y/N |  |  | 3개월 | 명 | 30 | 30
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  |  | Satisfaction | Choose same IOL again |  |  | Y/N |  |  | 3개월 | 명 | 30 | 30
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  |  | Satisfaction | recommend IOL |  |  | Y/N |  |  | 3개월 | 명 | 30 | 30
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  |  | Satisfaction | halo or glare |  |  | Y/N |  |  | 3개월 | 명 | 30 | 0
71_국내 | Kretz (2016) | Prospective non-comparative non-randomized clinical study | Cataract surgery | Refractive Bifocal IOL | Lentis-313MF15 |  |  |  | Complication | posterior capsule rupture, endophthalmitis, or corneal decompensation |  |  | Y/N |  |  | 3개월 | 명 | 30 | 0
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Satisfaction | 매우좋음 |  | near(40cm) | Y/N |  |  |  |  | 20 | 11
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Satisfaction | 좋음 |  | near(40cm) | Y/N |  |  |  |  | 20 | 8
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Satisfaction | 보통 |  | near(40cm) | Y/N |  |  |  |  | 20 | 1
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Satisfaction | 나쁨 |  | near(40cm) | Y/N |  |  |  |  | 20 | 0
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Satisfaction | 매우나쁨 |  | near(40cm) | Y/N |  |  |  |  | 20 | 0
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Satisfaction | 매우좋음 |  | intermediate(80cm) | Y/N |  |  |  |  | 20 | 9
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Satisfaction | 좋음 |  | intermediate(80cm) | Y/N |  |  |  |  | 20 | 7
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Satisfaction | 보통 |  | intermediate(80cm) | Y/N |  |  |  |  | 20 | 3
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Satisfaction | 나쁨 |  | intermediate(80cm) | Y/N |  |  |  |  | 20 | 1
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Satisfaction | 매우나쁨 |  | intermediate(80cm) | Y/N |  |  |  |  | 20 | 0
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Satisfaction | 매우좋음 |  | Distant(4m) | Y/N |  |  |  |  | 20 | 7
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Satisfaction | 좋음 |  | Distant(4m) | Y/N |  |  |  |  | 20 | 6
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Satisfaction | 보통 |  | Distant(4m) | Y/N |  |  |  |  | 20 | 5
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Satisfaction | 나쁨 |  | Distant(4m) | Y/N |  |  |  |  | 20 | 2
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Satisfaction | 매우나쁨 |  | Distant(4m) | Y/N |  |  |  |  | 20 | 0
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Satisfaction | Choose same IOL again |  |  | Y/N |  |  |  |  | 20 | 18
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Spectacle independence | 안경없이 작업이 가능한가? |  | Distant(4m) | [5 empty] | 20 | 20
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Spectacle independence | 안경없이 작업이 가능한가? |  | intermediate(80cm) | [5 empty] | 20 | 19
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Spectacle independence | 안경없이 작업이 가능한가? |  | near(40cm) | [5 empty] | 20 | 17
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Spectacle independence | 안경이 얼마나 자주 필요한가?(전혀) |  | Distant(4m) | [5 empty] | 20 | 20
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Spectacle independence | 안경이 얼마나 자주 필요한가?(가끔) |  | Distant(4m) | [5 empty] | 20 | 0
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Spectacle independence | 안경이 얼마나 자주 필요한가?(절반) |  | Distant(4m) | [5 empty] | 20 | 0
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Spectacle independence | 안경이 얼마나 자주 필요한가?(자주) |  | Distant(4m) | [5 empty] | 20 | 0
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Spectacle independence | 안경이 얼마나 자주 필요한가?(항상) |  | Distant(4m) | [5 empty] | 20 | 0
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Spectacle independence | 안경이 얼마나 자주 필요한가?(전혀) |  | intermediate(80cm) | [5 empty] | 20 | 19
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Spectacle independence | 안경이 얼마나 자주 필요한가?(가끔) |  | intermediate(80cm) | [5 empty] | 20 | 0
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Spectacle independence | 안경이 얼마나 자주 필요한가?(절반) |  | intermediate(80cm) | [5 empty] | 20 | 1
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Spectacle independence | 안경이 얼마나 자주 필요한가?(자주) |  | intermediate(80cm) | [5 empty] | 20 | 0
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Spectacle independence | 안경이 얼마나 자주 필요한가?(항상) |  | intermediate(80cm) | [5 empty] | 20 | 0
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Spectacle independence | 안경이 얼마나 자주 필요한가?(전혀) |  | near(40cm) | [5 empty] | 20 | 17
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Spectacle independence | 안경이 얼마나 자주 필요한가?(가끔) |  | near(40cm) | [5 empty] | 20 | 2
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Spectacle independence | 안경이 얼마나 자주 필요한가?(절반) |  | near(40cm) | [5 empty] | 20 | 1
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Spectacle independence | 안경이 얼마나 자주 필요한가?(자주) |  | near(40cm) | [5 empty] | 20 | 0
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | Diffractive Trifocal IOL | AT Lisa Tri 839MP |  |  |  | Spectacle independence | 안경이 얼마나 자주 필요한가?(항상) |  | near(40cm) | [5 empty] | 20 | 0
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis |  | AcrySof ReSTOR D1 |  | Satisfaction | 5점 |  |  | 1-5점(5점 best) |  |  | 2개월 | 명 | 11 | 6 | 10 | NA
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis |  | AcrySof ReSTOR D1 |  | Satisfaction | 4점 |  |  | 1-5점(5점 best) |  |  | 2개월 | 명 | 11 | 4 | 10 | NA
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis |  | AcrySof ReSTOR D1 |  | Satisfaction | 3점 |  |  | 1-5점(5점 best) |  |  | 2개월 | 명 | 11 | 1 | 10 | NA
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis |  | AcrySof ReSTOR D1 |  | Spectacle independence | [6 empty] | 2개월 | 명 | 11 | 11 | 10 | NA
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis |  | AcrySof ReSTOR D1 |  | Satisfaction | 달무리 | [5 empty] | 2개월 | 명 | 11 | 1 | 10 | 4
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis |  | AcrySof ReSTOR D1 |  | Satisfaction | 빛번짐 | [5 empty] | 2개월 | 명 | 11 | 1 | 10 | 0
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis |  | AcrySof ReSTOR D1 |  | Satisfaction | 눈부심 | [5 empty] | 2개월 | 명 | 11 | 0 | 10 | 4
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | Hybrid Bifocal IOL | OptiVis |  | AcrySof ReSTOR D1 |  | Complication | 수정체낭 혼탁 | [5 empty] | 2개월 | 명 | 11 | 0 | 10 | 0
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | CDVA | corrected | Distant(6m) | Snellen | Binocular | lost>2 line | 6개월 | 안 | 46 | 0
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | CDVA | corrected | Distant(6m) | Snellen | Binocular | lost 2 line | 6개월 | 안 | 46 | 0
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | CDVA | corrected | Distant(6m) | Snellen | Binocular | lost 1 line | 6개월 | 안 | 46 | 5
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | CDVA | corrected | Distant(6m) | Snellen | Binocular | 변화없음 | 6개월 | 안 | 46 | 24
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | CDVA | corrected | Distant(6m) | Snellen | Binocular | gain 1 line | 6개월 | 안 | 46 | 14
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | CDVA | corrected | Distant(6m) | Snellen | Binocular | gain 2 line | 6개월 | 안 | 46 | 3
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | CDVA | corrected | Distant(6m) | Snellen | Binocular | gain >2 line | 6개월 | 안 | 46 | 0
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | CNVA | corrected | Near(33cm) | Snellen | Binocular | lost>2 line | 6개월 | 안 | 46 | 0
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | CNVA | corrected | Near(33cm) | Snellen | Binocular | lost 2 line | 6개월 | 안 | 46 | 1
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | CNVA | corrected | Near(33cm) | Snellen | Binocular | lost 1 line | 6개월 | 안 | 46 | 3
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | CNVA | corrected | Near(33cm) | Snellen | Binocular | 변화없음 | 6개월 | 안 | 46 | 34
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | CNVA | corrected | Near(33cm) | Snellen | Binocular | gain 1 line | 6개월 | 안 | 46 | 8
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | CNVA | corrected | Near(33cm) | Snellen | Binocular | gain 2 line | 6개월 | 안 | 46 | 0
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | VA | CNVA | corrected | Near(33cm) | Snellen | Binocular | gain >2 line | 6개월 | 안 | 46 | 0
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | Spectacle dependence |  |  | Near | Y/N |  |  | 6개월 | 명 | 23 | 0
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | Spectacle dependence |  |  | Intermediate | Y/N |  |  | 6개월 | 명 | 23 | 4
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | Spectacle dependence |  |  | Distant | Y/N |  |  | 6개월 | 명 | 23 | 0
59 | Alfonso (2010) | Prospective, nonrandomized, masked, observational case series (self-controlled) | Emmetropic presbyopia | Diffractive multifocal IOL? | AcrySof ReSTOR Natural (SN60D3) multifocal IOL |  |  |  | Complication | [6 empty] | 6개월 | 명 | 23 | 0
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP |  | Satisfaction |  |  |  | Y/N |  |  | 6개월 | 명 | 93 | 89.28 | 89 | 86.33
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | first implanted eye | Complication | posterior capsule opacification | [5 empty] | 6개월 | 명 | 93 | 3 | 87 | 5
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | first implanted eye | Complication | halo | [5 empty] | 6개월 | 명 | 93 | 5 | 87 | 2
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | first implanted eye | Complication | virteous detachment | [5 empty] | 6개월 | 명 | 93 | 1 | 87 | 3
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | first implanted eye | Complication | dry eye | [5 empty] | 6개월 | 명 | 93 | 1 | 87 | 4
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | first implanted eye | Complication | corneal edema | [5 empty] | 6개월 | 명 | 93 | 2 | 87 | 2
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | first implanted eye | Complication | reduced visual acuity | [5 empty] | 6개월 | 명 | 93 | 2 | 87 | 1
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | first implanted eye | Complication | blepharitis | [5 empty] | 6개월 | 명 | 93 | 0 | 87 | 2
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | first implanted eye | Complication | cystoid macular edema | [5 empty] | 6개월 | 명 | 93 | 0 | 87 | 2
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | first implanted eye | Complication | Foreign body sensation | [5 empty] | 6개월 | 명 | 93 | 2 | 87 | 0
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | first implanted eye | Complication | iridocyclitis | [5 empty] | 6개월 | 명 | 93 | 2 | 87 | 0
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | second implanted eye | Complication | halo | [5 empty] | 6개월 | 명 | 93 | 5 | 87 | 2
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | second implanted eye | Complication | dry eye | [5 empty] | 6개월 | 명 | 93 | 2 | 87 | 3
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | second implanted eye | Complication | Increased IOP | [5 empty] | 6개월 | 명 | 93 | 1 | 87 | 3
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | second implanted eye | Complication | posterior capsule opacification | [5 empty] | 6개월 | 명 | 93 | 3 | 87 | 1
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | second implanted eye | Complication | virteous detachment | [5 empty] | 6개월 | 명 | 93 | 1 | 87 | 3
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | second implanted eye | Complication | Foreign body sensation | [5 empty] | 6개월 | 명 | 93 | 3 | 87 | 0
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | second implanted eye | Complication | iridocyclitis | [5 empty] | 6개월 | 명 | 93 | 2 | 87 | 1
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | second implanted eye | Complication | reduced visual acuity | [5 empty] | 6개월 | 명 | 93 | 2 | 87 | 1
202 | Lapid-Gortzak (2020) | Prospective, parallel-group, randomized, double-masked, postmarketing clinical study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix TFNT00 | Diffractive trifocal IOL | AT LISA tri 839MP | second implanted eye | Complication | cystoid macular edema | [5 empty] | 6개월 | 명 | 93 | 0 | 87 | 2
272 | Pedrotti (2017) | Prospective single-blind comparative study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF | Monofocal IOL | Tecnis 1-piece |  | Complication | posterior capsule opacification | [5 empty] | 12개월 | 명 | 21 | 3 | 21 | 4
272 | Pedrotti (2017) | Prospective single-blind comparative study | Cataract surgery | Refractive EDOF toric IOL | Lentis Comfort LS-313 MF | Monofocal IOL | Tecnis 1-piece |  | Complication | IOL decentration or tilt | [5 empty] | 12개월 | 명 | 21 | 0 | 21 | 0
323 | Gundersen (2016) | Randomized prospective study | Cataract surgery | Diffractive Trifocal toric IOL | FineVision IOL (PhysIOL ® ; FineVision Toric, Liège, Belgium) | Diffractive bifocal toric IOL | ReSTOR SND1T |  | Complication | [6 empty] | 3개월 | 명 | 11 | 0 | 11 | 0
323 | Gundersen (2016) | Randomized prospective study | Cataract surgery | Diffractive Trifocal toric IOL | FineVision IOL (PhysIOL ® ; FineVision Toric, Liège, Belgium) | Diffractive bifocal toric IOL | ReSTOR SND1T |  | VA | 20/25 | uncorrected | Intermediate(63cm) | Snellen | Binocular |  | 3개월 | 명 | 11 | 10 | 11 | 7
323 | Gundersen (2016) | Randomized prospective study | Cataract surgery | Diffractive Trifocal toric IOL | FineVision IOL (PhysIOL ® ; FineVision Toric, Liège, Belgium) | Diffractive bifocal toric IOL | ReSTOR SND1T |  | VA | better than 20/25 | distance-corrected | Intermediate(63cm) | Snellen | Binocular | 0.1 logMAR | 3개월 | 명 | 11 | 9 | 11 | 7
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | Complication | [6 empty] | 4개월 | 명 | 20 | 0 | 20 | 0 |  |  |  |  | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | Complication | [6 empty] | 4개월 | 명 | 20 | 0 | 20 | 0 |  |  |  |  | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for any purpose |  |  |  |  | always | 4개월 | 명 | 20 | 0 | 20 | 1 |  |  |  |  | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for any purpose |  |  |  |  | sometimes | 4개월 | 명 | 20 | 3 | 20 | 5 |  |  |  |  | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for any purpose |  |  |  |  | never | 4개월 | 명 | 20 | 17 | 20 | 14 |  |  |  |  | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for any purpose |  |  |  |  | always | 4개월 | 명 | 20 | 0 | 20 | 15 |  |  |  |  | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for any purpose |  |  |  |  | sometimes | 4개월 | 명 | 20 | 3 | 20 | 5 |  |  |  |  | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for any purpose |  |  |  |  | never | 4개월 | 명 | 20 | 17 | 20 | 0 |  |  |  |  | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for near tasks(eg. Reading print) |  |  |  |  | always | 4개월 | 명 | 20 | 0 | 20 | 1 |  |  |  |  | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for near tasks(eg. Reading print) |  |  |  |  | sometimes | 4개월 | 명 | 20 | 2 | 20 | 4 |  |  |  |  | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for near tasks(eg. Reading print) |  |  |  |  | never | 4개월 | 명 | 20 | 18 | 20 | 15 |  |  |  |  | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for near tasks(eg. Reading print) |  |  |  |  | always | 4개월 | 명 | 20 | 0 | 20 | 20 |  |  |  |  | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for near tasks(eg. Reading print) |  |  |  |  | sometimes | 4개월 | 명 | 20 | 2 | 20 | 0 |  |  |  |  | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for near tasks(eg. Reading print) |  |  |  |  | never | 4개월 | 명 | 20 | 18 | 20 | 0 |  |  |  |  | 3군으로 중재군 n수 50% 입력필요
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for intermediate tasks(eg. computer) |  |  |  |  | always | 4개월 | 명 | 20 | 0 | 20 | 0
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for intermediate tasks(eg. computer) |  |  |  |  | sometimes | 4개월 | 명 | 20 | 0 | 20 | 0
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for intermediate tasks(eg. computer) |  |  |  |  | never | 4개월 | 명 | 20 | 20 | 20 | 20
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for intermediate tasks(eg. computer) |  |  |  |  | always | 4개월 | 명 | 20 | 0 | 20 | 7
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for intermediate tasks(eg. computer) |  |  |  |  | sometimes | 4개월 | 명 | 20 | 0 | 20 | 10
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for intermediate tasks(eg. computer) |  |  |  |  | never | 4개월 | 명 | 20 | 20 | 20 | 3
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for distance tasks(eg. driving) |  |  |  |  | always | 4개월 | 명 | 20 | 0 | 20 | 0
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for distance tasks(eg. driving) |  |  |  |  | sometimes | 4개월 | 명 | 20 | 2 | 20 | 3
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for distance tasks(eg. driving) |  |  |  |  | never | 4개월 | 명 | 20 | 18 | 20 | 17
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for distance tasks(eg. driving) |  |  |  |  | always | 4개월 | 명 | 20 | 0 | 20 | 2
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for distance tasks(eg. driving) |  |  |  |  | sometimes | 4개월 | 명 | 20 | 2 | 20 | 3
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Monofocal IOL | SN60WF (Acrysof, Alcon Laboratories, Inc.) |  | Spectacle dependence | how often do you wear glasses for distance tasks(eg. driving) |  |  |  |  | never | 4개월 | 명 | 20 | 18 | 20 | 15
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | Complication | halo (frequent) |  |  |  |  | quite often | 4개월 | 명 | 20 | 3 | 20 | 5
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | Diffractive trifocal IOL | Acrysof IQ Panoptix TFNT00 | Diffractive EDOF IOL | ZXR00 (Tecnis Symfony, Abbott Laboratories, Inc.) |  | Complication | halo (symptom) |  |  |  |  | moderate and quite | 4개월 | 명 | 20 | 3 | 20 | 4
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular | 0.04(20/20 snellen) or better | 12개월 | 명 | 145 | 101.5
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Binocular | 0.04(20/20 snellen) or better | 12개월 | 명 | 145 | 63.8
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Binocular | 0.04(20/20 snellen) or better | 12개월 | 명 | 145 | 73.95
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Binocular | 0.04(20/20 snellen) or better | 12개월 | 명 | 145 | 63.8
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular | 0.14(20/25 snellen) or better | 12개월 | 명 | 145 | 127.6
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Binocular | 0.14(20/25 snellen) or better | 12개월 | 명 | 145 | 101.5
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Binocular | 0.14(20/25 snellen) or better | 12개월 | 명 | 145 | 117.45
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Binocular | 0.14(20/25 snellen) or better | 12개월 | 명 | 145 | 123.25
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular | 0.24(20/32 snellen) or better | 12개월 | 명 | 145 | 137.75
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Binocular | 0.24(20/32 snellen) or better | 12개월 | 명 | 145 | 129.05
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Binocular | 0.24(20/32 snellen) or better | 12개월 | 명 | 145 | 136.3
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Binocular | 0.24(20/32 snellen) or better | 12개월 | 명 | 145 | 137.75
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Binocular | 0.34(20/40 snellen) or better | 12개월 | 명 | 145 | 143.55
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Binocular | 0.34(20/40 snellen) or better | 12개월 | 명 | 145 | 139.2
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Binocular | 0.34(20/40 snellen) or better | 12개월 | 명 | 145 | 142.1
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Binocular | 0.34(20/40 snellen) or better | 12개월 | 명 | 145 | 140.65
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Binocular | 0.04(20/20 snellen) or better | 12개월 | 명 | 144 | 8.64
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Binocular | 0.14(20/25 snellen) or better | 12개월 | 명 | 144 | 3.456
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Binocular | 0.24(20/32 snellen) or better | 12개월 | 명 | 144 | 73.44
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  |  | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Binocular | 0.34(20/40 snellen) or better | 12개월 | 명 | 144 | 109.44
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | Retinal detachment |  |  |  |  | serious ocular AE | 12개월 | 명 | 149 | 2
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | retinopexy |  |  |  |  | serious ocular AE | 12개월 | 명 | 149 | 2
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | eye operation |  |  |  |  | serious ocular AE | 12개월 | 명 | 149 | 1
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | IOL extraction |  |  |  |  | serious ocular AE | 12개월 | 명 | 149 | 1
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | IOL repositioning |  |  |  |  | serious ocular AE | 12개월 | 명 | 149 | 0
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | Increased IOP |  |  |  |  | serious ocular AE | 12개월 | 명 | 149 | 1
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | Blurred vision |  |  |  |  | serious ocular AE | 12개월 | 명 | 149 | 1
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | visual field defect |  |  |  |  | serious ocular AE | 12개월 | 명 | 149 | 1
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | device dislocation |  |  |  |  | serious ocular AE | 12개월 | 명 | 149 | 0
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | optic nerve cup-to-disc ratio increase |  |  |  |  | serious ocular AE | 12개월 | 명 | 149 | 1
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | vitrectomy |  |  |  |  | serious ocular AE | 12개월 | 명 | 149 | 0
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | Retinal detachment |  |  |  |  | serious ocular AE | 12개월 | 명 | 148 | 2
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | retinopexy |  |  |  |  | serious ocular AE | 12개월 | 명 | 148 | 1
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | eye operation |  |  |  |  | serious ocular AE | 12개월 | 명 | 148 | 0
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | IOL extraction |  |  |  |  | serious ocular AE | 12개월 | 명 | 148 | 0
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | IOL repositioning |  |  |  |  | serious ocular AE | 12개월 | 명 | 148 | 1
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | Increased IOP |  |  |  |  | serious ocular AE | 12개월 | 명 | 148 | 0
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | Blurred vision |  |  |  |  | serious ocular AE | 12개월 | 명 | 148 | 0
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | visual field defect |  |  |  |  | serious ocular AE | 12개월 | 명 | 148 | 1
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | device dislocation |  |  |  |  | serious ocular AE | 12개월 | 명 | 148 | 1
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | optic nerve cup-to-disc ratio increase |  |  |  |  | serious ocular AE | 12개월 | 명 | 148 | 1
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | vitrectomy |  |  |  |  | serious ocular AE | 12개월 | 명 | 148 | 1
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | dry eye |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 149 | 15
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | posterior capsule opacification |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 149 | 6
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | Foreign body sensation |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 149 | 5
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | corneal edema |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 149 | 4
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | halo vision |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 149 | 4
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | blepharitis |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 149 | 3
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | glare |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 149 | 3
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | punctate keratitis |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 149 | 3
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | vitreous detachment |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 149 | 3
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | Increased IOP |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 149 | 2
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | halo vision |  |  |  |  | device-related ocular AE | 12개월 | 명 | 149 | 4
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | posterior capsule opacification |  |  |  |  | device-related ocular AE | 12개월 | 명 | 149 | 2
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | glare |  |  |  |  | device-related ocular AE | 12개월 | 명 | 149 | 2
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | first implanted eye | Complication | visual impairment |  |  |  |  | device-related ocular AE | 12개월 | 명 | 149 | 2
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | dry eye |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 148 | 13
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | posterior capsule opacification |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 148 | 8
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | Foreign body sensation |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 148 | 3
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | corneal edema |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 148 | 4
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | halo vision |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 148 | 4
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | blepharitis |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 148 | 3
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | glare |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 148 | 3
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | punctate keratitis |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 148 | 3
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | vitreous detachment |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 148 | 2
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | Increased IOP |  |  |  |  | Non-serious ocular AE | 12개월 | 명 | 148 | 3
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | halo vision |  |  |  |  | device-related ocular AE | 12개월 | 명 | 148 | 3
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | posterior capsule opacification |  |  |  |  | device-related ocular AE | 12개월 | 명 | 148 | 3
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | glare |  |  |  |  | device-related ocular AE | 12개월 | 명 | 148 | 2
442 | Kohnen (2020) | Prospective, single-arm, nonmasked, nonrandomized study | Cataract surgery | Diffractive trifocal IOL | AcrySof IQ PanOptix(model TFNT00) |  |  | second implanted eye | Complication | visual impairment |  |  |  |  | device-related ocular AE | 12개월 | 명 | 148 | 1
469 | Pedrotti (2020) | Prospective observational comparative clinical study | Presbyopia | Diffractive EDOF | Tecnis Symfony ZXR00 | Diffractive trifocal IOL | Acrysof IQ PanOptix TNFT00 |  | Complication | [6 empty] | 3개월 | 명 | 25 | 0 | 25 | 0
1409 | Altaie (2012) | Prospective, observational study | Cataract or CLE | Diffractive bifocal IOL | AcrySof Natural ReSTOR IOL (SN60D3) |  |  |  | VA | UNVA |  |  |  |  | 6/12 or better | 6개월 | 안 | 363 | 351
1409 | Altaie (2012) | Prospective, observational study | Cataract or CLE | Diffractive bifocal IOL | AcrySof Natural ReSTOR IOL (SN60D3) |  |  |  | VA | UNVA |  |  |  |  | achieved 6/6 | 6개월 | 안 | 363 | 188.397
1409 | Altaie (2012) | Prospective, observational study | Cataract or CLE | Diffractive bifocal IOL | AcrySof Natural ReSTOR IOL (SN60D3) |  |  | Cataract group | VA |  |  |  | Snellen |  | gain at least 1 line | 6개월 | 안 | 199 | 157.011
1409 | Altaie (2012) | Prospective, observational study | Cataract or CLE | Diffractive bifocal IOL | AcrySof Natural ReSTOR IOL (SN60D3) |  |  | CLE | VA |  |  |  | Snellen |  | gain at least 1 line | 6개월 | 안 | 164 | 116.932
1409 | Altaie (2012) | Prospective, observational study | Cataract or CLE | Diffractive bifocal IOL | AcrySof Natural ReSTOR IOL (SN60D3) |  |  |  | VA | [5 empty] | lost more than one line | 6개월 | 안 | 363 | 0
1409 | Altaie (2012) | Prospective, observational study | Cataract or CLE | Diffractive bifocal IOL | AcrySof Natural ReSTOR IOL (SN60D3) |  |  |  | VA | [5 empty] | unaided near vision of N5 | 6개월 | 안 | 363 | 342.672
1409 | Altaie (2012) | Prospective, observational study | Cataract or CLE | Diffractive bifocal IOL | AcrySof Natural ReSTOR IOL (SN60D3) |  |  |  | VA | [5 empty] | unaided near vision of N8 | 6개월 | 안 | 363 | 357.555
1409 | Altaie (2012) | Prospective, observational study | Cataract or CLE | Diffractive bifocal IOL | AcrySof Natural ReSTOR IOL (SN60D3) |  |  |  | VA | BCVA |  |  |  |  | achieved 6/6 | 6개월 | 안 | 363 | 363
1409 | Altaie (2012) | Prospective, observational study | Cataract or CLE | Diffractive bifocal IOL | AcrySof Natural ReSTOR IOL (SN60D3) |  |  |  | Spectacle independence |  |  |  | Y/N |  |  | 6개월 | 명 | 205 | 168
1409 | Altaie (2012) | Prospective, observational study | Cataract or CLE | Diffractive bifocal IOL | AcrySof Natural ReSTOR IOL (SN60D3) |  |  |  | Satisfaction | Choose same IOL again |  |  | Y/N |  |  | 6개월 | 명 | 205 | 188.6
1409 | Altaie (2012) | Prospective, observational study | Cataract or CLE | Diffractive bifocal IOL | AcrySof Natural ReSTOR IOL (SN60D3) |  |  |  | Complication | posterior capsular opacification |  |  | Y/N |  |  | 6개월 | 명 | 205 | 0
1409 | Altaie (2012) | Prospective, observational study | Cataract or CLE | Diffractive bifocal IOL | AcrySof Natural ReSTOR IOL (SN60D3) |  |  |  | Complication | halo |  |  | Y/N |  |  | 6개월 | 명 | 205 | 21
1409 | Altaie (2012) | Prospective, observational study | Cataract or CLE | Diffractive bifocal IOL | AcrySof Natural ReSTOR IOL (SN60D3) |  |  |  | Complication | symptom visually disabling or required IOL exchange |  |  | Y/N |  |  | 6개월 | 명 | 205 | 0
1409 | Altaie (2012) | Prospective, observational study | Cataract or CLE | Diffractive bifocal IOL | AcrySof Natural ReSTOR IOL (SN60D3) |  |  |  | Satisfaction | LASIK after IOL implantation | [5 empty] | 6개월 | 안 | 363 | 16
1409 | Altaie (2012) | Prospective, observational study | Cataract or CLE | Diffractive bifocal IOL | AcrySof Natural ReSTOR IOL (SN60D3) |  |  | CLE | Satisfaction | LRI procedure perioperatively and LASIK postoperatively | [5 empty] | 6개월 | 안 | 164 | 4
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  |  | Complication | [6 empty] | 6개월 | 명 | 33 | 0
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | BCVA | best-corrected |  |  |  | lost >2 line | 수술전 | 안 | 21 | 0
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | BCVA | best-corrected |  |  |  | lost 2 line | 수술전 | 안 | 21 | 0
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | BCVA | best-corrected |  |  |  | lost 1 line | 수술전 | 안 | 21 | 1.995
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | BCVA | best-corrected |  |  |  | 변화없음 | 수술전 | 안 | 21 | 9.009
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | BCVA | best-corrected |  | Snellen |  | gain 1 line | 수술전 | 안 | 21 | 4.011
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | BCVA | best-corrected |  | Snellen |  | gain 2 line | 수술전 | 안 | 21 | 3.003
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | BCVA | best-corrected |  | Snellen |  | gain >2 line | 수술전 | 안 | 21 | 3.003
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | BCVA | best-corrected |  |  |  | lost >2 line | 수술전 | 안 | 45 | 0
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | BCVA | best-corrected |  |  |  | lost 2 line | 수술전 | 안 | 45 | 0
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | BCVA | best-corrected |  |  |  | lost 1 line | 수술전 | 안 | 45 | 10.98
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | BCVA | best-corrected |  |  |  | 변화없음 | 수술전 | 안 | 45 | 18.99
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | BCVA | best-corrected |  | Snellen |  | gain 1 line | 수술전 | 안 | 45 | 8.01
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | BCVA | best-corrected |  | Snellen |  | gain 2 line | 수술전 | 안 | 45 | 7.02
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | BCVA | best-corrected |  | Snellen |  | gain >2 line | 수술전 | 안 | 45 | 0
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | lost >2 line | 6개월 | 안 | 45 | 0
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | lost 2 line | 6개월 | 안 | 45 | 0
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | lost 1 line | 6개월 | 안 | 45 | 0
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | 변화없음 | 6개월 | 안 | 45 | 35.01
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | gain 1 line | 6개월 | 안 | 45 | 7
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | gain 2 line | 6개월 | 안 | 45 | 3
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | hyperopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | gain >2 line | 6개월 | 안 | 45 | 0
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | lost >2 line | 6개월 | 안 | 21 | 0
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | lost 2 line | 6개월 | 안 | 21 | 0
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | lost 1 line | 6개월 | 안 | 21 | 0
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | 변화없음 | 6개월 | 안 | 21 | 16
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | gain 1 line | 6개월 | 안 | 21 | 4
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | gain 2 line | 6개월 | 안 | 21 | 1
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | gain >2 line | 6개월 | 안 | 21 | 0
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  |  | Complication | secondary intervention postoperatively |  |  |  | Monocular |  | 6개월 | 명 | 33 | 0
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  |  | Complication | sight-threating complication |  |  |  | Monocular |  | 6개월 | 명 | 33 | 0
1485 | Alfonso (2011) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.LISA 366D |  |  |  | Complication | need of Nd:YAG capsulotomy |  |  |  | Monocular |  | 6개월 | 명 | 33 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  |  | Complication | secondary intervention postoperatively |  |  |  | Monocular |  | 6개월 | 명 | 34 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  |  | Complication | sight-threating complication |  |  |  | Monocular |  | 6개월 | 명 | 34 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  |  | Complication | need of Nd:YAG capsulotomy |  |  |  | Monocular |  | 6개월 | 명 | 34 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | BCVA | best-corrected | Distant(6m) | Snellen |  | lost >2 line | 6개월 | 안 | 14 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | BCVA | best-corrected | Distant(6m) | Snellen |  | lost 2 line | 6개월 | 안 | 14 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | BCVA | best-corrected | Distant(6m) | Snellen |  | lost 1 line | 6개월 | 안 | 14 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | BCVA | best-corrected | Distant(6m) | Snellen |  | 변화없음 | 6개월 | 안 | 14 | 4
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | BCVA | best-corrected | Distant(6m) | Snellen |  | gain 1 line | 6개월 | 안 | 14 | 4
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | BCVA | best-corrected | Distant(6m) | Snellen |  | gain 2 line | 6개월 | 안 | 14 | 2
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | BCVA | best-corrected | Distant(6m) | Snellen |  | gain >2 line | 6개월 | 안 | 14 | 4
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | BCVA | best-corrected | Distant(6m) | Snellen |  | lost >2 line | 6개월 | 안 | 20 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | BCVA | best-corrected | Distant(6m) | Snellen |  | lost 2 line | 6개월 | 안 | 20 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | BCVA | best-corrected | Distant(6m) | Snellen |  | lost 1 line | 6개월 | 안 | 20 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | BCVA | best-corrected | Distant(6m) | Snellen |  | 변화없음 | 6개월 | 안 | 20 | 12
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | BCVA | best-corrected | Distant(6m) | Snellen |  | gain 1 line | 6개월 | 안 | 20 | 3
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | BCVA | best-corrected | Distant(6m) | Snellen |  | gain 2 line | 6개월 | 안 | 20 | 5
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | BCVA | best-corrected | Distant(6m) | Snellen |  | gain >2 line | 6개월 | 안 | 20 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | lost >2 line | 6개월 | 안 | 14 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | lost 2 line | 6개월 | 안 | 14 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | lost 1 line | 6개월 | 안 | 14 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | 변화없음 | 6개월 | 안 | 14 | 11
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | gain 1 line | 6개월 | 안 | 14 | 3
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | gain 2 line | 6개월 | 안 | 14 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | myopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | gain >2 line | 6개월 | 안 | 14 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | lost >2 line | 6개월 | 안 | 20 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | lost 2 line | 6개월 | 안 | 20 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | lost 1 line | 6개월 | 안 | 20 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | 변화없음 | 6개월 | 안 | 20 | 18
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | gain 1 line | 6개월 | 안 | 20 | 2
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | gain 2 line | 6개월 | 안 | 20 | 0
1544 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | 447D IOL |  |  | hyperopic | VA | BCNVA | best-corrected | Near(33cm) | Snellen |  | gain >2 line | 6개월 | 안 | 20 | 0
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | BCVA | best spectacle corrected | Distant(6m) | Snellen |  | lost >2 line | 6개월 | 명 | 26 | 0
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | BCVA | best spectacle corrected | Distant(6m) | Snellen |  | lost 2 line | 6개월 | 명 | 26 | 1
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | BCVA | best spectacle corrected | Distant(6m) | Snellen |  | lost 1 line | 6개월 | 명 | 26 | 1
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | BCVA | best spectacle corrected | Distant(6m) | Snellen |  | 변화없음 | 6개월 | 명 | 26 | 8
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | BCVA | best spectacle corrected | Distant(6m) | Snellen |  | gain 1 line | 6개월 | 명 | 26 | 6
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | BCVA | best spectacle corrected | Distant(6m) | Snellen |  | gain 2 line | 6개월 | 명 | 26 | 5
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | BCVA | best spectacle corrected | Distant(6m) | Snellen |  | gain >2 line | 6개월 | 명 | 26 | 5
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | hyperopic | VA | BCVA | best spectacle corrected | Distant(6m) | Snellen |  | lost >2 line | 6개월 | 명 | 39 | 0
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | hyperopic | VA | BCVA | best spectacle corrected | Distant(6m) | Snellen |  | lost 2 line | 6개월 | 명 | 39 | 1
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | hyperopic | VA | BCVA | best spectacle corrected | Distant(6m) | Snellen |  | lost 1 line | 6개월 | 명 | 39 | 2
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | hyperopic | VA | BCVA | best spectacle corrected | Distant(6m) | Snellen |  | 변화없음 | 6개월 | 명 | 39 | 25
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | hyperopic | VA | BCVA | best spectacle corrected | Distant(6m) | Snellen |  | gain 1 line | 6개월 | 명 | 39 | 5
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | hyperopic | VA | BCVA | best spectacle corrected | Distant(6m) | Snellen |  | gain 2 line | 6개월 | 명 | 39 | 5
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | hyperopic | VA | BCVA | best spectacle corrected | Distant(6m) | Snellen |  | gain >2 line | 6개월 | 명 | 39 | 1
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | BCNVA | best spectacle corrected | Near(33cm) | Snellen |  | lost >2 line | 6개월 | 명 | 26 | 0
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | BCNVA | best spectacle corrected | Near(33cm) | Snellen |  | lost 2 line | 6개월 | 명 | 26 | 0
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | BCNVA | best spectacle corrected | Near(33cm) | Snellen |  | lost 1 line | 6개월 | 명 | 26 | 2
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | BCNVA | best spectacle corrected | Near(33cm) | Snellen |  | 변화없음 | 6개월 | 명 | 26 | 24
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | BCNVA | best spectacle corrected | Near(33cm) | Snellen |  | gain 1 line | 6개월 | 명 | 26 | 0
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | BCNVA | best spectacle corrected | Near(33cm) | Snellen |  | gain 2 line | 6개월 | 명 | 26 | 0
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | myopic | VA | BCNVA | best spectacle corrected | Near(33cm) | Snellen |  | gain >2 line | 6개월 | 명 | 26 | 0
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | hyperopic | VA | BCNVA | best spectacle corrected | Near(33cm) | Snellen |  | lost >2 line | 6개월 | 명 | 39 | 0
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | hyperopic | VA | BCNVA | best spectacle corrected | Near(33cm) | Snellen |  | lost 2 line | 6개월 | 명 | 39 | 0
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | hyperopic | VA | BCNVA | best spectacle corrected | Near(33cm) | Snellen |  | lost 1 line | 6개월 | 명 | 39 | 0
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | hyperopic | VA | BCNVA | best spectacle corrected | Near(33cm) | Snellen |  | 변화없음 | 6개월 | 명 | 39 | 39
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | hyperopic | VA | BCNVA | best spectacle corrected | Near(33cm) | Snellen |  | gain 1 line | 6개월 | 명 | 39 | 0
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | hyperopic | VA | BCNVA | best spectacle corrected | Near(33cm) | Snellen |  | gain 2 line | 6개월 | 명 | 39 | 0
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  | hyperopic | VA | BCNVA | best spectacle corrected | Near(33cm) | Snellen |  | gain >2 line | 6개월 | 명 | 39 | 0
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  |  | Complication | secondary intervention postoperatively |  |  |  | Monocular |  | 6개월 | 명 | 65 | 0
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  |  | Complication | sight-threating complication |  |  |  | Monocular |  | 6개월 | 명 | 65 | 0
1556 | Alfonso (2010) | Prospective study | Presbyopia | Diffractive bifocal IOL | Acri.Twin |  |  |  | Complication | need of Nd:YAG capsulotomy |  |  |  | Monocular |  | 6개월 | 명 | 65 | 0
3093 | Tan (2019) | Prospective study | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Technis Symfony IOL |  |  |  | Complication | mild corneal edema | [5 empty] | 3개월 | 안 | 55 | 1
3093 | Tan (2019) | Prospective study | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Technis Symfony IOL |  |  |  | Complication | high IOP | [5 empty] | 3개월 | 안 | 55 | 2
3093 | Tan (2019) | Prospective study | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Technis Symfony IOL |  |  |  | Complication | obvious corneal edema | [5 empty] | 3개월 | 안 | 55 | 0
3093 | Tan (2019) | Prospective study | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Technis Symfony IOL |  |  |  | Complication | uncontrolled IOP | [5 empty] | 3개월 | 안 | 55 | 0
3093 | Tan (2019) | Prospective study | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Technis Symfony IOL |  |  |  | Complication | endophthalmitis | [5 empty] | 3개월 | 안 | 55 | 0
3093 | Tan (2019) | Prospective study | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Technis Symfony IOL |  |  |  | Complication | obvious posterior capsule opacification | [5 empty] | 3개월 | 안 | 55 | 0
3093 | Tan (2019) | Prospective study | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Technis Symfony IOL |  |  |  | Complication | pupil deformation | [5 empty] | 3개월 | 안 | 55 | 0
3093 | Tan (2019) | Prospective study | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Technis Symfony IOL |  |  |  | Complication | IOL displacement or tilting or signs of capsular contraction syndrome | [5 empty] | 3개월 | 안 | 55 | 0
3093 | Tan (2019) | Prospective study | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Technis Symfony IOL |  |  |  | Spectacle dependence | wear glasses for reading or using cell phone | [5 empty] | 3개월 | 명 | 40 | 3
3093 | Tan (2019) | Prospective study | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Technis Symfony IOL |  |  |  | Spectacle dependence | wear glasses for watching TV or using the computer | [5 empty] | 3개월 | 명 | 40 | 0
3093 | Tan (2019) | Prospective study | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Technis Symfony IOL |  |  |  | Spectacle independence | [6 empty] | 3개월 | 명 | 40 | 37.82
3093 | Tan (2019) | Prospective study | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Technis Symfony IOL |  |  |  | Satisfaction | complained of halos | [5 empty] | 3개월 | 명 | 40 | 2
3093 | Tan (2019) | Prospective study | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Technis Symfony IOL |  |  |  | Satisfaction | complained of glare during night driving | [5 empty] | 3개월 | 명 | 40 | 1
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | Decimal | Monocular | 0.6 | 3개월 | 안 | 82 | 14
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | Decimal | Monocular | 0.7 | 3개월 | 안 | 82 | 1
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | Decimal | Monocular | 0.8 | 3개월 | 안 | 82 | 22
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | Decimal | Monocular | 0.85 | 3개월 | 안 | 82 | 1
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | Decimal | Monocular | 0.9 | 3개월 | 안 | 82 | 2
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | Decimal | Monocular | 1 | 3개월 | 안 | 82 | 42
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UDVA | uncorrected | Distant | Decimal | Monocular | 0.5 | 3개월 | 안 | 82 | 3
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UDVA | uncorrected | Distant | Decimal | Monocular | 0.6 | 3개월 | 안 | 82 | 17
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UDVA | uncorrected | Distant | Decimal | Monocular | 0.8 | 3개월 | 안 | 82 | 25
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UDVA | uncorrected | Distant | Decimal | Monocular | 0.85 | 3개월 | 안 | 82 | 1
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UDVA | uncorrected | Distant | Decimal | Monocular | 1 | 3개월 | 안 | 82 | 36
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UNVA | uncorrected | near(40cm) | Decimal | Monocular | 0.6 | 3개월 | 안 | 82 | 6
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UNVA | uncorrected | Intermediate(66cm) | Decimal | Monocular | 0.8 | 3개월 | 안 | 82 | 23
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UNVA | uncorrected | Intermediate(66cm) | Decimal | Monocular | 0.85 | 3개월 | 안 | 82 | 1
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | VA | UNVA | uncorrected | Intermediate(66cm) | Decimal | Monocular | 1 | 3개월 | 안 | 82 | 52
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Complication | IOL decentration | [5 empty] | 3개월 | 안 | 82 | 0
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Complication | posterior capsule opacification | [5 empty] | 3개월 | 안 | 82 | 0
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | are you satisfied with your sight at the moment? |  |  |  |  | very dissatisifed | 1개월 | 명 | 41 | 0
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | are you satisfied with your sight at the moment? |  |  |  |  | fairly dissatisfied | 1개월 | 명 | 41 | 1
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | are you satisfied with your sight at the moment? |  |  |  |  | fairly satisfied | 1개월 | 명 | 41 | 24
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | are you satisfied with your sight at the moment? |  |  |  |  | very satisfied | 1개월 | 명 | 41 | 16
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | are you satisfied with your sight at the moment? |  |  |  |  | cannot decided | 1개월 | 명 | 41 | 0
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | are you satisfied with your sight at the moment? |  |  |  |  | very dissatisifed | 3개월 | 명 | 41 | 0
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | are you satisfied with your sight at the moment? |  |  |  |  | fairly dissatisfied | 3개월 | 명 | 41 | 0
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | are you satisfied with your sight at the moment? |  |  |  |  | fairly satisfied | 3개월 | 명 | 41 | 4
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | are you satisfied with your sight at the moment? |  |  |  |  | very satisfied | 3개월 | 명 | 41 | 37
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | are you satisfied with your sight at the moment? |  |  |  |  | cannot decided | 3개월 | 명 | 41 | 0
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | willingness to undergo surgery again |  |  | Y/N |  | Yes | 1개월 | 명 | 41 | 39
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | willingness to undergo surgery again |  |  | Y/N |  | Yes | 3개월 | 명 | 41 | 41
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | willingness to recommend surgery to other |  |  | Y/N |  | Yes | 1개월 | 명 | 41 | 40
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | willingness to recommend surgery to other |  |  | Y/N |  | Yes | 3개월 | 명 | 41 | 41
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Spectacle dependence |  |  |  | Y/N |  | No | 1개월 | 명 | 41 | 35
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Spectacle dependence |  |  |  | Y/N |  | No | 3개월 | 명 | 41 | 38
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Spectacle dependence | distant vision |  |  | Y/N |  | Yes | 1개월 | 명 | 41 | 4
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Spectacle dependence | distant vision |  |  | Y/N |  | Yes | 3개월 | 명 | 41 | 2
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Spectacle dependence | intermediate vision |  |  | Y/N |  | Yes | 1개월 | 명 | 41 | 2
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Spectacle dependence | intermediate vision |  |  | Y/N |  | Yes | 3개월 | 명 | 41 | 1
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | Glare |  |  | Y/N |  | Yes | 1개월 | 명 | 41 | 9
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | Glare |  |  | Y/N |  | Yes | 3개월 | 명 | 41 | 2
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | halo |  |  | Y/N |  | Yes | 1개월 | 명 | 41 | 27
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | halo |  |  | Y/N |  | Yes | 3개월 | 명 | 41 | 14
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | starburst |  |  | Y/N |  | Yes | 1개월 | 명 | 41 | 1
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | starburst |  |  | Y/N |  | Yes | 3개월 | 명 | 41 | 1
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | Blurred vision |  |  | Y/N |  | Yes | 1개월 | 명 | 41 | 2
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | Blurred vision |  |  | Y/N |  | Yes | 3개월 | 명 | 41 | 1
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | Do you find these symptoms disturbing |  |  | Y/N |  | Yes | 1개월 | 명 | 39 | 2
3220 | Almulhim (2018) | Prospective nonrandomized study | Cataract surgery, Presbyopia | Diffractive trifocal IOL | AT LISA tri 839MP, AT LISA toric 939 MP |  |  |  | Satisfaction | Do you find these symptoms disturbing |  |  | Y/N |  | Yes | 3개월 | 명 | 18 | 0
3241 | Zamora-De-La-Cruz (2018) | Observational longitudinal and prosective study | Cataract surgery, Presbyopia | Diffractive bifocal IOL | ReSTOR +2.50 | Diffractive trifocal IOL | Acrysof IQ PanOptix |  | Complication | posteior capsule opacification |  |  | Y/N |  |  |  | 명 | 12 | 0
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Monocular | -0.1 | 수술전 | 안 | 89 | 33.019
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Monocular | 0 | 수술전 | 안 | 89 | 59.007
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Monocular | 0.1 | 수술전 | 안 | 89 | 79.032
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Monocular | 0.2 | 수술전 | 안 | 89 | 87.042
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | CDVA | corrected | Distant | LogMAR | Monocular | 0.3 | 수술전 | 안 | 89 | 87.042
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | -0.1 | 3개월 | 안 | 89 | 21.983
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | 0 | 3개월 | 안 | 89 | 50.018
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | 0.1 | 3개월 | 안 | 89 | 79.032
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | 0.2 | 3개월 | 안 | 89 | 87.042
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | 0.3 | 3개월 | 안 | 89 | 89
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0 | 3개월 | 안 | 89 | 20.025
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 89 | 39.961
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 89 | 66.038
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 89 | 76.006
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 89 | 88.021
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Binocular | 0 | 3개월 | 안 | 89 | 29.014
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Binocular | 0.1 | 3개월 | 안 | 89 | 51.976
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Binocular | 0.2 | 3개월 | 안 | 89 | 73.959
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Binocular | 0.3 | 3개월 | 안 | 89 | 84.995
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Binocular | 0.4 | 3개월 | 안 | 89 | 89
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | VA | CDVA | corrected | Distant | Snellen | Monocular | lost more than 2line | 3개월 | 안 | 89 | 0
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | Complication | late iritis | [5 empty] | 3개월 | 명 | 58 | 2
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | Complication | posterior capsule opacification that requried neodynamic: YAG laser capsulotomy | [5 empty] | 3개월 | 명 | 58 | 0
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | Complication | further surgery (lasik or astigmatic keratotomy) to correct remaining refractive error | [5 empty] | 3개월 | 명 | 58 | 4
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | Satisfaction | starbust and halo |  |  |  |  | a lot of difficulty | 3개월 | 명 | 58 | 4.002
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | Satisfaction | starbust and halo |  |  |  |  | moderate difficulty | 3개월 | 명 | 58 | 8.004
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | Satisfaction | glare |  |  |  |  | a lot of difficulty | 3개월 | 명 | 58 | 3.016
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | Satisfaction | Glare |  |  |  |  | moderate difficulty | 3개월 | 명 | 58 | 8.99
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | Satisfaction | Driving at night |  |  |  |  | a lot of difficulty | 3개월 | 명 | 58 | 3.016
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | Satisfaction | Driving at night |  |  |  |  | moderate difficulty | 3개월 | 명 | 58 | 5.974
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | Satisfaction | Ghosting/Doubling |  |  |  |  | a lot of difficulty | 3개월 | 명 | 58 | 1.972
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | Satisfaction | Ghosting/Doubling |  |  |  |  | moderate difficulty | 3개월 | 명 | 58 | 4.002
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | Satisfaction | Reading small print(medicine bottles, telephone book, newspaper) |  |  |  |  | a lot of difficulty | 3개월 | 명 | 58 | 4.002
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | Satisfaction | Reading small print(medicine bottles, telephone book, newspaper) |  |  |  |  | moderate difficulty | 3개월 | 명 | 58 | 7.018
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | Satisfaction | intermediate vision(computer screen, dashboard) |  |  |  |  | a lot of difficulty | 3개월 | 명 | 58 | 1.972
4009 | Venter (2013) | Prospective Case series study | ametropic Presbyopia | Refractive bifocal toric IOL | Lentis Mplus toric LU-313 MFT |  |  |  | Satisfaction | intermediate vision(computer screen, dashboard) |  |  |  |  | moderate difficulty | 3개월 | 명 | 58 | 4.002
4240 | Van Der Linden (2012) | Comparative case series | Presbyopia | Refractive bifocal IOL | Lentis Mplus LS-312 | Diffractive bifocal IOL | ReSTOR SN6AD1 |  | Spectacle dependence | need reading glasses under good lighting conditions | [5 empty] | 3개월 | 명 | 45 | 0 | 72 | 0
4240 | Van Der Linden (2012) | Comparative case series | Presbyopia | Refractive bifocal IOL | Lentis Mplus LS-312 | Diffractive bifocal IOL | ReSTOR SN6AD1 |  | Satisfaction | satisfied with their postoperative vision | [5 empty] | 3개월 | 명 | 45 | 38 | 72 | 70.992
4240 | Van Der Linden (2012) | Comparative case series | Presbyopia | Refractive bifocal IOL | Lentis Mplus LS-312 | Diffractive bifocal IOL | ReSTOR SN6AD1 |  | Complication | IOL decentration greater than 1.0mm | [5 empty] | 3개월 | 명 | 45 | 0 | 72 | 0
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 50 | 0
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 50 | 1
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 50 | 14
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular | 0 | 3개월 | 안 | 50 | 34
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 50 | 43
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 50 | 47
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular | 3 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular | 0.5 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular | 0.6 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UDVA | uncorrected | Distant(4m) | LogMAR | Monocular | 0.7 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCVA | distance-corrected | Distant(4m) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 50 | 0
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCVA | distance-corrected | Distant(4m) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 50 | 6
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCVA | distance-corrected | Distant(4m) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 50 | 25
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCVA | distance-corrected | Distant(4m) | LogMAR | Monocular | 0 | 3개월 | 안 | 50 | 42
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCVA | distance-corrected | Distant(4m) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 50 | 49
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCVA | distance-corrected | Distant(4m) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCVA | distance-corrected | Distant(4m) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCVA | distance-corrected | Distant(4m) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCVA | distance-corrected | Distant(4m) | LogMAR | Monocular | 0.5 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCVA | distance-corrected | Distant(4m) | LogMAR | Monocular | 0.6 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCVA | distance-corrected | Distant(4m) | LogMAR | Monocular | 0.7 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | intermediate(80cm) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 50 | 0
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | intermediate(80cm) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 50 | 2
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | intermediate(80cm) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 50 | 8
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | intermediate(80cm) | LogMAR | Monocular | 0 | 3개월 | 안 | 50 | 23
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | intermediate(80cm) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 50 | 36
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | intermediate(80cm) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 50 | 45
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | intermediate(80cm) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 50 | 47
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | intermediate(80cm) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | intermediate(80cm) | LogMAR | Monocular | 0.5 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | intermediate(80cm) | LogMAR | Monocular | 0.6 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | intermediate(80cm) | LogMAR | Monocular | 0.7 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 50 | 0
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 50 | 0
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 50 | 3
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0 | 3개월 | 안 | 50 | 18
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 50 | 33
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 50 | 45
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 50 | 46
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.5 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.6 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.7 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 50 | 0
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 50 | 3
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Monocular | 0 | 3개월 | 안 | 50 | 14.5
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Monocular | 0.5 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 50 | 0
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 50 | 4
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0 | 3개월 | 안 | 50 | 19
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.5 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 50 | 0
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 50 | 0
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 50 | 5
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0 | 3개월 | 안 | 50 | 16
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 50 | 36
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 50 | 45
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0.5 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0.6 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0.7 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 50 | 0
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 50 | 0
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 50 | 11
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0 | 3개월 | 안 | 50 | 25
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 50 | 40
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 50 | 43
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.5 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.6 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.7 | 3개월 | 안 | 50 | 50
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | Spectacle independence |  |  |  | Y/N |  |  | 3개월 | 명 | 25 | 20
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | Satisfaction | recommend the IOL to others |  |  | Y/N |  |  | 3개월 | 명 | 25 | 22
471 | Kohnen (2020) | Prospective single-arm study | Cataract surgery, Presbyopia | Diffractive trifocal toric IOL | AcrySof IQ PanOptix TFNT20-60 |  |  |  | Complication | adverse or serious adverse event |  |  | Y/N |  |  | 3개월 | 명 | 25 | 0
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with the overall result of the procedure? |  |  |  |  | dissatified | 12개월 | 명 | 10 | 1
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with the overall result of the procedure? |  |  |  |  | a little | 12개월 | 명 | 10 | 0
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with the overall result of the procedure? |  |  |  |  | quite satisfied | 12개월 | 명 | 10 | 1
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with the overall result of the procedure? |  |  |  |  | very sastified | 12개월 | 명 | 10 | 8
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with your uncorrected vision? |  |  |  |  | No | 12개월 | 명 | 10 | 1
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with your uncorrected vision? |  |  |  |  | a little | 12개월 | 명 | 10 | 0
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with your uncorrected vision? |  |  |  |  | considerably | 12개월 | 명 | 10 | 4
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with your uncorrected vision? |  |  |  |  | a lot | 12개월 | 명 | 10 | 5
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with your near vision? |  |  |  |  | No | 12개월 | 명 | 10 | 1
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with your near vision? |  |  |  |  | a little | 12개월 | 명 | 10 | 2
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with your near vision? |  |  |  |  | considerably | 12개월 | 명 | 10 | 1
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with your near vision? |  |  |  |  | a lot | 12개월 | 명 | 10 | 6
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with your intermediate vision? |  |  |  |  | No | 12개월 | 명 | 10 | 1
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with your intermediate vision? |  |  |  |  | A little | 12개월 | 명 | 10 | 2
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with your intermediate vision? |  |  |  |  | Considerably | 12개월 | 명 | 10 | 5
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with your intermediate vision? |  |  |  |  | A lot | 12개월 | 명 | 10 | 4
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with your far vision? |  |  |  |  | No | 12개월 | 명 | 10 | 0
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with your far vision? |  |  |  |  | A little | 12개월 | 명 | 10 | 1
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with your far vision? |  |  |  |  | Considerably | 12개월 | 명 | 10 | 4
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | How satisfied are you with your far vision? |  |  |  |  | A lot | 12개월 | 명 | 10 | 5
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | Did you get the visual quality you wanted? |  |  |  |  | No | 12개월 | 명 | 10 | 1
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | Did you get the visual quality you wanted? |  |  |  |  | Maybe | 12개월 | 명 | 10 | 0
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | Did you get the visual quality you wanted? |  |  |  |  | Considerably | 12개월 | 명 | 10 | 1
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | Did you get the visual quality you wanted? |  |  |  |  | Yes | 12개월 | 명 | 10 | 8
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | Has this lens been a good choice? |  |  |  |  | No | 12개월 | 명 | 10 | 1
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | Has this lens been a good choice? |  |  |  |  | A little | 12개월 | 명 | 10 | 0
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | Has this lens been a good choice? |  |  |  |  | Considerably | 12개월 | 명 | 10 | 0
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | Has this lens been a good choice? |  |  |  |  | A lot | 12개월 | 명 | 10 | 5
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | With your current experience, would you operate again? |  |  |  |  | No | 12개월 | 명 | 10 | 1
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | With your current experience, would you operate again? |  |  |  |  | Maybe | 12개월 | 명 | 10 | 0
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | With your current experience, would you operate again? |  |  |  |  | Considerably | 12개월 | 명 | 10 | 0
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | With your current experience, would you operate again? |  |  |  |  | Yes | 12개월 | 명 | 10 | 9
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | Would you recommend the implantation of this lens to another person? |  |  |  |  | No | 12개월 | 명 | 10 | 1
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | Would you recommend the implantation of this lens to another person? |  |  |  |  | Maybe | 12개월 | 명 | 10 | 0
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | Would you recommend the implantation of this lens to another person? |  |  |  |  | Considerably | 12개월 | 명 | 10 | 0
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Satisfaction | Would you recommend the implantation of this lens to another person? |  |  |  |  | Yes | 12개월 | 명 | 10 | 9
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Spectacle independence | How often do you still wear your glasses or contact lenses? |  |  |  |  | Never | 12개월 | 명 | 10 | 6
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Spectacle independence | How often do you still wear your glasses or contact lenses? |  |  |  |  | Occasionally | 12개월 | 명 | 10 | 2
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Spectacle independence | How often do you still wear your glasses or contact lenses? |  |  |  |  | Often | 12개월 | 명 | 10 | 2
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Spectacle independence | How often do you still wear your glasses or contact lenses? |  |  |  |  | Most often | 12개월 | 명 | 10 | 0
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Spectacle independence | If you wear glasses, for what distance? |  |  |  |  | Not available | 12개월 | 명 | 4 | 2.4
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Spectacle independence | If you wear glasses, for what distance? |  |  |  |  | Near vision | 12개월 | 명 | 4 | 1.6
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Spectacle independence | If you wear glasses, for what distance? |  |  |  |  | Intermediate vision | 12개월 | 명 | 4 | 0
488 | Alio (2020) | Prospective consecutive case series | Cataract surgery, Presbyopia | Refractive multifocal IOL | Precizon Presbyopic IOL NVA model 570 (OPTHEC BV) |  |  |  | Spectacle independence | If you wear glasses, for what distance? |  |  |  |  | Far vision | 12개월 | 명 | 4 | 0
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | >1D(high preopoerative corneal astigmatism) | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | 0-0.1 | 4개월 | 안 | 14 | 5.6
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | >1D(high preopoerative corneal astigmatism) | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | 0.2-0.3 | 4개월 | 안 | 14 | 5.6
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | >1D(high preopoerative corneal astigmatism) | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | 0.4+ | 4개월 | 안 | 14 | 2.8
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | 0-0.99D(low preoperative corneal astigmatism) | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0-0.1 | 4개월 | 안 | 34 | 26.86
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | 0-0.99D(low preoperative corneal astigmatism) | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0.2-0.3 | 4개월 | 안 | 34 | 6.46
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | 0-0.99D(low preoperative corneal astigmatism) | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0.4+ | 4개월 | 안 | 34 | 0.68
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | <1D(UNVA low group) | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0.2 | 4개월 | 안 | 80 | 7.2
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | <1D(UNVA low group) | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0.3 | 4개월 | 안 | 80 | 48.8
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | <1D(UNVA low group) | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0.4 | 4개월 | 안 | 80 | 16
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | <1D(UNVA low group) | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | 0.5 | 4개월 | 안 | 80 | 4.8
513 | Power (2018) | Prospective cohort analysis | Cataract surgery, Presbyopia | Diffractive EDOF IOL | Symfony |  |  | <1D(UNVA low group) | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | >0.6 | 4개월 | 안 | 80 | 3.2
621 | Friedrich (2012) | Prospective consecutive study | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Complication | [6 empty] | 6개월 | 명 | 19 | 0
621 | Friedrich (2012) | Prospective consecutive study | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | BCVA | best-corrected | Distant | LogMAR | Monocular | 0.1 or better | 6개월 | 안 | 31 | 30
621 | Friedrich (2012) | Prospective consecutive study | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UCVA | uncorrected | Distant | LogMAR | Monocular | 0.1 or better | 6개월 | 안 | 31 | 24
621 | Friedrich (2012) | Prospective consecutive study | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UCVA | uncorrected | Distant | LogMAR | Binocular | 0.1 or better | 6개월 | 명 | 19 | 18
621 | Friedrich (2012) | Prospective consecutive study | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UCVA | uncorrected | near | LogMAR | Monocular | 0.0 or better | 6개월 | 안 | 31 | 21
621 | Friedrich (2012) | Prospective consecutive study | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UCVA | uncorrected | near | LogMAR | Monocular | 0.1 or better | 6개월 | 안 | 31 | 29
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | BCDVA | best-corrected | Distant(4m) | Snellen | Binocular | 20/20 or better | 수술전 | 명 | 44 | 21
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | BCDVA | best-corrected | Distant(4m) | Snellen | Binocular | 20/25 or better | 수술전 | 명 | 44 | 26
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | BCDVA | best-corrected | Distant(4m) | Snellen | Binocular | 20/32 or better | 수술전 | 명 | 44 | 31
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | BCDVA | best-corrected | Distant(4m) | Snellen | Binocular | 20/24 or better | 수술전 | 명 | 44 | 40
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | BCDVA | best-corrected | Distant(4m) | Snellen | Binocular | 20/20 or better | 3개월 | 명 | 44 | 36
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | BCDVA | best-corrected | Distant(4m) | Snellen | Binocular | 20/25 or better | 3개월 | 명 | 44 | 43
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | BCDVA | best-corrected | Distant(4m) | Snellen | Binocular | 20/32 or better | 3개월 | 명 | 44 | 44
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | BCDVA | best-corrected | Distant(4m) | Snellen | Binocular | 20/24 or better | 3개월 | 명 | 44 | 44
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | UCDVA | uncorrected | Distant(4m) | Snellen | Binocular | 20/20 or better | 3개월 | 명 | 44 | 23
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | UCDVA | uncorrected | Distant(4m) | Snellen | Binocular | 20/25 or better | 3개월 | 명 | 44 | 35
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | UCDVA | uncorrected | Distant(4m) | Snellen | Binocular | 20/32 or better | 3개월 | 명 | 44 | 42
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | UCDVA | uncorrected | Distant(4m) | Snellen | Binocular | 20/24 or better | 3개월 | 명 | 44 | 44
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | UCIVA | uncorrected | Intermediate(60cm) | Snellen | Binocular | 20/20 or better | 3개월 | 명 | 44 | 29
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | UCIVA | uncorrected | Intermediate(60cm) | Snellen | Binocular | 20/25 or better | 3개월 | 명 | 44 | 40
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | UCIVA | uncorrected | Intermediate(60cm) | Snellen | Binocular | 20/32 or better | 3개월 | 명 | 44 | 42
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | UCIVA | uncorrected | Intermediate(60cm) | Snellen | Binocular | 20/24 or better | 3개월 | 명 | 44 | 44
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | UCNVA | uncorrected | near(40cm) | Snellen | Binocular | 20/20 or better | 3개월 | 명 | 44 | 25
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | UCNVA | uncorrected | near(40cm) | Snellen | Binocular | 20/25 or better | 3개월 | 명 | 44 | 35
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | UCNVA | uncorrected | near(40cm) | Snellen | Binocular | 20/32 or better | 3개월 | 명 | 44 | 41
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | VA | UCNVA | uncorrected | near(40cm) | Snellen | Binocular | 20/24 or better | 3개월 | 명 | 44 | 42
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your vision for seeing objects at near distance? |  |  |  | Binocular | Dissatisfied/very dissatisfied | 수술전 | 명 | 44 | 39
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your vision for seeing objects at near distance? |  |  |  | Binocular | Neither satisfied nor dissatisfied | 수술전 | 명 | 44 | 3
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your vision for seeing objects at near distance? |  |  |  | Binocular | Satisfied/very satisfied | 수술전 | 명 | 44 | 2
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at near distance? |  |  |  | Binocular | None of the time | 수술전 | 명 | 44 | 7
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at near distance? |  |  |  | Binocular | Some of the time | 수술전 | 명 | 44 | 7
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at near distance? |  |  |  | Binocular | Most of the time | 수술전 | 명 | 44 | 11
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at near distance? |  |  |  | Binocular | All of the time | 수술전 | 명 | 44 | 19
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your vision for seeing objects at intermediate distance? |  |  |  | Binocular | Dissatisfied/very dissatisfied | 수술전 | 명 | 44 | 38
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your vision for seeing objects at intermediate distance? |  |  |  | Binocular | Neither satisfied nor dissatisfied | 수술전 | 명 | 44 | 3
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your vision for seeing objects at intermediate distance? |  |  |  | Binocular | Satisfied/very satisfied | 수술전 | 명 | 44 | 3
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at intermediated distance? |  |  |  | Binocular | None of the time | 수술전 | 명 | 44 | 9
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at intermediated distance? |  |  |  | Binocular | Some of the time | 수술전 | 명 | 44 | 12
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at intermediated distance? |  |  |  | Binocular | Most of the time | 수술전 | 명 | 44 | 11
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at intermediated distance? |  |  |  | Binocular | All of the time | 수술전 | 명 | 44 | 12
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your vision for seeing objects at distance? |  |  |  | Binocular | Dissatisfied/very dissatisfied | 수술전 | 명 | 44 | 28
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your vision for seeing objects at distance? |  |  |  | Binocular | Neither satisfied nor dissatisfied | 수술전 | 명 | 44 | 11
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your vision for seeing objects at distance? |  |  |  | Binocular | Satisfied/very satisfied | 수술전 | 명 | 44 | 5
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at distance? |  |  |  | Binocular | None of the time | 수술전 | 명 | 44 | 14
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at distance? |  |  |  | Binocular | Some of the time | 수술전 | 명 | 44 | 9
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at distance? |  |  |  | Binocular | Most of the time | 수술전 | 명 | 44 | 5
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at distance? |  |  |  | Binocular | All of the time | 수술전 | 명 | 44 | 16
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How often do you experience halos? |  |  |  | Binocular | None of the time | 수술전 | 명 | 44 | 10
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How often do you experience halos? |  |  |  | Binocular | Some of the time | 수술전 | 명 | 44 | 20
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How often do you experience halos? |  |  |  | Binocular | Most of the time | 수술전 | 명 | 44 | 10
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How often do you experience halos? |  |  |  | Binocular | All of the time | 수술전 | 명 | 44 | 4
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How severe were these halos? |  |  |  | Binocular | none | 수술전 | 명 | 44 | 9
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How severe were these halos? |  |  |  | Binocular | mild | 수술전 | 명 | 44 | 13
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How severe were these halos? |  |  |  | Binocular | moderate | 수술전 | 명 | 44 | 15
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How severe were these halos? |  |  |  | Binocular | severe | 수술전 | 명 | 44 | 7
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | If you currently drive, how much difficulty do you have driving at night? |  |  |  | Binocular | No difficulty at all | 수술전 | 명 | 44 | 5
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | If you currently drive, how much difficulty do you have driving at night? |  |  |  | Binocular | A little difficulty | 수술전 | 명 | 44 | 5
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | If you currently drive, how much difficulty do you have driving at night? |  |  |  | Binocular | Moderate difficulty | 수술전 | 명 | 44 | 13
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | If you currently drive, how much difficulty do you have driving at night? |  |  |  | Binocular | Extreme difficulty | 수술전 | 명 | 44 | 11
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | If you currently drive, how much difficulty do you have driving at night? |  |  |  | Binocular | I do not drive at night | 수술전 | 명 | 44 | 10
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | If you do not drive at night, what is the reason? |  |  |  | Binocular | Because of your current eyesight | 수술전 | 명 | 44 | 12
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | If you do not drive at night, what is the reason? |  |  |  | Binocular | Because you are not interested in driving | 수술전 | 명 | 44 | 5
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | If you do not drive at night, what is the reason? |  |  |  | Binocular | Because you have other reasons | 수술전 | 명 | 44 | 10
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | If you do not drive at night, what is the reason? |  |  |  | Binocular | I drive at night | 수술전 | 명 | 44 | 17
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your cataract surgery result? |  |  |  | Binocular | Dissatisfied/very dissatisfied | 수술전 | 명 | 44 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your cataract surgery result? |  |  |  | Binocular | Neither satisfied nor dissatisfied | 수술전 | 명 | 44 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your cataract surgery result? |  |  |  | Binocular | Satisfied/very satisfied | 수술전 | 명 | 44 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | Would you recommend the cataract surgery and the new lenses that were put into your eyes to other people? |  |  | Y/N | Binocular | Yes | 수술전 | 명 | 44 | NA
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your vision for seeing objects at near distance? |  |  |  | Binocular | Dissatisfied/very dissatisfied | 3개월 | 명 | 44 | 3
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your vision for seeing objects at near distance? |  |  |  | Binocular | Neither satisfied nor dissatisfied | 3개월 | 명 | 44 | 4
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your vision for seeing objects at near distance? |  |  |  | Binocular | Satisfied/very satisfied | 3개월 | 명 | 44 | 37
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at near distance? |  |  |  | Binocular | None of the time | 3개월 | 명 | 44 | 37
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at near distance? |  |  |  | Binocular | Some of the time | 3개월 | 명 | 44 | 6
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at near distance? |  |  |  | Binocular | Most of the time | 3개월 | 명 | 44 | 0
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at near distance? |  |  |  | Binocular | All of the time | 3개월 | 명 | 44 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your vision for seeing objects at intermediate distance? |  |  |  | Binocular | Dissatisfied/very dissatisfied | 3개월 | 명 | 44 | 5
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your vision for seeing objects at intermediate distance? |  |  |  | Binocular | Neither satisfied nor dissatisfied | 3개월 | 명 | 44 | 5
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your vision for seeing objects at intermediate distance? |  |  |  | Binocular | Satisfied/very satisfied | 3개월 | 명 | 44 | 34
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at intermediated distance? |  |  |  | Binocular | None of the time | 3개월 | 명 | 44 | 40
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at intermediated distance? |  |  |  | Binocular | Some of the time | 3개월 | 명 | 44 | 3
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at intermediated distance? |  |  |  | Binocular | Most of the time | 3개월 | 명 | 44 | 0
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at intermediated distance? |  |  |  | Binocular | All of the time | 3개월 | 명 | 44 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your vision for seeing objects at distance? |  |  |  | Binocular | Dissatisfied/very dissatisfied | 3개월 | 명 | 44 | 4
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your vision for seeing objects at distance? |  |  |  | Binocular | Neither satisfied nor dissatisfied | 3개월 | 명 | 44 | 9
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your vision for seeing objects at distance? |  |  |  | Binocular | Satisfied/very satisfied | 3개월 | 명 | 44 | 31
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at distance? |  |  |  | Binocular | None of the time | 3개월 | 명 | 44 | 42
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at distance? |  |  |  | Binocular | Some of the time | 3개월 | 명 | 44 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at distance? |  |  |  | Binocular | Most of the time | 3개월 | 명 | 44 | 0
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Spectacle independence | How often do you wear eyeglasses or contact lenses for seeing objects at distance? |  |  |  | Binocular | All of the time | 3개월 | 명 | 44 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How often do you experience halos? |  |  |  | Binocular | None of the time | 3개월 | 명 | 44 | 3
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How often do you experience halos? |  |  |  | Binocular | Some of the time | 3개월 | 명 | 44 | 11
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How often do you experience halos? |  |  |  | Binocular | Most of the time | 3개월 | 명 | 44 | 15
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How often do you experience halos? |  |  |  | Binocular | All of the time | 3개월 | 명 | 44 | 15
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How severe were these halos? |  |  |  | Binocular | none | 3개월 | 명 | 44 | 3
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How severe were these halos? |  |  |  | Binocular | mild | 3개월 | 명 | 44 | 5
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How severe were these halos? |  |  |  | Binocular | moderate | 3개월 | 명 | 44 | 22
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How severe were these halos? |  |  |  | Binocular | severe | 3개월 | 명 | 44 | 14
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | If you currently drive, how much difficulty do you have driving at night? |  |  |  | Binocular | No difficulty at all | 3개월 | 명 | 44 | 4
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | If you currently drive, how much difficulty do you have driving at night? |  |  |  | Binocular | A little difficulty | 3개월 | 명 | 44 | 5
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | If you currently drive, how much difficulty do you have driving at night? |  |  |  | Binocular | Moderate difficulty | 3개월 | 명 | 44 | 11
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | If you currently drive, how much difficulty do you have driving at night? |  |  |  | Binocular | Extreme difficulty | 3개월 | 명 | 44 | 10
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | If you currently drive, how much difficulty do you have driving at night? |  |  |  | Binocular | I do not drive at night | 3개월 | 명 | 44 | 14
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | If you do not drive at night, what is the reason? |  |  |  | Binocular | Because of your current eyesight | 3개월 | 명 | 44 | 11
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | If you do not drive at night, what is the reason? |  |  |  | Binocular | Because you are not interested in driving | 3개월 | 명 | 44 | 5
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | If you do not drive at night, what is the reason? |  |  |  | Binocular | Because you have other reasons | 3개월 | 명 | 44 | 12
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | If you do not drive at night, what is the reason? |  |  |  | Binocular | I drive at night | 3개월 | 명 | 44 | 16
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your cataract surgery result? |  |  |  | Binocular | Dissatisfied/very dissatisfied | 3개월 | 명 | 44 | 3
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your cataract surgery result? |  |  |  | Binocular | Neither satisfied nor dissatisfied | 3개월 | 명 | 44 | 8
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | How satisfied are you with your cataract surgery result? |  |  |  | Binocular | Satisfied/very satisfied | 3개월 | 명 | 44 | 33
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  |  | Satisfaction | Would you recommend the cataract surgery and the new lenses that were put into your eyes to other people? |  |  | Y/N | Binocular | Yes | 3개월 | 명 | 44 | 28
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | dry eye |  |  |  | Monocular |  | 3개월 | 안 | 45 | 11
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | Glare |  |  |  | Monocular |  | 3개월 | 안 | 45 | 10
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | Visual impairment |  |  |  | Monocular |  | 3개월 | 안 | 45 | 3
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | Halo vision |  |  |  | Monocular |  | 3개월 | 안 | 45 | 3
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | Foreign body sensation in eyes |  |  |  | Monocular |  | 3개월 | 안 | 45 | 3
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | Vitreous floaters |  |  |  | Monocular |  | 3개월 | 안 | 45 | 2
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | Posterior capsule opacification |  |  |  | Monocular |  | 3개월 | 안 | 45 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | Vision blurred |  |  |  | Monocular |  | 3개월 | 안 | 45 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | Conjunctivitis allergic |  |  |  | Monocular |  | 3개월 | 안 | 45 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | Corneal abrasion |  |  |  | Monocular |  | 3개월 | 안 | 45 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | Corneal edema |  |  |  | Monocular |  | 3개월 | 안 | 45 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | Meibomian gland dysfunction |  |  |  | Monocular |  | 3개월 | 안 | 45 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | Conjunctivitis |  |  |  | Monocular |  | 3개월 | 안 | 45 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | Corneal opacity |  |  |  | Monocular |  | 3개월 | 안 | 45 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | Device dislocation |  |  |  | Monocular |  | 3개월 | 안 | 45 | 0
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | Myopia |  |  |  | Monocular |  | 3개월 | 안 | 45 | 0
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | Optic disc hemorrhage |  |  |  | Monocular |  | 3개월 | 안 | 45 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | Photopsia |  |  |  | Monocular |  | 3개월 | 안 | 45 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | first implanted eye | Complication | Surgery |  |  |  | Monocular |  | 3개월 | 안 | 45 | 0
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | dry eye |  |  |  | Monocular |  | 3개월 | 안 | 44 | 10
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | Glare |  |  |  | Monocular |  | 3개월 | 안 | 44 | 9
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | Visual impairment |  |  |  | Monocular |  | 3개월 | 안 | 44 | 3
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | Halo vision |  |  |  | Monocular |  | 3개월 | 안 | 44 | 3
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | Foreign body sensation in eyes |  |  |  | Monocular |  | 3개월 | 안 | 44 | 2
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | Vitreous floaters |  |  |  | Monocular |  | 3개월 | 안 | 44 | 3
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | Posterior capsule opacification |  |  |  | Monocular |  | 3개월 | 안 | 44 | 2
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | Vision blurred |  |  |  | Monocular |  | 3개월 | 안 | 44 | 2
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | Conjunctivitis allergic |  |  |  | Monocular |  | 3개월 | 안 | 44 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | Corneal abrasion |  |  |  | Monocular |  | 3개월 | 안 | 44 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | Corneal edema |  |  |  | Monocular |  | 3개월 | 안 | 44 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | Meibomian gland dysfunction |  |  |  | Monocular |  | 3개월 | 안 | 44 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | Conjunctivitis |  |  |  | Monocular |  | 3개월 | 안 | 44 | 0
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | Corneal opacity |  |  |  | Monocular |  | 3개월 | 안 | 44 | 0
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | Device dislocation |  |  |  | Monocular |  | 3개월 | 안 | 44 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | Myopia |  |  |  | Monocular |  | 3개월 | 안 | 44 | 1
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | Optic disc hemorrhage |  |  |  | Monocular |  | 3개월 | 안 | 44 | 0
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | Photopsia |  |  |  | Monocular |  | 3개월 | 안 | 44 | 0
653 | Kim (2020) | Prospective, single-arm, unmasked, nonrandomized, multicenter study | Cataract surgery | Diffractive Trifocal IOL | AcrySof IQ® PanOptix IOL (TFNT00) |  |  | second implanted eye | Complication | Surgery |  |  |  | Monocular |  | 3개월 | 안 | 44 | 1
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 54 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 54 | 4.86
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 54 | 24.84
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0 | 3개월 | 안 | 54 | 51.84
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  | Trifocal | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 54 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 840MP |  |  | Trifocal | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 54 | 3.24
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 841MP |  |  | Trifocal | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 54 | 25.92
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 842MP |  |  | Trifocal | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0 | 3개월 | 안 | 54 | 44.82
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 843MP |  |  | Trifocal | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 844MP |  |  | Trifocal | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 845MP |  |  | Trifocal | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 846MP |  |  | Trifocal | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR00 |  |  | EDOF | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 52 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR01 |  |  | EDOF | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 52 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR02 |  |  | EDOF | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 52 | 29.12
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR03 |  |  | EDOF | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0 | 3개월 | 안 | 52 | 47.94
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR04 |  |  | EDOF | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 52 | 52
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR05 |  |  | EDOF | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 52 | 52
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR06 |  |  | EDOF | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 52 | 52
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR07 |  |  | EDOF | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 52 | 52
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF30 |  |  | segRef | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 50 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF31 |  |  | segRef | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 50 | 3
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF32 |  |  | segRef | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 50 | 20
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF33 |  |  | segRef | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0 | 3개월 | 안 | 50 | 43
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF34 |  |  | segRef | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 50 | 49
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF35 |  |  | segRef | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 50 | 49
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF36 |  |  | segRef | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 50 | 50
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF37 |  |  | segRef | VA | CDVA | corrected | Distant(4m) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 50 | 50
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 54 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 54 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 54 | 1.08
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0 | 3개월 | 안 | 54 | 11.88
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 54 | 35.1
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 54 | 43.74
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 54 | 46.98
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 54 | 51.84
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.5 | 3개월 | 안 | 54 | 52.92
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.6 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  | Trifocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 54 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 840MP |  |  | Trifocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 54 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 841MP |  |  | Trifocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 54 | 11.88
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 842MP |  |  | Trifocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0 | 3개월 | 안 | 54 | 30.24
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 843MP |  |  | Trifocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 54 | 45.9
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 844MP |  |  | Trifocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 54 | 52.92
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 845MP |  |  | Trifocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 54 | 52.92
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 846MP |  |  | Trifocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 845MP |  |  | Trifocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.5 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 846MP |  |  | Trifocal | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.6 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR00 |  |  | EDOF | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 52 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR01 |  |  | EDOF | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 52 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR02 |  |  | EDOF | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 52 | 6.76
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR03 |  |  | EDOF | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0 | 3개월 | 안 | 52 | 33.8
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR04 |  |  | EDOF | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 52 | 46.8
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR05 |  |  | EDOF | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 52 | 52
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR06 |  |  | EDOF | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 52 | 52
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR07 |  |  | EDOF | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 52 | 52
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR07 |  |  | EDOF | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.5 | 3개월 | 안 | 52 | 52
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR07 |  |  | EDOF | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.6 | 3개월 | 안 | 52 | 52
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF30 |  |  | segRef | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 50 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF31 |  |  | segRef | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 50 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF32 |  |  | segRef | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 50 | 1
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF33 |  |  | segRef | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0 | 3개월 | 안 | 50 | 14
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF34 |  |  | segRef | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 50 | 33
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF35 |  |  | segRef | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 50 | 37
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF36 |  |  | segRef | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 50 | 42
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF37 |  |  | segRef | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 50 | 47
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF37 |  |  | segRef | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.5 | 3개월 | 안 | 50 | 48
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF37 |  |  | segRef | VA | DCIVA | distance-corrected | intermediate(80cm) | LogMAR | Monocular | 0.6 | 3개월 | 안 | 50 | 50
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 54 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 54 | 2.16
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 54 | 11.88
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0 | 3개월 | 안 | 54 | 25.92
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 54 | 43.74
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 54 | 51.84
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.5 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.6 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.7 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR00 |  |  | EDOF | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 52 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR01 |  |  | EDOF | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 52 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR02 |  |  | EDOF | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 52 | 6.24
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR03 |  |  | EDOF | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0 | 3개월 | 안 | 52 | 28.08
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR04 |  |  | EDOF | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 52 | 39
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR05 |  |  | EDOF | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 52 | 45.76
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR06 |  |  | EDOF | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 52 | 46.8
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR07 |  |  | EDOF | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 52 | 49.92
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR07 |  |  | EDOF | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.5 | 3개월 | 안 | 52 | 52
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR07 |  |  | EDOF | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.6 | 3개월 | 안 | 52 | 52
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR07 |  |  | EDOF | VA | DCIVA | distance-corrected | Intermediate(60cm) | LogMAR | Monocular | 0.7 | 3개월 | 안 | 52 | 52
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 54 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 54 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 54 | 5.94
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0 | 3개월 | 안 | 54 | 27
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 54 | 49.14
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 54 | 50.76
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.5 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.6 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 839MP |  |  | Trifocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 54 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 840MP |  |  | Trifocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 54 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 841MP |  |  | Trifocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 54 | 11.88
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 842MP |  |  | Trifocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0 | 3개월 | 안 | 54 | 25.92
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 843MP |  |  | Trifocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 54 | 42.12
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 844MP |  |  | Trifocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 54 | 50.22
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 845MP |  |  | Trifocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 54 | 52.92
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 846MP |  |  | Trifocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 54 | 52.92
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 845MP |  |  | Trifocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.5 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 846MP |  |  | Trifocal | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.6 | 3개월 | 안 | 54 | 54
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR00 |  |  | EDOF | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 52 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR01 |  |  | EDOF | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 52 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR02 |  |  | EDOF | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 52 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR03 |  |  | EDOF | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0 | 3개월 | 안 | 52 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR04 |  |  | EDOF | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 52 | 5.2
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR05 |  |  | EDOF | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 52 | 22.88
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR06 |  |  | EDOF | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 52 | 37.96
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR07 |  |  | EDOF | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 52 | 44.2
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR07 |  |  | EDOF | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.5 | 3개월 | 안 | 52 | 45.76
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR07 |  |  | EDOF | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.6 | 3개월 | 안 | 52 | 52
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF30 |  |  | segRef | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | -0.3 | 3개월 | 안 | 50 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF31 |  |  | segRef | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | -0.2 | 3개월 | 안 | 50 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF32 |  |  | segRef | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | -0.1 | 3개월 | 안 | 50 | 3
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF33 |  |  | segRef | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0 | 3개월 | 안 | 50 | 17
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF34 |  |  | segRef | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.1 | 3개월 | 안 | 50 | 30
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF35 |  |  | segRef | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.2 | 3개월 | 안 | 50 | 42
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF36 |  |  | segRef | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.3 | 3개월 | 안 | 50 | 47
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF37 |  |  | segRef | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.4 | 3개월 | 안 | 50 | 50
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF37 |  |  | segRef | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.5 | 3개월 | 안 | 50 | 50
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF37 |  |  | segRef | VA | DCNVA | distance-corrected | near(40cm) | LogMAR | Monocular | 0.6 | 3개월 | 안 | 50 | 50
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  | Panfocal | Spectacle independence |  |  |  | Y/N |  |  | 3개월 | 명 | 27 | 26
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 846MP |  |  | Trifocal | Spectacle independence |  |  |  | Y/N |  |  | 3개월 | 명 | 26 | 25
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR00 |  |  | EDOF | Spectacle independence |  |  |  | Y/N |  |  | 3개월 | 명 | 24 | 17
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF37 |  |  | segRef | Spectacle independence |  |  |  | Y/N |  |  | 3개월 | 명 | 25 | 23
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AcrySof PanOptix |  |  |  | Complication |  |  |  | Y/N |  |  | 3개월 | 명 | 27 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive Trifocal IOL | AT LISA tri 846MP |  |  |  | Complication |  |  |  | Y/N |  |  | 3개월 | 명 | 26 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Diffractive EDOF toric IOL | TECNIS Symfony ZXR00 |  |  |  | Complication |  |  |  | Y/N |  |  | 3개월 | 명 | 24 | 0
691 | Bohm (2019) | Prospective case series | Presbyopia | Refractive Bifocal IOL | LENTIS Mplus X LS-313 MF37 |  |  |  | Complication |  |  |  | Y/N |  |  | 3개월 | 명 | 25 | 0
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  | operated eye | VA | UDVA | uncorrected | Distant | Snellen | Monocular | 20/25 or better | 6개월 | 안 | 26 | 18
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | VA | UDVA | uncorrected | Distant | Snellen | Binocular | 20/25 or better | 6개월 | 명 | 26 | 24
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | VA | UDVA | uncorrected | Distant | Snellen | Binocular | 20/30 | 6개월 | 명 | 26 | 2
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  | operated eye | VA | UNVA | uncorrected | Near(30cm) | Snellen | Monocular | 20/25 or better | 6개월 | 안 | 26 | 23
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  | operated eye | VA | UNVA | uncorrected | Near(30cm) | Snellen | Monocular | 20/30 | 6개월 | 안 | 26 | 3
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | VA | CDVA | corrected | Distant | Snellen |  | 변화없음 | 6개월 | 명 | 26 | 24
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | VA | CIVA | corrected | Intermediate(60cm) | Snellen |  | 변화없음 | 6개월 | 명 | 26 | 23
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | VA | CNVA | corrected | Near(30cm) | Snellen |  | 변화없음 | 6개월 | 명 | 26 | 24
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | VA | CDVA | corrected | Distant | Snellen |  | loss of 1 line | 6개월 | 명 | 26 | 2
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | VA | CIVA | corrected | Intermediate(60cm) | Snellen |  | loss of 1 line | 6개월 | 명 | 26 | 3
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | VA | CNVA | corrected | Near(30cm) | Snellen |  | loss of 1 line | 6개월 | 명 | 26 | 2
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | VA | CDVA | corrected | Distant | Snellen |  | loss more than one line | 6개월 | 명 | 26 | 0
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | VA | CIVA | corrected | Intermediate(60cm) | Snellen |  | loss more than one line | 6개월 | 명 | 26 | 0
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | VA | CNVA | corrected | Near(30cm) | Snellen |  | loss more than one line | 6개월 | 명 | 26 | 0
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Complication | intraopeartive complication | [6 empty] | 명 | 26 | 0
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Complication | IOL exchange | [6 empty] | 명 | 26 | 0
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | Glare |  |  | 1-6점(1점 best) |  | 1 | 6개월 | 명 | 26 | 8.06
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | Glare |  |  | 1-6점(1점 best) |  | 2 | 6개월 | 명 | 26 | 1.04
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | Glare |  |  | 1-6점(1점 best) |  | 3 | 6개월 | 명 | 26 | 3.9
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | Glare |  |  | 1-6점(1점 best) |  | 4 | 6개월 | 명 | 26 | 3.12
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | Glare |  |  | 1-6점(1점 best) |  | 5 | 6개월 | 명 | 26 | 2.08
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | Glare |  |  | 1-6점(1점 best) |  | 6 | 6개월 | 명 | 26 | 3.12
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | Halos |  |  | 1-6점(1점 best) |  | 1 | 6개월 | 명 | 26 | 7.02
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | Halos |  |  | 1-6점(1점 best) |  | 2 | 6개월 | 명 | 26 | 1.04
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | Halos |  |  | 1-6점(1점 best) |  | 3 | 6개월 | 명 | 26 | 4.94
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | Halos |  |  | 1-6점(1점 best) |  | 4 | 6개월 | 명 | 26 | 3.12
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | Halos |  |  | 1-6점(1점 best) |  | 5 | 6개월 | 명 | 26 | 1.04
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | Halos |  |  | 1-6점(1점 best) |  | 6 | 6개월 | 명 | 26 | 3.9
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | starburst |  |  | 1-6점(1점 best) |  | 1 | 6개월 | 명 | 26 | 14.04
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | starburst |  |  | 1-6점(1점 best) |  | 2 | 6개월 | 명 | 26 | 2.08
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | starburst |  |  | 1-6점(1점 best) |  | 3 | 6개월 | 명 | 26 | 1.04
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | starburst |  |  | 1-6점(1점 best) |  | 4 | 6개월 | 명 | 26 | 2.08
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | starburst |  |  | 1-6점(1점 best) |  | 5 | 6개월 | 명 | 26 | 0
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | starburst |  |  | 1-6점(1점 best) |  | 6 | 6개월 | 명 | 26 | 2.08
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | double vision |  |  | 1-6점(1점 best) |  | 1 | 6개월 | 명 | 26 | 16.12
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | double vision |  |  | 1-6점(1점 best) |  | 2 | 6개월 | 명 | 26 | 0
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | double vision |  |  | 1-6점(1점 best) |  | 3 | 6개월 | 명 | 26 | 2.08
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | double vision |  |  | 1-6점(1점 best) |  | 4 | 6개월 | 명 | 26 | 2.08
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | double vision |  |  | 1-6점(1점 best) |  | 5 | 6개월 | 명 | 26 | 0
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | double vision |  |  | 1-6점(1점 best) |  | 6 | 6개월 | 명 | 26 | 1.04
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | overall satisfaction |  |  | Y/N |  |  | 6개월 | 명 | 26 | 22.88
702 | Levinger (2019) | Interventional case series | Presbyopia | Diffractive Trifocal IOL | FineVision Micro F |  |  |  | Satisfaction | recommend this procedure to friend and family |  |  | Y/N |  |  | 6개월 | 명 | 26 | 24.96
732 | Vounotrypidis (2017) | Prospective longitudinal case series | Cataract surgery | Refractive EDOF IOL | Lentis comfort LS-313 MF15 |  |  |  | Satisfaction | photic phenomena independent of light conditions |  |  | Y/N |  |  | 3개월 | 명 | 22 | 2
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | 20/12.5 | 수술전 | 안 | 70 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | 20/16 | 수술전 | 안 | 70 | 13.3
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | 20/20 | 수술전 | 안 | 70 | 39.9
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | 20/25 | 수술전 | 안 | 70 | 58.1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | 20/32 | 수술전 | 안 | 70 | 65.8
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | 20/40 | 수술전 | 안 | 70 | 69.3
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | 20/63 | 수술전 | 안 | 70 | 69.3
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | 20/150 | 수술전 | 안 | 70 | 70
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UDVA | uncorrected | Distant(3m) | Snellen | Monocular | 20/12.5 | 6개월 | 안 | 72 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UDVA | uncorrected | Distant(3m) | Snellen | Monocular | 20/16 | 6개월 | 안 | 72 | 23.76
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UDVA | uncorrected | Distant(3m) | Snellen | Monocular | 20/20 | 6개월 | 안 | 72 | 48.96
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UDVA | uncorrected | Distant(3m) | Snellen | Monocular | 20/25 | 6개월 | 안 | 72 | 59.76
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UDVA | uncorrected | Distant(3m) | Snellen | Monocular | 20/32 | 6개월 | 안 | 72 | 69.12
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UDVA | uncorrected | Distant(3m) | Snellen | Monocular | 20/40 | 6개월 | 안 | 72 | 71.28
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UDVA | uncorrected | Distant(3m) | Snellen | Monocular | 20/63 | 6개월 | 안 | 72 | 72
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UDVA | uncorrected | Distant(3m) | Snellen | Monocular | 20/150 | 6개월 | 안 | 72 | 72
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | loss ≥ 3 | 6개월 | 안 | 72 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | loss 2 | 6개월 | 안 | 72 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | loss 1 | 6개월 | 안 | 72 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | 변화없음 | 6개월 | 안 | 72 | 22.32
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | gain 1 | 6개월 | 안 | 72 | 23.76
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | gain 2 | 6개월 | 안 | 72 | 14.4
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | gain ≥ 3 | 6개월 | 안 | 72 | 11.52
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UDVA | uncorrected | Distant(3m) | Snellen | Monocular | 20/20 or better | 6개월 | 안 | 72 | 48.96
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UDVA | uncorrected | Distant(3m) | Snellen | Monocular | 20/25 or better | 6개월 | 안 | 72 | 59.76
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UDVA | uncorrected | Distant(3m) | Snellen | Monocular | 20/32 or better | 6개월 | 안 | 72 | 69.12
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UDVA | uncorrected | Distant(3m) | Snellen | Monocular | 20/40 or better | 6개월 | 안 | 72 | 71.28
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UDVA | uncorrected | Distant(3m) | Snellen | Monocular | 20/50 or better | 6개월 | 안 | 72 | 72
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | 20/20 or better | 6개월 | 안 | 72 | 72
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | 20/25 or better | 6개월 | 안 | 72 | 72
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | 20/32 or better | 6개월 | 안 | 72 | 72
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | 20/40 or better | 6개월 | 안 | 72 | 72
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Monocular | 20/50 or better | 6개월 | 안 | 72 | 72
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UDVA | uncorrected | Distant(3m) | Snellen | Binocular | 20/20 or better | 6개월 | 명 | 36 | 33.12
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UDVA | uncorrected | Distant(3m) | Snellen | Binocular | 20/25 or better | 6개월 | 명 | 36 | 33.12
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UDVA | uncorrected | Distant(3m) | Snellen | Binocular | 20/32 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Binocular | 20/20 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Binocular | 20/25 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | CDVA | corrected | Distant(3m) | Snellen | Binocular | 20/32 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mesopic | VA | CDVA | corrected | Distant(3m) | Snellen | Binocular | 20/20 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mesopic | VA | CDVA | corrected | Distant(3m) | Snellen | Binocular | 20/25 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mesopic | VA | CDVA | corrected | Distant(3m) | Snellen | Binocular | 20/32 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UIVA | uncorrected | Intermediate(67cm) | Snellen | Monocular | 20/20 or better | 6개월 | 안 | 72 | 2.88
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UIVA | uncorrected | Intermediate(67cm) | Snellen | Monocular | 20/25 or better | 6개월 | 안 | 72 | 12.96
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UIVA | uncorrected | Intermediate(67cm) | Snellen | Monocular | 20/32 or better | 6개월 | 안 | 72 | 28.8
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UIVA | uncorrected | Intermediate(67cm) | Snellen | Monocular | 20/40 or better | 6개월 | 안 | 72 | 41.76
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UIVA | uncorrected | Intermediate(67cm) | Snellen | Monocular | 20/50 or better | 6개월 | 안 | 72 | 58.32
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UIVA | uncorrected | Intermediate(67cm) | Snellen | Monocular | 20/63 or better | 6개월 | 안 | 72 | 64.8
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UIVA | uncorrected | Intermediate(67cm) | Snellen | Monocular | 20/80 or better | 6개월 | 안 | 72 | 69.84
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UIVA | uncorrected | Intermediate(67cm) | Snellen | Monocular | 20/100 or better | 6개월 | 안 | 72 | 72
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Monocular | 20/20 or better | 6개월 | 안 | 72 | 0.72
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Monocular | 20/25 or better | 6개월 | 안 | 72 | 7.92
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Monocular | 20/32 or better | 6개월 | 안 | 72 | 20.16
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Monocular | 20/40 or better | 6개월 | 안 | 72 | 28.08
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Monocular | 20/50 or better | 6개월 | 안 | 72 | 56.16
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Monocular | 20/63 or better | 6개월 | 안 | 72 | 64.8
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Monocular | 20/80 or better | 6개월 | 안 | 72 | 69.12
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Monocular | 20/100 or better | 6개월 | 안 | 72 | 72
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UIVA | uncorrected | Intermediate(67cm) | Snellen | Binocular | 20/20 or better | 6개월 | 명 | 36 | 6.12
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UIVA | uncorrected | Intermediate(67cm) | Snellen | Binocular | 20/25 or better | 6개월 | 명 | 36 | 15.84
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UIVA | uncorrected | Intermediate(67cm) | Snellen | Binocular | 20/32 or better | 6개월 | 명 | 36 | 27
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UIVA | uncorrected | Intermediate(67cm) | Snellen | Binocular | 20/40 or better | 6개월 | 명 | 36 | 30.96
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UIVA | uncorrected | Intermediate(67cm) | Snellen | Binocular | 20/50 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UIVA | uncorrected | Intermediate(67cm) | Snellen | Binocular | 20/63 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UIVA | uncorrected | Intermediate(67cm) | Snellen | Binocular | 20/80 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Binocular | 20/20 or better | 6개월 | 명 | 36 | 2.16
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Binocular | 20/25 or better | 6개월 | 명 | 36 | 12.96
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Binocular | 20/32 or better | 6개월 | 명 | 36 | 19.08
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Binocular | 20/40 or better | 6개월 | 명 | 36 | 27
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Binocular | 20/50 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Binocular | 20/63 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Binocular | 20/80 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mesopic | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Binocular | 20/20 or better | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mesopic | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Binocular | 20/25 or better | 6개월 | 명 | 36 | 1.08
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mesopic | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Binocular | 20/32 or better | 6개월 | 명 | 36 | 1.08
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mesopic | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Binocular | 20/40 or better | 6개월 | 명 | 36 | 5.04
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mesopic | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Binocular | 20/50 or better | 6개월 | 명 | 36 | 20.88
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mesopic | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Binocular | 20/63 or better | 6개월 | 명 | 36 | 30.96
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mesopic | VA | DCIVA | distance-corrected | Intermediate(67cm) | Snellen | Binocular | 20/80 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UNVA | uncorrected | Near(33cm) | Snellen | Monocular | 20/20 or better | 6개월 | 안 | 72 | 5.76
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UNVA | uncorrected | Near(33cm) | Snellen | Monocular | 20/25 or better | 6개월 | 안 | 72 | 41.76
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UNVA | uncorrected | Near(33cm) | Snellen | Monocular | 20/32 or better | 6개월 | 안 | 72 | 63.36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UNVA | uncorrected | Near(33cm) | Snellen | Monocular | 20/40 or better | 6개월 | 안 | 72 | 71.28
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UNVA | uncorrected | Near(33cm) | Snellen | Monocular | 20/50 or better | 6개월 | 안 | 72 | 72
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCNVA | distance-corrected | Near(33cm) | Snellen | Monocular | 20/20 or better | 6개월 | 안 | 72 | 12.96
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCNVA | distance-corrected | Near(33cm) | Snellen | Monocular | 20/25 or better | 6개월 | 안 | 72 | 54.72
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCNVA | distance-corrected | Near(33cm) | Snellen | Monocular | 20/32 or better | 6개월 | 안 | 72 | 66.24
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCNVA | distance-corrected | Near(33cm) | Snellen | Monocular | 20/40 or better | 6개월 | 안 | 72 | 72
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCNVA | distance-corrected | Near(33cm) | Snellen | Monocular | 20/50 or better | 6개월 | 안 | 72 | 72
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UNVA | uncorrected | Near(33cm) | Snellen | Binocular | 20/20 or better | 6개월 | 명 | 36 | 7.92
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UNVA | uncorrected | Near(33cm) | Snellen | Binocular | 20/25 or better | 6개월 | 명 | 36 | 29.88
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UNVA | uncorrected | Near(33cm) | Snellen | Binocular | 20/32 or better | 6개월 | 명 | 36 | 34.92
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UNVA | uncorrected | Near(33cm) | Snellen | Binocular | 20/40 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | UNVA | uncorrected | Near(33cm) | Snellen | Binocular | 20/50 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCNVA | distance-corrected | Near(33cm) | Snellen | Binocular | 20/20 or better | 6개월 | 명 | 36 | 10.08
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCNVA | distance-corrected | Near(33cm) | Snellen | Binocular | 20/25 or better | 6개월 | 명 | 36 | 32.04
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCNVA | distance-corrected | Near(33cm) | Snellen | Binocular | 20/32 or better | 6개월 | 명 | 36 | 34.92
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCNVA | distance-corrected | Near(33cm) | Snellen | Binocular | 20/40 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | VA | DCNVA | distance-corrected | Near(33cm) | Snellen | Binocular | 20/50 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mesopic | VA | DCNVA | distance-corrected | Near(33cm) | Snellen | Binocular | 20/20 or better | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mesopic | VA | DCNVA | distance-corrected | Near(33cm) | Snellen | Binocular | 20/25 or better | 6개월 | 명 | 36 | 9
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mesopic | VA | DCNVA | distance-corrected | Near(33cm) | Snellen | Binocular | 20/32 or better | 6개월 | 명 | 36 | 29.88
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mesopic | VA | DCNVA | distance-corrected | Near(33cm) | Snellen | Binocular | 20/40 or better | 6개월 | 명 | 36 | 34.92
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | mesopic | VA | DCNVA | distance-corrected | Near(33cm) | Snellen | Binocular | 20/50 or better | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Complication | intraopeartive complication | [5 empty] | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Complication | IOL exchange | [5 empty] | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | During the day | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Very Satisfied | 3개월 | 명 | 36 | 13
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | During the day | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Satisfied | 3개월 | 명 | 36 | 21
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | During the day | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Neutral | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | During the day | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Dissatisfied | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | During the day | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Very Dissatisfied | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | at night | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Very Satisfied | 3개월 | 명 | 36 | 9
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | at night | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Satisfied | 3개월 | 명 | 36 | 23
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | at night | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Neutral | 3개월 | 명 | 36 | 4
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | at night | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Dissatisfied | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | at night | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Very Dissatisfied | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Very Satisfied | 3개월 | 명 | 36 | 10
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Satisfied | 3개월 | 명 | 36 | 24
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Neutral | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Dissatisfied | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Very Dissatisfied | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | During the day | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Very Satisfied | 6개월 | 명 | 36 | 16
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | During the day | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Satisfied | 6개월 | 명 | 36 | 18
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | During the day | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Neutral | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | During the day | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Dissatisfied | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | During the day | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Very Dissatisfied | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | at night | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Very Satisfied | 6개월 | 명 | 36 | 12
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | at night | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Satisfied | 6개월 | 명 | 36 | 21
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | at night | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Neutral | 6개월 | 명 | 36 | 2
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | at night | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Dissatisfied | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | at night | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Very Dissatisfied | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Very Satisfied | 6개월 | 명 | 36 | 13
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Satisfied | 6개월 | 명 | 36 | 20
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Neutral | 6개월 | 명 | 36 | 2
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Dissatisfied | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how satisfied have you been with your vision without contact lenses and/or glasses? |  |  |  |  | Very Dissatisfied | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | glass | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Never | 3개월 | 명 | 36 | 31
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | glass | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Hardly Ever | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | glass | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Sometimes | 3개월 | 명 | 36 | 2
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | glass | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Often | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | glass | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Always | 3개월 | 명 | 36 | 2
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | glass | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Never | 6개월 | 명 | 36 | 30
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | glass | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Hardly Ever | 6개월 | 명 | 36 | 3
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | glass | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Sometimes | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | glass | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Often | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | glass | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Always | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Contact lens | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Never | 3개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Contact lens | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Hardly Ever | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Contact lens | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Sometimes | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Contact lens | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Often | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Contact lens | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Always | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Contact lens | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Never | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Contact lens | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Hardly Ever | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Contact lens | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Sometimes | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Contact lens | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Often | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Contact lens | Spectacle independence | Over the past month, how often did you wear the following to correct your vision? |  |  |  |  | Always | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | During the past month, how much more convenient has it been to see without glasses or contact lenses while performing your normal daily activities? |  |  |  |  | Much More | 3개월 | 명 | 36 | 31
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | During the past month, how much more convenient has it been to see without glasses or contact lenses while performing your normal daily activities? |  |  |  |  | More | 3개월 | 명 | 36 | 4
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | During the past month, how much more convenient has it been to see without glasses or contact lenses while performing your normal daily activities? |  |  |  |  | The Same | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | During the past month, how much more convenient has it been to see without glasses or contact lenses while performing your normal daily activities? |  |  |  |  | Less | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | During the past month, how much more convenient has it been to see without glasses or contact lenses while performing your normal daily activities? |  |  |  |  | Much Less | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | During the past month, how much more convenient has it been to see without glasses or contact lenses while performing your normal daily activities? |  |  |  |  | Much More | 6개월 | 명 | 36 | 32
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | During the past month, how much more convenient has it been to see without glasses or contact lenses while performing your normal daily activities? |  |  |  |  | More | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | During the past month, how much more convenient has it been to see without glasses or contact lenses while performing your normal daily activities? |  |  |  |  | The Same | 6개월 | 명 | 36 | 2
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | During the past month, how much more convenient has it been to see without glasses or contact lenses while performing your normal daily activities? |  |  |  |  | Less | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | During the past month, how much more convenient has it been to see without glasses or contact lenses while performing your normal daily activities? |  |  |  |  | Much Less | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Does having intraocular lenses give you the freedom to perform multiple activities without depending upon glasses and/or contact lenses? |  |  |  |  | Always | 3개월 | 명 | 36 | 34
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Does having intraocular lenses give you the freedom to perform multiple activities without depending upon glasses and/or contact lenses? |  |  |  |  | Usually | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Does having intraocular lenses give you the freedom to perform multiple activities without depending upon glasses and/or contact lenses? |  |  |  |  | Sometimes | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Does having intraocular lenses give you the freedom to perform multiple activities without depending upon glasses and/or contact lenses? |  |  |  |  | Not Often | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Does having intraocular lenses give you the freedom to perform multiple activities without depending upon glasses and/or contact lenses? |  |  |  |  | Never | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Does having intraocular lenses give you the freedom to perform multiple activities without depending upon glasses and/or contact lenses? |  |  |  |  | Always | 6개월 | 명 | 36 | 34
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Does having intraocular lenses give you the freedom to perform multiple activities without depending upon glasses and/or contact lenses? |  |  |  |  | Usually | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Does having intraocular lenses give you the freedom to perform multiple activities without depending upon glasses and/or contact lenses? |  |  |  |  | Sometimes | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Does having intraocular lenses give you the freedom to perform multiple activities without depending upon glasses and/or contact lenses? |  |  |  |  | Not Often | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Does having intraocular lenses give you the freedom to perform multiple activities without depending upon glasses and/or contact lenses? |  |  |  |  | Never | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Very Easy | 3개월 | 명 | 36 | 12
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Easy | 3개월 | 명 | 36 | 10
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Difficult | 3개월 | 명 | 36 | 13
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Very Difficult | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Impossible | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Very Easy | 6개월 | 명 | 36 | 16
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Easy | 6개월 | 명 | 36 | 9
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Difficult | 6개월 | 명 | 36 | 10
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Very Difficult | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Impossible | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Very Easy | 3개월 | 명 | 36 | 8
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Easy | 3개월 | 명 | 36 | 13
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Difficult | 3개월 | 명 | 36 | 13
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Very Difficult | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Impossible | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Very Easy | 6개월 | 명 | 36 | 14
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Easy | 6개월 | 명 | 36 | 10
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Difficult | 6개월 | 명 | 36 | 11
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Very Difficult | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Impossible | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Very Easy | 3개월 | 명 | 36 | 22
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Easy | 3개월 | 명 | 36 | 13
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Difficult | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Very Difficult | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Impossible | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Very Easy | 6개월 | 명 | 36 | 24
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Easy | 6개월 | 명 | 36 | 11
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Difficult | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Very Difficult | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Impossible | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Driving at night | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Very Easy | 3개월 | 명 | 15 | 6
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Driving at night | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Easy | 3개월 | 명 | 15 | 5
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Driving at night | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Difficult | 3개월 | 명 | 15 | 2
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Driving at night | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Very Difficult | 3개월 | 명 | 15 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Driving at night | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Impossible | 3개월 | 명 | 15 | 2
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Driving at night | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Very Easy | 6개월 | 명 | 15 | 6
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Driving at night | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Easy | 6개월 | 명 | 15 | 7
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Driving at night | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Difficult | 6개월 | 명 | 15 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Driving at night | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Very Difficult | 6개월 | 명 | 15 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Driving at night | Spectacle independence | Since your surgery, how easy is it to perform the following activities without glasses and/or contact lenses? |  |  |  |  | Impossible | 6개월 | 명 | 15 | 2
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how would you rate the distance at which it is comfortable for you to read? |  |  |  |  | Much Too Close | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how would you rate the distance at which it is comfortable for you to read? |  |  |  |  | Slightly Too Close | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how would you rate the distance at which it is comfortable for you to read? |  |  |  |  | Perfect | 3개월 | 명 | 36 | 34
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how would you rate the distance at which it is comfortable for you to read? |  |  |  |  | Slightly Too Far | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how would you rate the distance at which it is comfortable for you to read? |  |  |  |  | Much Too Far | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how would you rate the distance at which it is comfortable for you to read? |  |  |  |  | Much Too Close | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how would you rate the distance at which it is comfortable for you to read? |  |  |  |  | Slightly Too Close | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how would you rate the distance at which it is comfortable for you to read? |  |  |  |  | Perfect | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how would you rate the distance at which it is comfortable for you to read? |  |  |  |  | Slightly Too Far | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | Since your surgery, how would you rate the distance at which it is comfortable for you to read? |  |  |  |  | Much Too Far | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Never | 3개월 | 명 | 36 | 27
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Hardly Ever | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Sometimes | 3개월 | 명 | 36 | 7
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Often | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Always | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Never | 6개월 | 명 | 36 | 30
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Hardly Ever | 6개월 | 명 | 36 | 4
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Sometimes | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Often | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Reading small print in dim lights | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Always | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Never | 3개월 | 명 | 36 | 30
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Hardly Ever | 3개월 | 명 | 36 | 3
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Sometimes | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Often | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Always | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Never | 6개월 | 명 | 36 | 32
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Hardly Ever | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Sometimes | 6개월 | 명 | 36 | 3
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Often | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing at arm’s length | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Always | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Never | 3개월 | 명 | 36 | 35
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Hardly Ever | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Sometimes | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Often | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Always | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Never | 6개월 | 명 | 36 | 35
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Hardly Ever | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Sometimes | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Often | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Seeing far away | Satisfaction | How often have you been bothered by fluctuations in your vision when |  |  |  |  | Always | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Halo around lights | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Not Present | 3개월 | 명 | 36 | 10
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Halo around lights | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Present But Not Bothersome | 3개월 | 명 | 36 | 23
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Halo around lights | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Mildly Bothersome | 3개월 | 명 | 36 | 3
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Halo around lights | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Moderately Bothersome | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Halo around lights | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Very Bothersome | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Halo around lights | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Not Present | 6개월 | 명 | 36 | 15
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Halo around lights | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Present But Not Bothersome | 6개월 | 명 | 36 | 17
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Halo around lights | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Mildly Bothersome | 6개월 | 명 | 36 | 3
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Halo around lights | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Moderately Bothersome | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Halo around lights | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Very Bothersome | 6개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty with glare at night | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Not Present | 3개월 | 명 | 36 | 11
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty with glare at night | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Present But Not Bothersome | 3개월 | 명 | 36 | 18
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty with glare at night | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Mildly Bothersome | 3개월 | 명 | 36 | 3
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty with glare at night | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Moderately Bothersome | 3개월 | 명 | 36 | 4
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty with glare at night | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Very Bothersome | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty with glare at night | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Not Present | 6개월 | 명 | 36 | 15
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty with glare at night | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Present But Not Bothersome | 6개월 | 명 | 36 | 16
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty with glare at night | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Mildly Bothersome | 6개월 | 명 | 36 | 3
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty with glare at night | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Moderately Bothersome | 6개월 | 명 | 36 | 2
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty with glare at night | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Very Bothersome | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Double images/ghosting | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Not Present | 3개월 | 명 | 36 | 35
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Double images/ghosting | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Present But Not Bothersome | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Double images/ghosting | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Mildly Bothersome | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Double images/ghosting | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Moderately Bothersome | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Double images/ghosting | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Very Bothersome | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Double images/ghosting | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Not Present | 6개월 | 명 | 36 | 36
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Double images/ghosting | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Present But Not Bothersome | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Double images/ghosting | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Mildly Bothersome | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Double images/ghosting | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Moderately Bothersome | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Double images/ghosting | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Very Bothersome | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty reading small print | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Not Present | 3개월 | 명 | 36 | 22
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty reading small print | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Present But Not Bothersome | 3개월 | 명 | 36 | 10
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty reading small print | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Mildly Bothersome | 3개월 | 명 | 36 | 3
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty reading small print | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Moderately Bothersome | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty reading small print | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Very Bothersome | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty reading small print | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Not Present | 6개월 | 명 | 36 | 25
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty reading small print | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Present But Not Bothersome | 6개월 | 명 | 36 | 7
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty reading small print | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Mildly Bothersome | 6개월 | 명 | 36 | 2
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty reading small print | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Moderately Bothersome | 6개월 | 명 | 36 | 2
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty reading small print | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Very Bothersome | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty seeing far away | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Not Present | 3개월 | 명 | 36 | 32
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty seeing far away | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Present But Not Bothersome | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty seeing far away | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Mildly Bothersome | 3개월 | 명 | 36 | 3
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty seeing far away | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Moderately Bothersome | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty seeing far away | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Very Bothersome | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty seeing far away | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Not Present | 6개월 | 명 | 36 | 32
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty seeing far away | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Present But Not Bothersome | 6개월 | 명 | 36 | 2
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty seeing far away | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Mildly Bothersome | 6개월 | 명 | 36 | 2
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty seeing far away | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Moderately Bothersome | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  | Difficulty seeing far away | Satisfaction | Please indicate any of the following symptoms you currently experience related to your vision |  |  |  |  | Very Bothersome | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | If given the opportunity, would you elect to be implanted with the same intraocular lenses? |  |  |  |  | Definitely | 3개월 | 명 | 36 | 32
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | If given the opportunity, would you elect to be implanted with the same intraocular lenses? |  |  |  |  | Most Likely | 3개월 | 명 | 36 | 3
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | If given the opportunity, would you elect to be implanted with the same intraocular lenses? |  |  |  |  | Maybe | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | If given the opportunity, would you elect to be implanted with the same intraocular lenses? |  |  |  |  | Unlikely | 3개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | If given the opportunity, would you elect to be implanted with the same intraocular lenses? |  |  |  |  | Definitely Not | 3개월 | 명 | 36 | 1
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | If given the opportunity, would you elect to be implanted with the same intraocular lenses? |  |  |  |  | Definitely | 6개월 | 명 | 36 | 32
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | If given the opportunity, would you elect to be implanted with the same intraocular lenses? |  |  |  |  | Most Likely | 6개월 | 명 | 36 | 3
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | If given the opportunity, would you elect to be implanted with the same intraocular lenses? |  |  |  |  | Maybe | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | If given the opportunity, would you elect to be implanted with the same intraocular lenses? |  |  |  |  | Unlikely | 6개월 | 명 | 36 | 0
782 | Chang (2019) | Prospective, observational case series | Cataract surgery, Presbyopia | Diffractive Bifocal IOL | TECNIS ZMB00 |  |  |  | Satisfaction | If given the opportunity, would you elect to be implanted with the same intraocular lenses? |  |  |  |  | Definitely Not | 6개월 | 명 | 36 | 1
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.00 | final visit(2-4개월) | 안 | 48 | 19
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.10 | final visit(2-4개월) | 안 | 48 | 30
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.20 | final visit(2-4개월) | 안 | 48 | 41
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.30 | final visit(2-4개월) | 안 | 48 | 44
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.00 | final visit(2-4개월) | 안 | 46 | 24
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.10 | final visit(2-4개월) | 안 | 46 | 30
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.20 | final visit(2-4개월) | 안 | 46 | 40
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.30 | final visit(2-4개월) | 안 | 46 | 44
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.00 | final visit(2-4개월) | 안 | 34 | 1
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.10 | final visit(2-4개월) | 안 | 34 | 13
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.20 | final visit(2-4개월) | 안 | 34 | 24
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.30 | final visit(2-4개월) | 안 | 34 | 26
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.00 | final visit(2-4개월) | 안 | 13 | 6
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.10 | final visit(2-4개월) | 안 | 13 | 7
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.20 | final visit(2-4개월) | 안 | 13 | 12
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.30 | final visit(2-4개월) | 안 | 12 | 13
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.00 | final visit(2-4개월) | 안 | 12 | 7
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.10 | final visit(2-4개월) | 안 | 12 | 9
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.20 | final visit(2-4개월) | 안 | 12 | 10
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.30 | final visit(2-4개월) | 안 | 12 | 12
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.00 | final visit(2-4개월) | 안 | 9 | 1
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.10 | final visit(2-4개월) | 안 | 9 | 3
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.20 | final visit(2-4개월) | 안 | 9 | 6
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.30 | final visit(2-4개월) | 안 | 9 | 7
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.00 | final visit(2-4개월) | 안 | 13 | 6
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.10 | final visit(2-4개월) | 안 | 13 | 10
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.20 | final visit(2-4개월) | 안 | 13 | 11
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.30 | final visit(2-4개월) | 안 | 13 | 12
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.00 | final visit(2-4개월) | 안 | 13 | 7
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.10 | final visit(2-4개월) | 안 | 13 | 9
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.20 | final visit(2-4개월) | 안 | 13 | 12
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.30 | final visit(2-4개월) | 안 | 13 | 12
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.00 | final visit(2-4개월) | 안 | 10 | 0
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.10 | final visit(2-4개월) | 안 | 10 | 7
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.20 | final visit(2-4개월) | 안 | 10 | 9
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.30 | final visit(2-4개월) | 안 | 10 | 9
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.00 | final visit(2-4개월) | 안 | 9 | 2
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.10 | final visit(2-4개월) | 안 | 9 | 4
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.20 | final visit(2-4개월) | 안 | 9 | 6
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.30 | final visit(2-4개월) | 안 | 9 | 7
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.00 | final visit(2-4개월) | 안 | 9 | 2
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.10 | final visit(2-4개월) | 안 | 9 | 3
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.20 | final visit(2-4개월) | 안 | 9 | 6
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.30 | final visit(2-4개월) | 안 | 9 | 8
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.00 | final visit(2-4개월) | 안 | 9 | 0
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.10 | final visit(2-4개월) | 안 | 9 | 2
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.20 | final visit(2-4개월) | 안 | 9 | 3
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.30 | final visit(2-4개월) | 안 | 9 | 4
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.00 | final visit(2-4개월) | 안 | 13 | 5
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.10 | final visit(2-4개월) | 안 | 13 | 9
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.20 | final visit(2-4개월) | 안 | 13 | 12
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | UDVA | uncorrected | Distant | LogMAR | Monocular | ≤0.30 | final visit(2-4개월) | 안 | 13 | 12
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.00 | final visit(2-4개월) | 안 | 12 | 8
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.10 | final visit(2-4개월) | 안 | 12 | 9
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.20 | final visit(2-4개월) | 안 | 12 | 12
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | UNVA | uncorrected | near(40cm) | LogMAR | Monocular | ≤0.30 | final visit(2-4개월) | 안 | 12 | 12
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.00 | final visit(2-4개월) | 안 | 8 | 0
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.10 | final visit(2-4개월) | 안 | 8 | 2
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.20 | final visit(2-4개월) | 안 | 8 | 6
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | VA | UIVA | uncorrected | Intermediate(66cm) | LogMAR | Monocular | ≤0.30 | final visit(2-4개월) | 안 | 8 | 6
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | spectacle dependence |  |  |  | Y/N |  |  |  | 안 | 57 | 5.985
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | spectacle dependence |  |  |  | Y/N |  |  |  | 안 | 17 | 2.431
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | spectacle dependence |  |  |  | Y/N |  |  |  | 안 | 15 | 5.625
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | spectacle dependence |  |  |  | Y/N |  |  |  | 안 | 12 | 0
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | spectacle dependence |  |  |  | Y/N |  |  |  | 안 | 13 | 2.886
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | Satisfaction | photic phenomena |  |  | Y/N |  |  |  | 안 | 57 | 35.91
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | Satisfaction | photic phenomena |  |  | Y/N |  |  |  | 안 | 17 | 13.073
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | Satisfaction | photic phenomena |  |  | Y/N |  |  |  | 안 | 15 | 11.25
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | Satisfaction | photic phenomena |  |  | Y/N |  |  |  | 안 | 12 | 6.672
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | Satisfaction | photic phenomena |  |  | Y/N |  |  |  | 안 | 13 | 5.421
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | Satisfaction | [5 empty] | not satisfied |  | 안 | 57 | 3.933
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | Satisfaction | [5 empty] | satisfied |  | 안 | 57 | 11.799
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  |  | Satisfaction | [5 empty] | very satisfied |  | 안 | 57 | 41.268
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | Satisfaction | [5 empty] | not satisfied |  | 안 | 17 | 0
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | Satisfaction | [5 empty] | satisfied |  | 안 | 17 | 4.25
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT150 | Satisfaction | [5 empty] | very satisfied |  | 안 | 17 | 12.75
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | Satisfaction | [5 empty] | not satisfied |  | 안 | 15 | 0
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | Satisfaction | [5 empty] | satisfied |  | 안 | 15 | 4.29
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT225 | Satisfaction | [5 empty] | very satisfied |  | 안 | 15 | 12.255
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | Satisfaction | [5 empty] | not satisfied |  | 안 | 12 | 0
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | Satisfaction | [5 empty] | satisfied |  | 안 | 12 | 2.4
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT300 | Satisfaction | [5 empty] | very satisfied |  | 안 | 12 | 9.6
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | Satisfaction | [5 empty] | not satisfied |  | 안 | 13 | 2.899
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | Satisfaction | [5 empty] | satisfied |  | 안 | 13 | 1.443
1149 | Kretz (2015) | Prospective, non-randomised multicentre clinical study | Cataract surgery | Diffractive Bifocal toric IOL | TECNIS ZMT IOL(ZMT150, ZMT225, ZMT300, ZMT400) |  |  | ZMT400 | Satisfaction | [5 empty] | very satisfied |  | 안 | 13 | 8.658
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | VA | UDVA | uncorrected | Distant | Snellen | Monocular | 6/6 or better | 3개월 | 안 | 106 | 93
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | VA | UDVA | uncorrected | Distant | Snellen | Binocular | 6/6 or better | 3개월 | 안 | 106 | 100
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | VA | UNVA | uncorrected | near(40cm) | Snellen | Monocular | 6/9 or better | 3개월 | 안 | 106 | 93
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | VA | UNVA | uncorrected | near(40cm) | Snellen | Binocular | 6/9 or better | 3개월 | 안 | 106 | 104
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | VA | UIVA | uncorrected | Intermediate(70cm) | Snellen | Monocular | 6/9 or better | 3개월 | 안 | 106 | 85
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | VA | UIVA | uncorrected | Intermediate(70cm) | Snellen | Binocular | 6/9 or better | 3개월 | 안 | 106 | 96
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | Complication | [6 empty] | 3개월 | 명 | 53 | 0
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | Satisfaction | How satisfied are you with the outcome of your procedure? |  |  |  |  | Very Satisfied | 3개월 | 명 | 53 | 39.26
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | Satisfaction | How satisfied are you with the outcome of your procedure? |  |  |  |  | Satisfied | 3개월 | 명 | 53 | 10.017
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | Satisfaction | How satisfied are you with the outcome of your procedure? |  |  |  |  | Neither | 3개월 | 명 | 53 | 3.021
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | Satisfaction | How satisfied are you with the outcome of your procedure? |  |  |  |  | Dissatisfied | 3개월 | 명 | 53 | 0
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | Satisfaction | How satisfied are you with the outcome of your procedure? |  |  |  |  | Very dissatisfied | 3개월 | 명 | 53 | 0
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | Satisfaction | How much difficulty do you have performing tasks that require good close up vision (such as cooking, fixing things around the house, sewing, using hand tools, and reading or working with a computer?) |  |  |  |  | No difficulty at all | 3개월 | 명 | 53 | 36.994
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | Satisfaction | How much difficulty do you have performing tasks that require good close up vision (such as cooking, fixing things around the house, sewing, using hand tools, and reading or working with a computer?) |  |  |  |  | A little difficulty | 3개월 | 명 | 53 | 9.01
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | Satisfaction | How much difficulty do you have performing tasks that require good close up vision (such as cooking, fixing things around the house, sewing, using hand tools, and reading or working with a computer?) |  |  |  |  | Moderate difficulty | 3개월 | 명 | 53 | 5.989
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | Satisfaction | How much difficulty do you have performing tasks that require good close up vision (such as cooking, fixing things around the house, sewing, using hand tools, and reading or working with a computer?) |  |  |  |  | A lot of difficulty | 3개월 | 명 | 53 | 1.007
1162 | Venter (2014) | Prospective consecutive case series | Presbyopia | Refractive bifocal IOL | SBL-3 |  |  |  | Satisfaction | Would you recommend this procedure to your family/friends? |  |  | Y/N |  | Yes | 3개월 | 명 | 53 | 51.993
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | Complication | Posterior capsule rupture | [5 empty] | 6개월 | 명 | 30 | 0
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | Complication | endophthalmitis | [5 empty] | 6개월 | 명 | 30 | 0
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | Complication | corneal decompensation | [5 empty] | 6개월 | 명 | 30 | 0
1220 | Mojzis (2014) | Case series | Presbyopia | Diffractive trifocal IOL | AT Lisa tri 839MP |  |  |  | Complication | Increased IOP | [5 empty] | 수술직후 | 명 | 30 | 2
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | BCVA | best spectacle corrected | Distant(4m) | Snellen | Binocular | lost>2 | 6개월 | 안 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | BCVA | best spectacle corrected | Distant(4m) | Snellen | Binocular | lost 2 | 6개월 | 안 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | BCVA | best spectacle corrected | Distant(4m) | Snellen | Binocular | lost 1 | 6개월 | 안 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | BCVA | best spectacle corrected | Distant(4m) | Snellen | Binocular | 변화없음 | 6개월 | 안 | 30 | 20.001
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | BCVA | best spectacle corrected | Distant(4m) | Snellen | Binocular | gained 1 | 6개월 | 안 | 30 | 3.999
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | BCVA | best spectacle corrected | Distant(4m) | Snellen | Binocular | gained 2 | 6개월 | 안 | 30 | 3
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | BCVA | best spectacle corrected | Distant(4m) | Snellen | Binocular | gained >2 | 6개월 | 안 | 30 | 3
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | BCVA | best spectacle corrected | Distant(4m) | Snellen | Binocular | lost>2 | 6개월 | 안 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | BCVA | best spectacle corrected | Distant(4m) | Snellen | Binocular | lost 2 | 6개월 | 안 | 30 | 2.001
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | BCVA | best spectacle corrected | Distant(4m) | Snellen | Binocular | lost 1 | 6개월 | 안 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | BCVA | best spectacle corrected | Distant(4m) | Snellen | Binocular | 변화없음 | 6개월 | 안 | 30 | 21
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | BCVA | best spectacle corrected | Distant(4m) | Snellen | Binocular | gained 1 | 6개월 | 안 | 30 | 3.999
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | BCVA | best spectacle corrected | Distant(4m) | Snellen | Binocular | gained 2 | 6개월 | 안 | 30 | 3
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | BCVA | best spectacle corrected | Distant(4m) | Snellen | Binocular | gained >2 | 6개월 | 안 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | BCNVA | best spectacle corrected | near(40cm) | Snellen | Binocular | lost>2 | 6개월 | 안 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | BCNVA | best spectacle corrected | near(40cm) | Snellen | Binocular | lost 2 | 6개월 | 안 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | BCNVA | best spectacle corrected | near(40cm) | Snellen | Binocular | lost 1 | 6개월 | 안 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | BCNVA | best spectacle corrected | near(40cm) | Snellen | Binocular | 변화없음 | 6개월 | 안 | 30 | 18
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | BCNVA | best spectacle corrected | near(40cm) | Snellen | Binocular | gained 1 | 6개월 | 안 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | BCNVA | best spectacle corrected | near(40cm) | Snellen | Binocular | gained 2 | 6개월 | 안 | 30 | 12
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | myopic | VA | BCNVA | best spectacle corrected | near(40cm) | Snellen | Binocular | gained >2 | 6개월 | 안 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | BCNVA | best spectacle corrected | near(40cm) | Snellen | Binocular | lost>2 | 6개월 | 안 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | BCNVA | best spectacle corrected | near(40cm) | Snellen | Binocular | lost 2 | 6개월 | 안 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | BCNVA | best spectacle corrected | near(40cm) | Snellen | Binocular | lost 1 | 6개월 | 안 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | BCNVA | best spectacle corrected | near(40cm) | Snellen | Binocular | 변화없음 | 6개월 | 안 | 30 | 23.01
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | BCNVA | best spectacle corrected | near(40cm) | Snellen | Binocular | gained 1 | 6개월 | 안 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | BCNVA | best spectacle corrected | near(40cm) | Snellen | Binocular | gained 2 | 6개월 | 안 | 30 | 6.99
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  | hyperopic | VA | BCNVA | best spectacle corrected | near(40cm) | Snellen | Binocular | gained >2 | 6개월 | 안 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  |  | Complication | Corneal edema | [5 empty] | 6개월 | 명 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  |  | Complication | pupillary block | [5 empty] | 6개월 | 명 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  |  | Complication | RD | [5 empty] | 6개월 | 명 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  |  | Complication | endophthalmitis | [5 empty] | 6개월 | 명 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  |  | Complication | retreatment | [5 empty] | 6개월 | 명 | 30 | 0
1350 | Ferrer-Blasco (2012) | Prospective study | Presbyopia | Diffractive multifocal IOL | AcrySof ® ReSTOR ® SN6AD3 |  |  |  | Complication | posterior capsulotomy | [5 empty] | 6개월 | 명 | 30 | 0

## 비뚤림 위험평가(RoB)_IOL
4. 비뚤림위험평가
L: 낮음, H: 높음, U: 불확실, N: 해당없음
no. | 1저자 (연도) | 연구설계 | 대상자 | 1. 무작위 배정순서 생성 | 1. 판단근거 | 2. 배정순서 은폐 | 2. 판단근거 | 3. 연구참여자, 연구자에 대한 눈가림 | 3. 판단근거 | 4. 결과평가에 대한 눈가림 | 4. 판단근거 | 5. 불충분한 결과자료 | 5. 판단근거 | 6. 선택적 보고 | 6. 판단근거 | 7. 민간연구지원 비뚤림 | 7. 판단근거
323 | Gundersen (2016) | Randomized prospective study | Cataract surgery | U | This study was a single-site, single-surgeon, randomized prospective study | U | This study was a single-site, single-surgeon, randomized prospective study | L | 눈가림이 결과에 영향을 미치지 않을 것으로 판단됨 | L | 눈가림이 결과에 영향을 미치지 않을 것으로 판단됨 | L | 결측치 없음 | H | 불완전한 결과보고로 인해 메타분석에 포함시킬 수 없는 경우 | H | The study was funded with a grant from FineVision, Liège, Belgium
365 | Monaco (2017) | Prospective randomized double-blind controlled clinical study | Cataract surgery | L | Randomization was performed using Minim, a free minimization program that incorporates a single factor (ie, the surgeon). | U | 언급없음 | L | This study was double masked. Patients were not told which IOL was implanted until all data had been collected and study recorded at the outcome assessment. Examiners performing refractive outcome assessments were blinded to which IOL was implanted to minimize a risk for bias caused by IOL identification during examinations. Statisticians were also masked to IOL identity until after the primary outcome data were analyzed. | L | Examiners performing refractive outcome assessments were blinded to which IOL was implanted to minimize a risk for bias caused by IOL identification during examinations. Statisticians were also masked to IOL identity until after the primary outcome data were analyzed. | L | 결측치 없음 | L | 불완전한 결과보고로 인해 메타분석에 포함시킬 수 없는 경우 | L | None of the authors has a financial or proprietary interest in any material or method mentioned.
391 | Alio (2011) | Prospective randomized controlled clinical study | Presbyopia | U | This multicenter prospective randomized controlled clinical study | U | Each patient was randomly assigned to 1 | L | 눈가림이 결과에 영향을 미치지 않을 것으로 판단됨 | L | 눈가림이 결과에 영향을 미치지 않을 것으로 판단됨 | L | 결측치 없음 | L | 방법에 기술된방법대로 결과보함 | L | Supported in part by a grant from the Spanish Ministry of Health, Instituto Carlos III, Red Tem atica de Investigaci on Cooperativa en Salud Patolog ıa Ocular del Envejecimiento, Calidad Visual y Calidad de Vida, Subproyecto de Calidad Visual (RD07/0062), and a research grant from the European Society of Cataract & Refractive Surgeons.

## 비뚤림 위험평가 (RoBANS)_IOL
4. 비뚤림위험평가
L: 낮음, H: 높음, U: 불확실, N: 해당없음
no. | 1저자 (연도) | 연구설계 | 질환 | 1. 대상군 비교가능성 | 1. 판단근거 | 2. 대상군 선정 | 2. 판단근거 | 3. 교란변수 | 3. 판단근거 | 4. 노출측정 | 4. 판단근거 | 5. 평가자의 눈가림(시력) | 5. 판단근거 | 6. 결과평가 | 6. 판단근거 | 7. 불완전한 결과자료 | 7. 판단근거 | 8. 선택적 결과보고(시력) | 8. 판단근거 | 9. 민간연구지원 비뚤림 | 9. 판단근거
10_국내 | Kretz (2016) | Prospective study | 백내장, 노안, RLE에 적절한 pre-presbyopia 환자 | L | 중재에 대한 노출전후의 인구집단이 동일함 | U | 전향적 수집이나 연속적으로 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | U | 환자만족도 조사도구가 신뢰도와 타당도 입증여부 확인불가함 | L | 탈락률에 대한 기술없음 | L | 방법에 기술된 결과를 포함함 | L | This study was supported by a resarch grant form the Klaus Tschira Foundation
28_국내 | 곽준영 (2012) | Prospective study | 백내장수술 | L | 중재에 대한 노출전후의 인구집단이 동일함 | U | 전향적 수집이나 연속적으로 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | L | 객관적 방법을 사용하여 결과평가함 | L | 탈락률에 대한 기술없음 | L | 방법에 기술된 결과를 포함함 | U | 관련기술없음
46_국내 | 최문정 (2020) | Prospective study | 백내장수술 | L | 중재에 대한 노출전후의 인구집단이 동일함 | U | 전향적 수집이나 연속적으로 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | L | 객관적 방법을 사용하여 결과평가함 | U | 탈락률 4%, 탈락자와 완료자의 기저상태에 대한 보고가 없음 | L | 방법에 기술된 결과를 포함함 | L | This research was supported by a grant from the Korea Health Technology R&D Project through the Korea Health Industry Development Institute (KHIDI) funded by the Ministry of Health & Welfare, Republic of Korea (grant no. HI18C1111), and by the Basic Science Research Program through the National Research Foundation of Korea (NRF) funded by the Ministry of Education, Science and Technology (NRF-2019R1F1A1062468).
71_국내 | Kretz (2016)b | Prospective non-comparative non-randomized clinical study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | U | 전향적 수집이나 연속적으로 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | U | 환자만족도 조사도구가 신뢰도와 타당도 입증여부 확인불가함 | L | 탈락률에 대한 기술없음 | L | 방법에 기술된 결과를 포함함 | L | The David J Apple International Laboratory for Ocular Pathology received research grants from the Klaus Tschira Foundation.
78_국내 | 권영기 (2015) | Retrospective study | 백내장수술 | L | 중재에 대한 노출전후의 인구집단이 동일함 | U | 전향적 수집이나 연속적으로 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | U | 환자만족도 조사도구가 신뢰도와 타당도 입증여부 확인불가함 | U | 환자만족도에서 2명 탈락, 탈락자와 완료자의 차이에 대한 설명없음 | L | 방법에 기술된 결과를 포함함 | U | 관련기술없음
198_국내 | 박율리 (2018) | 전향, 무작위배정 | 백내장수술 | L | 무작위배정 | L | 참여자 모집전략이 대상군 모두 동일함 | U | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | U | 환자만족도 조사도구가 신뢰도와 타당도 입증여부 확인불가함 | L | 탈락률에 대한 기술없음 | H | 두 시술의 결과를 모두 기술하지 않은 경우있음 | U | 관련기술없음
59 | Alfonso (2010) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | L | 전향적 수집, 연속적 수집 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | L | 선행연구에서 사용된 설문문항 이용 | L | 탈락률에 대한 기술없음 | L | 방법에 기술된 결과를 포함함 | L | The authors have no proprietary interest in the materials presented herein
272 | Pedrotti (2017) | Prospective study | Cataract surgery | L | 군간 질병중등도 등에 차이가 없어 비교할만한 인구집단임 | L | 전향적 수집, 연속적 수집, 참여자 모집전략이 대상군 모두 동일함 | U | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | L | 객관적 방법을 사용하여 결과평가함 | L | 탈락률 없음 | L | 불완전한 결과보로 인해 메타분석에 포함시킬 수 없는 경우 | L | The authors declare that they have no conflict of interest.
442 | Kohnen (2020) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | U | 전향적 수집이나 연속적으로 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | L | 객관적 방법을 사용하여 결과평가함 | L | 탈락률 3% | L | 방법에 기술된 결과를 포함함 | H | Sponsored by Alcon Research LLC (Fort Worth, Texas, USA). The sponsor participated in study design; data management, analysis and interpretation; and manuscript preparation, review, and approval. Medical writing support was provided by Natalia Zhukovskaya, PhD, of ICON plc (North Wales, Pennsylvania, USA), and was funded by Alcon Laboratories, Inc.
469 | Pedrotti (2020) | Prospective study | Cataract surgery | L | Twenty- five consecutive patients were prospectively recruited the day after the implantation of the second eye with the Tecnis Symfony IOL (Symfony group) and 25 consecutive patients were prospectively enrolled the day after the implantation of the second eye with the Acrysoft IQ PanOptix IOL (PanOptix group), being careful to enrol patients who respected inclusion criteria and in order to create two homogeneous group for age and ocular parameters with a randomisation of 1:1 | L | Twenty- five consecutive patients were prospectively recruited the day after the implantation of the second eye with the Tecnis Symfony IOL (Symfony group) and 25 consecutive patients were prospectively enrolled the day after the implantation of the second eye with the Acrysoft IQ PanOptix IOL (PanOptix group), being careful to enrol patients who respected inclusion criteria and in order to create two homogeneous group for age and ocular parameters with a randomisation of 1:1 | U | Twenty- five consecutive patients were prospectively recruited the day after the implantation of the second eye with the Tecnis Symfony IOL (Symfony group) and 25 consecutive patients were prospectively enrolled the day after the implantation of the second eye with the Acrysoft IQ PanOptix IOL (PanOptix group), being careful to enrol patients who respected inclusion criteria and in order to create two homogeneous group for age and ocular parameters with a randomisation of 1:1 | L | 수술기록으로 노출확인 가능함 | L | Only the investigators carrying out the follow-­ups were masked to the implanted IOLs. Neither surgery treatment nor outcomes were masked. | L | 객관적 방법을 사용하여 결과평가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | L | This study was funded by Università degli Studi di Verona.
471 | Kohnen (2020)b | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | U | 전향적 수집이나 연속적으로 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | L | 객관적 방법을 사용하여 결과평가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | H | Supported by a research grant from Alcon Research LLC, Fort Worth, Texas.
488 | Alio (2020) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | L | 전향적 수집, 연속적 수집 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | L | 객관적 방법을 사용하여 결과평가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | L | Network for Cooperative Research in Health “OFTARED” – Reference: RD16/ 0008/0012. Funded by Instituto de Salud Carlos III and co-funded by European Regional Development Fund (ERDF), “A way to make Europe”.
513 | Power (2018) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | U | 전향적 수집이나 연속적으로 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | L | 객관적 방법을 사용하여 결과평가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | L | The authors declare no conflict of interest, financial or otherwise.
536 | Escandon-Garcia (2018) | Prospective study | Cataract surgery | L | 군간 질병중등도 등에 차이가 없어 비교할만한 인구집단임 | L | 참여자 모집전략이 대상군 모두 동일함 | U | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | L | 객관적 방법을 사용하여 결과평가함 | L | 탈락률 없음 | L | 불완전한 결과보로 인해 메타분석에 포함시킬 수 없는 경우 | L | This study has been funded by FEDER through the COMPTETE Program and by the Portuguese Foundation for Science and Technology (FCT) in the framework of projects PTDC/SAU-BEB/098391/2008 and PTDC/SAUBEB/098392/2008 and the Strategic Project PEST-C/FIS/ UI607/2011.
621 | Friedrich (2012) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | U | 전향적 수집이나 연속적으로 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | L | 객관적 방법을 사용하여 결과평가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | L | The authors have no proprietary or commercial interest in the medical devices that are involved in this work. This research received no specific grant from any funding agency in the public, commercial or not-for-profit sectors.
653 | Kim (2020) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | U | 전향적 수집이나 연속적으로 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | U | 환자만족도 조사도구가 신뢰도와 타당도 입증여부 확인불가함 | L | 탈락률 없음 | H | 불완전한 결과보로 인해 메타분석에 포함시킬 수 없는 경우 | H | This study was funded by Alcon Research LLC, Fort Worth, TX, USA, and Alcon Korea Ltd., Seoul, South Korea.
666 | Song (2020) | Prospective study | Cataract surgery | L | 군간 질병중등도 등에 차이가 없어 비교할만한 인구집단임 | L | 참여자 모집전략이 대상군 모두 동일함 | U | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | Participant-and examiner-blinded cohort study | L | QoV 도구 validated | L | 탈락률 없음 | H | 불완전한 결과보로 인해 메타분석에 포함시킬 수 없는 경우 | L | This work was partially supported by the Key Research and Development Program of Zhejiang Province, China (2015C03042), and the National Natural Science Foundation of China (81570822 and 81600716).
691 | Bohm (2019) | Prospective study | Cataract surgery | L | 군간 질병중등도 등에 차이가 없어 비교할만한 인구집단임 | L | 참여자 모집전략이 대상군 모두 동일함 | U | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | L | 객관적 방법을 사용하여 결과평가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | U | Dr. Bohm is a consultant to and conducts research for Santen GmbH and Santhera Pharmaceuticals AG. Dr. Kohnen is a consultant to and conducts research for Abbott/Johnson & Johnson, Alcon/Novartis AG, Avedro Inc., Karl Zeiss Meditec AG, Oculentis GmbH, Oculus Optikger€ate GmbH, Presbia PLC, and Schwind eye-tech-solutions GmbH & Co. KG; is a consultant to Allergan, Inc., Bausch & Lomb, Inc., Dompe Farmaceutici S.p.A., Geuder AG, Me- dupdate Europe GmbH, Merck Sharp & Dohme Corp., Rayner Intraocular Lenses Ltd., Santen GmbH, STAAR Surgical AG, TearLab Corp., The a Pharma GmbH, Thieme Compliance GmbH, Uni-Med Verlag AG, and Ziemer Ophthalmology GmbH. He does research for Hoya Surgical Optics GmbH. None of the other authors has a financial or proprietary interest in any material or methods mentioned.
702 | Levinger (2019) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | L | 전향적 수집, 연속적 수집 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | U | 환자만족도 조사도구가 신뢰도와 타당도 입증여부 확인불가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | L | No conflicting relationship exists for any author.
732 | Vounotrypidis (2017) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | U | 전향적 수집이나 연속적으로 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | U | 환자만족도 조사도구가 신뢰도와 타당도 입증여부 확인불가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | L | None of the authors has a financial or proprietary interest in any material or method mentioned.
782 | Chang (2019) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | U | 전향적 수집이나 연속적으로 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | U | 환자만족도 조사도구가 신뢰도와 타당도 입증여부 확인불가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | H | The study was sponsored by a grant by Abbott Medical Optics, Inc for the cost of assessments.
1149 | Kretz (2015) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | L | 전향적 수집, 연속적 수집 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | U | 환자만족도 조사도구가 신뢰도와 타당도 입증여부 확인불가함 | U | 탈락률에 대한 정보없음 | L | 방법에 기술된 결과를 포함함 | H | he study was supported by an unrestricted grant by Abbott Medical Optics, Ettlingen, Germany.
1162 | Venter (2014) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | L | 전향적 수집, 연속적 수집 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | U | 환자만족도 조사도구가 신뢰도와 타당도 입증여부 확인불가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | L | The authors have no financial or proprietary interest in the materials presented herein.
1220 | Mojzis (2014) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | U | 전향적 수집이나 연속적으로 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | U | 환자만족도 조사도구가 신뢰도와 타당도 입증여부 확인불가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | L | No author has a financial or proprietary interest in any material or method mentioned.
1288 | Tsaousis (2013) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | U | 전향적 수집이나 연속적으로 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | L | 객관적 방법을 사용하여 결과평가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | L | The authors have no financial or proprietary interest in the materials presented herein.
1350 | Ferrer-Blasco (2012) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | L | 전향적 수집, 연속적 수집 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | L | 객관적 방법을 사용하여 결과평가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | L | This research was supported in part by Ministerio de Ciencia e Innovación Research Grants to Robert Montés-Micó (#SAF2008-01114# and #SAF2009-13342#)
1409 | Altaie (2012) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | L | 전향적 수집, 연속적 수집 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | U | 환자만족도 조사도구가 신뢰도와 타당도 입증여부 확인불가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | L | Funding sources: None declared
1485 | Alfonso (2011) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | L | 전향적 수집, 연속적 수집 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | L | 객관적 방법을 사용하여 결과평가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | L | This research was supported in part by Ministerio de Ciencia e Innovación Research Grants to Robert Montés-Micó (SAF2008-01114 and SAF2009-13342).
1544 | Alfonso (2010)b | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | L | 전향적 수집, 연속적 수집 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | L | 객관적 방법을 사용하여 결과평가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | L | This research was supported in part by a Ministerio de Ciencia e Innovación Research Grants to Robert Montés-Micó (#SAF2008-01114# and #SAF2009-13342#).
1556 | Alfonso (2010)c | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | L | 전향적 수집, 연속적 수집 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | All examinations were performed 6 months after IOL implantation by the same ophthalmic technician, who was unaware of the study’s objective | L | 객관적 방법을 사용하여 결과평가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | L | This research was supported in part by Ministerio de Ciencia e Innovación research grants to Robert Montés-Micó (SAF2008-01114 and SAF2009-13342)
3093 | Tan (2019) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | U | 전향적 수집이나 연속적으로 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | U | 환자만족도 조사도구가 신뢰도와 타당도 입증여부 확인불가함 | U | 환자만족도 조사도구가 신뢰도와 타당도 입증여부 확인불가함 | L | 방법에 기술된 결과를 포함함 | U | 관련기술없음
3220 | Almulhim (2018) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | U | 전향적 수집이나 연속적으로 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | U | 환자만족도 조사도구가 신뢰도와 타당도 입증여부 확인불가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | L | The authors declared that there is no conflict of interest.
3241 | Zamora-De-La-Cruz (2018) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | L | 전향적 수집, 연속적 수집 | U | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | U | 환자만족도 조사도구가 신뢰도와 타당도 입증여부 확인불가함 | L | 탈락률 없음 | H | 불완전한 결과보로 인해 메타분석에 포함시킬 수 없는 경우 | L | No funding has been grandted for this research and the authors declare no commercial interest in this protocol
3522 | Plaza-Puche (2016) | Prospective study | Cataract surgery | L | 나이, IOL powe에서 군간 차이가 없음 | L | 전향적 수집, 연속적 수집 | U | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | Follw-up was performed in an unmatched way by an independent optometrist specialized and certificated in good clinical practice | L | 객관적 방법을 사용하여 결과평가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | L | This publication was carried out in the framework of the Red Tematical de invetigacion cooperativa en Salud (RETICS), reference number RD 12/0034/007, finaced by the instituo carlos III-General subdirection of Networks and cooperative investigation centers and the european regisonal development fund.
4009 | Venter (2013) | Prospective study | Cataract surgery | L | 중재에 대한 노출전후의 인구집단이 동일함 | L | 전향적 수집, 연속적 수집 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | U | 환자만족도 조사도구가 신뢰도와 타당도 입증여부 확인불가함 | L | 탈락률 없음 | L | 방법에 기술된 결과를 포함함 | L | Neither author has a financial or proprietary interest in any material or method mentioned.
4240 | Van Der Linden (2012) | Prospective study | Cataract surgery | L | 군간 질병중등도 등에 차이가 없어 비교할만한 인구집단임 | L | 전향적 수집, 연속적 수집 | U | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있음 | L | 수술기록으로 노출확인 가능함 | L | 결과평가자의 눈가림 없지만 눈가림 여부가 결과측정에 영향을 미치지 않음 | U | 환자만족도 조사도구가 신뢰도와 타당도 입증여부 확인불가함 | L | 탈락률 없음 | H | 불완전한 결과보로 인해 메타분석에 포함시킬 수 없는 경우 | U | Dr. Lapid-Gortzak has received speaker’s fees from Alcon, Simovision, Oculentis, and Medtechnika. Dr. Lapid-Gortzak and Dr. van der Meulen have received an unrestricted research grant from Oculentis. Dr. van Velthoven has received speaker’s fees from Novartis (Netherlands). Drs. van der Linden, Nieuwendaal, and Mourits have no financial or proprietary interest in any material or method mentioned.

## 문헌특성_RS
1. 선택문헌의 특성
no. | 1저자 (연도) | 연구설계 (RCT, NRCT, Obs) | Trial명(NCT no.) (있는 경우) 예. | 연구국가 | 연구대상자 |  |  |  | 환자특성 | [9 empty] | 중재군(I) |  | 대조군(C) |  | 3군 |  | 4군 |  | 결과변수 (1. VA, 2.CS, 3. Satisfaction, 4. Independence, 5. Complicaiton | 최대추적관찰기간 | 저자결론
[5 empty] | 대상자 모집기간 | 대상자 | 이전시술여부 | 선택/배제기준 | 연구대상자 수: 명(안) | [6 empty] | 기저특성 |  |  | 중재군 |  | 대조군(C) |  | 대조/중재군명 |  | 대조/중재군명
[9 empty] | 총 | 비교군 수 (2군, 3군 등) | 중재군 (N명(안)) | 대조군 (N명(안)) | 3군 (N명(안)) | 4군 (N명(안)) | 탈락률(%) 전체(중재 vs. 대조) | 평균연령 | 남성 (%) | 양안시술여부(Y/N) |  | 수술방법 |  | 수술방법 |  | 수술방법 |  | 수술방법
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial |  | 캐나다 | 2004.8-2006.8 | Hyperopic presbyopia | N | 포함기준: hyperopia up to +4.00 diopters (D) sphere and astigmatism up to +1.50 D cylinder; age 45 to 70 years; and уϩ1.25 D add at 16 inches at the time of enrollment. All eyes were required to demonstrate refractive stability, defined as ≤0.50-D change in sphere or cylinder that could be confirmed by clinical records. All patients were required to have CDVA of at least 20/20 and CNVA of at least 20/25. 배제기준: All patients had a normal ocular surface on slit-lamp examination with no evidence of dry eye or lid or corneal abnormalities. Pupil size as measured by the aberrometer had to be ≥4.0 mm without pharmacologic agents, and the anticipated postoperative K-reading had to be <50.00 D. The remaining exclusion criteria followed the standard refractive surgery guidelines. | 33명(66안) | 1 | 33명(66안) | NA | NA | NA | 24.24% | 55.1 | 34% | Y | Aspheric Wavefront-guided LASIK | VISX STAR S4 or STAR S4 IR excimer laser system (Abbott Medical Optics) | NA | [5 empty] | 1, 2, 3, 4 | 12개월 | 근기능 시력을 확장하도록 설계된 aspheric ablation은 12개월동안 효과적이고 안정적임
329 | Khalifa (2011) | Prospective consecutive comparative randomized double-masked study |  | 이집트 | 2009.1-2009.7 | Myopic presbyopia | N | 포함기준: all aged >40 years with symptoms of presbyopia 배제기준: <40 years old, any abnormal or suspicious patterns on corneal topography, central pachymetric measurements ,<500 µm, and history of chronic eye diseases. Patients with lens opacities were also excluded from the study. | 26명(52안) | 2 | 15명(30안) | 11명(22안) | NA | NA | 0% | 중재: 47.46, 대조: 45.2 | 34.62% | Y | all high-order aberrations were treated | SCHWIND ESIRIS LASIK done using the Moria M2 | Vertical coma aberrations were selectively left untreated, while all other HOAs were treated | SCHWIND ESIRIS LASIK done using the Moria M2 |  |  |  |  | 1, 2 | 3개월 | HOA의 선택적 치료를 포함하여 wavefront-guided ablation으로 중등도 근시를 치료하는 것은 시각 기능을 개선하는데 매우 유용함. 모든 HOA를 치료하지만 vertical coma는 노안환자의 근거리 나안시력을 개선함
430 | Kohnen (2020) | Prospective, nonrandomized trial |  | 독일 | 2012.11-2016.6 | Presbyopia | N | 포함기준: 40세 이상, manifest presbyopia with CNVA ≤0.2 logMAR with near addition of ≥+1.5 D, refraction errors between -5.00 D and +5.00 D in SE, and astigmatism of ≤3.00 D. Patients could only be enrolled in the study if they had avoided the use of contact lenses for 2 weeks with soft contact lenses or 6 weeks with rigid contact lenses before the examination. 배제기준: Patients with ocular pathologies or impaired vision were excluded from participation as well as professionals whose jobs demand a high distance VA | 29명(58안) | 2 | 14명(28안) | 15명(30안) | NA | NA | 3.333% | 54.3 | 62.07% | Y | PresbyMAX (Hybrid micro-Monovision) | hybrid profile with less depth of focus on the dominant eye to improve distance vision without compromising on UNVA | PresbyMAX (micro-Monovision) | LASIK multifocal aspheric laser profile to induce depth of focus was used on both eyes combined with a micromonovision, aiming for emmetropia on the dominant eye and making the nondominant eye slightly myopic |  |  |  |  | 1, 3, 4 | 12개월 | PresbyMAX는 CDVA의 감소를 보였지만 대부분의 환자를 만족시킴. 근거리 시력의 증가와 이에 따라 안경의 필요성을 줄인다는 것이 이러한 감소를 보상하는 것으로 보임
457 | Xu (2020) | Prospective, non-randomized, non-comparative study |  | 중국 | 2014.4-2016.8 | Emmetropic presbyopia | U | 포함기준: at least 50 years of age, MRSE of ±0.50 D for distance vision with astigmatism ≤1.00 D in each eye, UDVA of 20/40 (logMAR 0.30) or better in each eye, CDVA of 20/25 (logMAR 0.10) or better in each eye, and presbyopia ≥1.50 D at 40 cm. 배제기준: a history of ocular trauma or prior ocular surgery except for corneal refractive surgery, ocular pathology other than cataract, history of nuclear sclerosis LOCS III grade 2 or worse or other cataracts reducing CDVA, and acute or chronic disease or illness that could increase the operative risk or confound the study outcomes. | 4명(8안) | 1 | 4명(8안) | NA |  |  | 0% | 58.5 | 0% | Y | LaserACE | [7 empty] | 1, 4, 5 | 12개월 | LaserACE 절차는 원거리 시력이나 시각적 질에 영향을 주지 않으면서 노안 교정을 시행하는데 효과적이고 안전한 기술임. 그러나 그 메커니즘과 장기적인 효과는 추가 검증이 필요함
497 | Liu (2020) | Prospective, non-comparative case series |  | 중국 | NA | Myopic presbyopia | N | 포함기준: suitable for LASIK, presbyopic with a CDVA of 20/25 or better in both eyes, tolerance of at least -0.75 D anisometropia, photopic pupil diameter smaller than 3.5 mm and mesopic pupil diameter lager than 4.5 mm [the photopic pupil data from topografic (offset) information but mesopic/scotopic from pupillometry], had stable refraction for at least 1y before the study [<0.5 D change in mean spherical equivalent (SE)], and discontinued contact lenses for at least 2wk before the preoperative evaluation 배제기준: systemic illness, previous ocular surgery, abnormal corneal topography and clinically relevant lens opacity, any signs of binocular vision anomalies at distance and near and a pupil offset of 0.7 mm or more. | 37명(74안) | 1 | 37명(74안) | NA |  |  | 8.108% | 43.8 | 29.73% | Y | PresbyMAX with micro-Monovision | PresbyLASIK using AMARIS 750S excimer laser | [6 empty] | 1, 2, 3, 5 | 12개월 | micro-Monovision이 있는 Hybrid bi-aspheric multifocal central presbyLASIK은 근시 노안을 위한 효과적인 옵션임. 비교적 젊은 노안 인구에서 1년 후 수술 결과는 높은 만족도와 함께 원거리 및 근거리 시력이 모두 개선되었음.
695 | Romero (2019) | Prospective case series |  | 스페인 | NA | Myopic presbyopia, hyperopic presbyopia | N | 포함기준: 40-55세, preoperative refraction between +0.50 D and +3.00 D or between -0.50 D and -6.00 D, preoperative astigmatism between 0.00 D and 3.00 D, no previous eye surgery, manifest presbyopia symptoms, a transparent crystalline lens, a normal preoperative examination, and good adaptation on the micromonovision tolerance test. Patients also had to meet the criteria for residual corneal thickness eye surgery, and stromal bed. 배제기준: previous eye surgery, central corneal pachymetry less than 480 mm or cases in which surgery could not be safely performed because of corneal thickness, moderate to severe dry eye or other types of corneal disease, strabismus, phoria or amblyopia that would prevent corrected binocular vision, unstable refraction over the past year, autoimmune system disease, macular degeneration or other retinopathies, glaucoma, and history of ocular inflammation. | 50명(100안) | 1 | 50명(100안) | NA |  |  | 0% | 46.8 | 42% | Y | Presbyopia correction using micro-Monovision and aspheric aberration patterns | [7 empty] | 1, 2, 5 | 6개월 | micro-Monovision 및 excimer laser를 이용한 aspheric aberration ablation에 의한 노안교정은 입체시, 대비감도, 광학품질의 변화를 최소화함
734 | Taneri (2019) | Prospective, non-randomized, comparative case series |  | 독일 | NA | Hyperopic presbyopia |  | 포함기준: Hyperopic patients older than 47 years with CDVA of 20/25 or better in each eye were included in this study. Distance refraction had to be stable (< 0.50 diopters [D] change in mean spherical equivalent) for at least 1 year prior to the treatment with astigmatism between 0.00 and 2.50 D. A maximum difference between manifest and cycloplegic spherical equivalent of 0.50 D was allowed. All eyes were required to have normal keratometry and topography and a preoperative central corneal thickness of more than 500 µm. Angle kappa had to be less than 10° in the Orbscan examination. 배제기준: Patients who suffered from systemic illness or had previous ocular surgery were excluded from the study. Additional exclusion criteria were clinically relevant lens opacities or any signs of binocular vision anomalies, including (relative) amblyopia at distance or near. | 15명(30안) | 2 | 8명(16안) | 7명(14안) |  |  | 20% | 중앙값 중재: 52, 대조: 53 | 26.67% | Y | Varifocal LASIK | SUPRACOR | Monofocal LASIK | [5 empty] | 1 | 3개월 | 목표굴절로 정시를 사용하여 varifocal ablations은 원시노안 LASIK에서 monofocal ablation과 비교하여 추가 이점을 제공하지 않았음. 저자들이 epithelial remodeling이 varifocal ablation의 패턴의 영향을 가리고 근시수술 후 굴절(modified monovision)이 근거리 및 중간시력을 더욱 향상시킬 필요가 있음 이러한 결과는 노안교정을 평가하는 연구에서 대조군의 가치를 보여줌
897 | Pajic (2017) | Prospective, single-surgeon study |  | 스위스 | NA | Myopic presbyopia | NA | NA | 36명(72안) | 1 | 36명(72안) |  |  |  | 0% | NA | NA | Y | multifocal presbyLASIK | SUPRACOR refractive surgery | [6 empty] | 1, 5 | 6개월 | SUPRACOR가 노안에 대해 안전하고 신뢰할 수 있는 굴절수술 치료법임
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study |  | 스위스 | 2013-2014 | Hyperopic presbyopia |  | 포함기준: age ≥47 years, near addition ≥ 1.75 D, maximal difference between mean refractive spherical equivalent and cycloplegic refraction 0.75 D, CDVA ≥0.8, mean refractive sperical equivalent +1.0 to +4.0 and ≤2 D astigmatism. 배제기준: patients with irregular pattern of corneal topograpy, history of ocular disease, ectopic pupil, and signs of binocular visin anomalies were excluded | 20명(39안) | 1 | 20명(39안) |  |  |  | 0% | 59 | 50% | Y | SUPRACOR | Multifocal central presbyopic LASIK using femtosecond laser for flap creation | [6 empty] | 1, 2, 3, 4 | 12개월 | SUPRACOR 이후 1년시점에서 UDVA, UNVA는 크게 개선되었고 photopic contrast sensitivity은 정상 연령상관범위에 있었고 대부분의 환자(75%)가 결과에 만족함. 원거리 시력상실, 시력변동, 눈부심 증가와 같은 부작용은 spectacle-free near vision의 개선보다 환자만족도에 더 큰 영향을 미침
958 | Wang Yin (2016) | Prospective cohort study |  | 프랑스 | 2013.2-2015.1 | Hyperopic presbyopia | N | 포함기준: 40-65세, spherical refractive error between 0.0 D and +6.0D, corneal cylinder up to 3.0D, preoperative CDVA of 20/20 or better, central corneal thickness more than 500 μm, and no other ocular pathology or previous ocular surgery | 69명(138안) | 1 | 69명(138안) |  |  |  |  | 53.84 | 40.58% | Y | central presbyopic LASIK with Custom Q algorithm | wavelight EX500 | [6 empty] | 1, 3 | 12개월 | Monovision correctin algorithm을 이용한 corneal asphericity modulation이 가능한 central presbyopic LASIK 은 노안치료에 효과적이고 안전함
963 | Zhang (2016) | Prospective, non-comparative case series |  | 중국 | NA | Myopic presbyopia | N | 포함기준: CDVA of 20/25 or better in either eye. 제외기준: previous refractive surgery, corneal and lens opacities that may affect vision, acute or chronic systemic disease, or any kind of immunosuppressive disorder. | 40명(80안) | 1 | 40명(80안) |  |  |  | 0 | 43.4 | NA | Y | Aspheric Micro-monovision LASIK (Laser Blended Vision) | [7 empty] | 1, 2, 3, 4 | 3개월 | aspheric micro-monovision protocol 은 myopic astigmatism를 치료하고 동시에 노안증상을 완화하는데 안전하고 효과적인 수단임
1123 | Soler Tomas (2015) | Prospective, longitudinal case series |  | 스페인 | NA | Hyperopic presbyopia | N | 포함기준: Hyperopic presbyopes seeking independence from reading glasses who understood and signed an informed consent form, +1 to +2.5 D mean refractive spherical equivalent (SE), up to 1-D astigmatism, maximum difference between subjective and cycloplegic refraction ≤ +0.5, and tolerance to the simulation test of decreased distance vision. | 30명(40안) | 2 | 14명(21안) | 16명(19안) |  |  | 0 | 중재: 51.9, 대조: 53.5 | 63.33% | N | asymmetrical presbyLASIK |  | symmetrical presbyLASIK | [5 empty] | 1, 3 | 18개월 | Symmetrical, asymmetrical presbyLASIK은 수술 18개월 후 원거리 UCVA, 근거리 UCVA를 유의하게 개선시킴
1225 | Cosar (2014) | Prospective, nonrandomized, clinical trial |  | 터키 | 2011.7-2012.9 | Hyperopic presbyopia | N | 포함기준: presbyopia of at least +1.50 D with hyperopia up to +4.00 D, cylinder up to +2.00 D, CDVA of at least 20/20, and Scheimpflug corneal thickness of at least 500 μm (thinnest point). 배제기준: unstable refraction, previous ocular surgery, suspicion of keratoconus, any ocular disease, and systemic conditions that could alter the woundhealing process such as diabetes and connective tissue disorders. | 68명(123안) | 1 | 68명(123안) |  |  |  |  | 57.3 | 50% | N | SUPRACOR | ablation was performed using the Technolas excimer laser 217P using the Supracor treatment algorithm | [6 empty] | 1, 5 | 6개월 | SUPRACOR 노안시술은 6개월 추적관찰에서 근거리 시력결과가 좋았으나 CDVA 감소가 39.1 %의 눈에서 발생함
1268 | Ryan (2013) | Prospective case series |  | 아일랜드 | NA | Hyperopic presbyopia |  | 포함기준: manifest refraction spherical equivalent (MRSE) between +0.75 diopter (D) and +3.25 D, 2.00 D or less of astigmatism, a maximum difference between the MRSE and the cycloplegic refraction of 0.75 D, mean keratometry readings of 41.0 to 45.0 D, age 47 years or older, ablility to tolerate a near addition of +1.75 D, CDVA in each eye of 0.2 logMAR or better, central corneal thickness of 500 μm or more, and a mesopic pupil of 3.0 to 6.0mm 배제기준: presence of ocular surface disease, abnormal corneal topography, less than 5.0mm wetting on Schirmer testing over 5 min. without topical anesthesia and clinically singificant lenticular opacity | 23명(46안) | 1 | 23명(46안) |  |  |  |  | 57 | 17.39% | Y | SUPRACOR | SUPRACOR multifocal excimer LASIK procedure | [6 empty] | 1, 3, 4, 5 | 6개월 | SUPRACOR는 근거리 시력에 대한 높은 수준의 안경 독립성을 제공하며, UDVA은 재시술을 줄이기 위해 nomogram 조정을 개선되어야 함. 우세안의 정시와 비우세안의 0.50 D근시를 목표로하는 것이 효과적인 해결책이 될 수 있음
1309 | Luger (2013) | Prospective study |  | NA | NA | Presbyopia | N | 포함기준: Patients included in the study had manifest spherical refractive error ranging from -7.00 to +3.25 diopters (D) with up to 3.00 D of astigmatism, with presbyopic adds of up to +2.75 D. Patients were enrolled in the study if they had CDVA of 20/25 or better using the ETDRS chart, had stable refraction for 1 year before the study, and discontinued contact lenses for at least 2 to 4 weeks before the preoperative evaluation. Patients were required to have normal keratometry and topography. For comparative analyses, all aberrations were analyzed for a diameter of 6 mm. 배제기준: Patients suffering from systemic illness who had a calculated postoperative corneal bed thickness less than 300 mm after ablation, had preoperative central corneal thickness of less than 500 mm, had previous ocular surgery, or had abnormal corneal topography were excluded from the study. | 33명(66안) | 1 | 33명(66안) |  |  |  | 6.061% | 53 | NA | Y | Schwind PresbyMAX | Presby aberration-free mode using FemtoLASIK for Sphere from -7.00 to +3.25 diopters (D), astigmatism up to 3.00 D, and addition up to +2.75 D | [6 empty] | 1, 2 | 1년 | 환자 선택 및 기대 관리는 환자 만족도를 달성하는 데 필수적임. 광학적으로 결과가 예측 가능하고 좋더라도 일부 환자는 타협이 어려우며 다른 일부 환자는 근거리 시력의 사소한 손실로 불만족함. 특정 환자군이 PresbyMAX에 가장 적합함. 약간 defocused된 이미지를 생성하는 다초점 콘택트 렌즈 또는 시험 프레임을 사용한 테스트를 통해 수술 후 시각적 인상을 시뮬레이션하고 환자 수용성을 확인할 수 있음
1310 | Oh (2013) | Prospective comparative study |  | 한국 | 2007.6.-2010.12 | Presbyopia | N | 포함기준: We enrolled all qualified presbyopic patients who had a spherical equivalent manifest refraction that ranged between 0.75 and -0.75 D. 배제기준: pseudophakia; previous laser in situ keratomileusis, photorefractive keratectomy, or other refractive surgery; keratoconus or other corneal ectasia; and any other corneal opacifying disease. | 23명(46안) | 2 | 12명(24안) | 11명(22안) |  |  |  | 52.8 | 47.83% | Y | Aspheric corneal ablation with the CST | aspheric corneal ablation while preserving the central cornea (CST group) by amniotic membrane patching to protect the center of the cornea and encourage early wound healing | conventional presbyopic excimer laser (NCST) | [5 empty] | 1, 2 | 2년 | CST를 이용한 레이저각막절제술은 근시와 원거리 시력에서 더 빠른 상처회복과 더 긴 효과유지를 보여 정시 노안환자를 치료하는데 효과적일 수 있음
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation |  | 독일 | NA | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia |  | 포함기준: sphere within −4.0 and +4.0 diopters and an astigmatism equal to or smaller than 2 diopters. 제외기준: irregular patterns of corneal topography or history of ocular disease, clinically-relevant lens opacity, an ectopic pupil more than 1 mm off-centre and any signs of binocular vision anomalies at distance and near | 30명(60안) | 1 | 30명(60안) |  |  |  |  | 원시 54세, 정시 52세, 근시 51세 | 30% | Y | PresbyMAX | simultaneous correction of presbyopia and ametropia by a bi-aspheric cornea modulation technique | [6 empty] | 1, 2 | 6개월 | 백내장 증상이 없지만 굴절 이상이 없는 노안 환자에서 PresbyMAX는 노안 증상을 감소시키고 원거리 굴절을 교정하여 대부분의 환자에서 일상 생활에서 안경없는 시력을 제공함

## 결과지표_연속형_RS
2. 연속형 결과변수
no. | 1저자 (연도) | 연구설계 | 대상자 | 중재군명 | 대조군명 | 하위그룹 | 결과그룹 | 결과지표명 | 시력구분(Uncorrected, corrected) | 시력구분(Distant, Intermediate, Near) | 단위(LogMAR, Snellen, et al.,) | 안구 (Binocular, Monocular) | 비고 | 측정시점 (개월) | 분석단위(명/안) | 중재군 |  |  | 대조군 |  |  | 최종-기저평균차(변화량) |  |  |  | 두 군간 차이 | 비고
[16 empty] | N | mean | SD | N | mean | SD | 중재군 | 변화량 SD | 대조군 | 변화량 SD | P value
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | #N/A | Aspheric Wavefront-guided LASIK |  |  | VA | CDVA | Corrected | Distant | Snellen |  |  | 시술전 | 안 | 66 | 20/15 | 3 line
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | DCNVA | Distance corrected | Near(40cm) | Jaeger notation |  |  | 시술전 | 안 | 66 | J5 | 1 line
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | CDVA | Corrected | Distant | Snellen |  |  | 6개월 | 안 | 60 | 20/20 +1 | 1 line
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | CNVA | Corrected | Near(40cm) | Snellen |  |  | 6개월 | 안 | 60 | 20/20 +1 | 1 line
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | DCNVA | Distance corrected | Near(40cm) | Snellen | Binocular |  | 6개월 | 명 | 30 | 2.7 | 1.7
329 | Khalifa (2011) | Prospective consecutive comparative randomized double-masked study | Myopic presbyopia | all high-order aberrations were treated | Vertical coma aberrations were selectively left untreated, while all other HOAs were treated |  | VA | BCNVA | Best-corrected | Near(40cm) | Decimal notation |  |  | 수술전 | 안 | 30 | 0.9 | NA | 22 | 0.9 | NA
329 | Khalifa (2011) | Prospective consecutive comparative randomized double-masked study | Myopic presbyopia | all high-order aberrations were treated | Vertical coma aberrations were selectively left untreated, while all other HOAs were treated |  | VA | BCNVA | Best-corrected | Near(40cm) | Decimal notation |  |  | 3개월 | 안 | 30 | 0.8 | NA | 22 | 1 | NA
329 | Khalifa (2011) | Prospective consecutive comparative randomized double-masked study | Myopic presbyopia | all high-order aberrations were treated | Vertical coma aberrations were selectively left untreated, while all other HOAs were treated |  | VA | UNVA | Uncorrected | Near(40cm) | Decimal notation |  |  | 수술전 | 안 | 30 | 0.8 | NA | 22 | 0.8 | NA
329 | Khalifa (2011) | Prospective consecutive comparative randomized double-masked study | Myopic presbyopia | all high-order aberrations were treated | Vertical coma aberrations were selectively left untreated, while all other HOAs were treated |  | VA | UNVA | Uncorrected | Near(40cm) | Decimal notation |  |  | 3개월 | 안 | 30 | 0.3 | NA | 22 | 0.7 | NA |  |  |  |  | <0.001 | p값으로 SD 추정
329 | Khalifa (2011) | Prospective consecutive comparative randomized double-masked study | Myopic presbyopia | all high-order aberrations were treated | Vertical coma aberrations were selectively left untreated, while all other HOAs were treated |  | VA | DCNVA | Distance corrected | Near(40cm) | Decimal notation |  |  | 수술전 | 안 | 30 | 0.3 | NA | 22 | 0.3 | NA
329 | Khalifa (2011) | Prospective consecutive comparative randomized double-masked study | Myopic presbyopia | all high-order aberrations were treated | Vertical coma aberrations were selectively left untreated, while all other HOAs were treated |  | VA | DCNVA | Distance corrected | Near(40cm) | Decimal notation |  |  | 3개월 | 안 | 30 | 0.3 | NA | 22 | 0.6 | NA
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 수술전 | 안 | 58 | 0.35 | 0.35
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | Corrected | Distant | LogMAR | Binocular |  | 수술전 | 안 | 58 | -0.07 | 0.13
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Binocular |  | 수술전 | 안 | 58 | 0.53 | 0.27
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Binocular |  | 수술전 | 안 | 58 | 0.51 | 0.14
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 수술전 | 안 | 58 | 0.14 | 0.32
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | Corrected | Distant | LogMAR | Binocular |  | 수술전 | 안 | 58 | -0.16 | 0.1
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Binocular |  | 수술전 | 안 | 58 | 0.48 | 0.26
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Binocular |  | 수술전 | 안 | 58 | 0.47 | 0.13
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCIVA | Distance corrected | Intermediate | LogMAR | Binocular |  | 수술전 | 안 | 58 | 0.24 | 0.14
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 28 | 0.42 | 0.37 | 30 | 0.29 | 0.31 |  |  |  |  | >.05
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | CDVA | Corrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 28 | -0.08 | 0.13 | 30 | -0.07 | 0.12 |  |  |  |  | >.05
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular |  | 수술전 | 안 | 28 | 0.48 | 0.3 | 30 | 0.58 | 0.24 |  |  |  |  | >.05
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Monocular |  | 수술전 | 안 | 28 | 0.47 | 0.16 | 30 | 0.55 | 0.11 |  |  |  |  | >.05
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 28 | 0.12 | 0.36 | 30 | 0.16 | 0.27 |  |  |  |  | >.05
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | CDVA | Corrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 28 | -0.15 | 0.11 | 30 | -0.16 | 0.09 |  |  |  |  | >.05
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular |  | 수술전 | 안 | 28 | 0.46 | 0.26 | 30 | 0.5 | 0.25 |  |  |  |  | >.05
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Monocular |  | 수술전 | 안 | 28 | 0.43 | 0.12 | 30 | 0.51 | 0.12 |  |  |  |  | >.05
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular | dominant/distant eye | 1개월 | 안 | 58 | 0.12 | 0.19
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | Corrected | Distant | LogMAR | Monocular | dominant/distant eye | 1개월 | 안 | 58 | 0 | 0.12
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular | dominant/distant eye | 1개월 | 안 | 58 | 0.27 | 0.17
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Monocular | dominant/distant eye | 1개월 | 안 | 58 | 0.34 | 0.18
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular | nondominant/near eye | 1개월 | 안 | 58 | 0.29 | 0.18
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | Corrected | Distant | LogMAR | Monocular | nondominant/near eye | 1개월 | 안 | 58 | 0.05 | 0.13
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular | nondominant/near eye | 1개월 | 안 | 58 | 0.16 | 0.17
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Monocular | nondominant/near eye | 1개월 | 안 | 58 | 0.31 | 0.18
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular | dominant/distant eye | 1개월 | 안 | 28 | 0.09 | 0.19 | 30 | 0.15 | 0.19
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | CDVA | Corrected | Distant | LogMAR | Monocular | dominant/distant eye | 1개월 | 안 | 28 | -0.06 | 0.05 | 30 | 0 | 0.12
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular | dominant/distant eye | 1개월 | 안 | 28 | 0.25 | 0.18 | 30 | 0.29 | 0.17
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Monocular | dominant/distant eye | 1개월 | 안 | 28 | 0.36 | 0.17 | 30 | 0.33 | 0.2
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular | nondominant/near eye | 1개월 | 안 | 28 | 0.28 | 0.2 | 30 | 0.31 | 0.17
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | CDVA | Corrected | Distant | LogMAR | Monocular | nondominant/near eye | 1개월 | 안 | 28 | 0.22 | 0.18 | 30 | 0.06 | 0.11
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular | nondominant/near eye | 1개월 | 안 | 28 | 0.08 | 0.11 | 30 | 0.22 | 0.18
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Monocular | nondominant/near eye | 1개월 | 안 | 28 | 0.22 | 0.18 | 30 | 0.38 | 0.15
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular | dominant/distant eye | 3개월 | 안 | 58 | 0.21 | 0.23
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | Corrected | Distant | LogMAR | Monocular | dominant/distant eye | 3개월 | 안 | 58 | 0.05 | 0.18
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular | dominant/distant eye | 3개월 | 안 | 58 | 0.23 | 0.16
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Monocular | dominant/distant eye | 3개월 | 안 | 58 | 0.34 | 0.14
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular | nondominant/near eye | 3개월 | 안 | 58 | 0.31 | 0.18
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | Corrected | Distant | LogMAR | Monocular | nondominant/near eye | 3개월 | 안 | 58 | 0.06 | 0.17
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular | nondominant/near eye | 3개월 | 안 | 58 | 0.14 | 0.15
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Monocular | nondominant/near eye | 3개월 | 안 | 58 | 0.36 | 0.18
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 3개월 | 안 | 58 | 0.15 | 0.2
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | Corrected | Distant | LogMAR | Binocular |  | 3개월 | 안 | 58 | 0 | 0.12
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Binocular |  | 3개월 | 안 | 58 | 0.09 | 0.94
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Binocular |  | 3개월 | 안 | 58 | 0.24 | 0.13
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular | dominant/distant eye | 3개월 | 안 | 28 | 0.19 | 0.23 | 30 | 0.23 | 0.23
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | CDVA | Corrected | Distant | LogMAR | Monocular | dominant/distant eye | 3개월 | 안 | 28 | 0 | 0.13 | 30 | 0.09 | 0.2
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular | dominant/distant eye | 3개월 | 안 | 28 | 0.25 | 0.19 | 30 | 0.21 | 0.13
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Monocular | dominant/distant eye | 3개월 | 안 | 28 | 0.35 | 0.16 | 30 | 0.33 | 0.11
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular | nondominant/near eye | 3개월 | 안 | 28 | 0.28 | 0.2 | 30 | 0.33 | 0.18
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | CDVA | Corrected | Distant | LogMAR | Monocular | nondominant/near eye | 3개월 | 안 | 28 | 0.05 | 0.17 | 30 | 0.06 | 0.16
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular | nondominant/near eye | 3개월 | 안 | 28 | 0.07 | 0.13 | 30 | 0.19 | 0.15
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Monocular | nondominant/near eye | 3개월 | 안 | 28 | 0.28 | 0.2 | 30 | 0.41 | 0.15
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 3개월 | 안 | 28 | 0.12 | 0.19 | 30 | 0.17 | 0.17
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | CDVA | Corrected | Distant | LogMAR | Binocular |  | 3개월 | 안 | 28 | -0.03 | 0.09 | 30 | 0.02 | 0.14
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Binocular |  | 3개월 | 안 | 28 | 0.08 | 0.11 | 30 | 0.1 | 0.07
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Binocular |  | 3개월 | 안 | 28 | 0.23 | 0.14 | 30 | 0.26 | 0.12
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular | dominant/distant eye | 6개월 | 안 | 58 | 0.14 | 0.21
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | Corrected | Distant | LogMAR | Monocular | dominant/distant eye | 6개월 | 안 | 58 | 0.02 | 0.13
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular | dominant/distant eye | 6개월 | 안 | 58 | 0.17 | 0.17
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Monocular | dominant/distant eye | 6개월 | 안 | 58 | 0.28 | 0.18
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular | nondominant/near eye | 6개월 | 안 | 58 | 0.27 | 0.18
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | Corrected | Distant | LogMAR | Monocular | nondominant/near eye | 6개월 | 안 | 58 | 0.06 | 0.1
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular | nondominant/near eye | 6개월 | 안 | 58 | 0.08 | 0.17
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Monocular | nondominant/near eye | 6개월 | 안 | 58 | 0.26 | 0.17
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 6개월 | 안 | 58 | 0.09 | 0.16
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | Corrected | Distant | LogMAR | Binocular |  | 6개월 | 안 | 58 | -0.05 | 0.11
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Binocular |  | 6개월 | 안 | 58 | 0.02 | 0.14
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Binocular |  | 6개월 | 안 | 58 | 0.14 | 0.12
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular | dominant/distant eye | 6개월 | 안 | 28 | 0.18 | 0.25 | 30 | 0.11 | 0.18
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | CDVA | Corrected | Distant | LogMAR | Monocular | dominant/distant eye | 6개월 | 안 | 28 | 0.02 | 0.12 | 30 | 0.01 | 0.14
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular | dominant/distant eye | 6개월 | 안 | 28 | 0.2 | 0.14 | 30 | 0.15 | 0.2
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Monocular | dominant/distant eye | 6개월 | 안 | 28 | 0.31 | 0.15 | 30 | 0.32 | 0.25
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular | nondominant/near eye | 6개월 | 안 | 28 | 0.27 | 0.22 | 30 | 0.27 | 0.18
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | CDVA | Corrected | Distant | LogMAR | Monocular | nondominant/near eye | 6개월 | 안 | 28 | 0.05 | 0.1 | 30 | 0.08 | 0.1
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular | nondominant/near eye | 6개월 | 안 | 28 | 0.11 | 0.16 | 30 | 0.06 | 0.17
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Monocular | nondominant/near eye | 6개월 | 안 | 28 | 0.3 | 0.18 | 30 | 0.23 | 0.15
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 6개월 | 안 | 28 | 0.09 | 0.17 | 30 | 0.09 | 0.14
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | CDVA | Corrected | Distant | LogMAR | Binocular |  | 6개월 | 안 | 28 | -0.06 | 0.1 | 30 | -0.04 | 0.12
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Binocular |  | 6개월 | 안 | 28 | 0.08 | 0.16 | 30 | -0.04 | 0.08
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Binocular |  | 6개월 | 안 | 28 | 0.17 | 0.12 | 30 | 0.11 | 0.11
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular | dominant/distant eye | 12개월 | 안 | 58 | 0.08 | 0.21
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | Corrected | Distant | LogMAR | Monocular | dominant/distant eye | 12개월 | 안 | 58 | 0 | 0.18
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular | dominant/distant eye | 12개월 | 안 | 58 | 0.28 | 0.19
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Monocular | dominant/distant eye | 12개월 | 안 | 58 | 0.22 | 0.2
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular | nondominant/near eye | 12개월 | 안 | 58 | 0.22 | 0.26
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | Corrected | Distant | LogMAR | Monocular | nondominant/near eye | 12개월 | 안 | 58 | 0.03 | 0.15
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular | nondominant/near eye | 12개월 | 안 | 58 | 0.13 | 0.22
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Monocular | nondominant/near eye | 12개월 | 안 | 58 | 0.27 | 0.17
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 12개월 | 안 | 58 | 0.05 | 0.2
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | Corrected | Distant | LogMAR | Binocular |  | 12개월 | 안 | 58 | -0.06 | 0.14
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Binocular |  | 12개월 | 안 | 58 | 0.05 | 0.15
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Binocular |  | 12개월 | 안 | 58 | 0.19 | 0.19
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | UIVA | Uncorrected | Intermediate | LogMAR | Binocular |  | 12개월 | 안 | 58 | -0.05 | 0.15
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular | dominant/distant eye | 12개월 | 안 | 28 | 0.13 | 0.18 | 30 | 0.03 | 0.24
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | CDVA | Corrected | Distant | LogMAR | Monocular | dominant/distant eye | 12개월 | 안 | 28 | 0.02 | 0.12 | 30 | -0.02 | 0.22
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular | dominant/distant eye | 12개월 | 안 | 28 | 0.33 | 0.17 | 30 | 0.11 | 0.18
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Monocular | dominant/distant eye | 12개월 | 안 | 28 | 0.41 | 0.14 | 30 | 0.17 | 0.14
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular | nondominant/near eye | 12개월 | 안 | 28 | 0.26 | 0.25 | 30 | 0.18 | 0.26
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | CDVA | Corrected | Distant | LogMAR | Monocular | nondominant/near eye | 12개월 | 안 | 28 | 0.07 | 0.07 | 30 | 0 | 0.19
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular | nondominant/near eye | 12개월 | 안 | 28 | 0.13 | 0.2 | 30 | 0.14 | 0.25
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UIVA | Uncorrected | Intermediate | LogMAR | Monocular | nondominant/near eye | 12개월 | 안 | 28 | -0.01 | 0.17 | 30 | -0.08 | 0.11
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Monocular | nondominant/near eye | 12개월 | 안 | 28 | 0.34 | 0.17 | 30 | 0.21 | 0.15
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 12개월 | 안 | 28 | 0.09 | 0.17 | 30 | 0.01 | 0.21
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | CDVA | Corrected | Distant | LogMAR | Binocular |  | 12개월 | 안 | 28 | -0.06 | 0.1 | 30 | -0.06 | 0.17
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Binocular |  | 12개월 | 안 | 28 | 0.11 | 0.16 | 30 | -0.02 | 0.1
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Binocular |  | 12개월 | 안 | 28 | 0.3 | 0.17 | 30 | 0.09 | 0.15
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | Satisfaction | Distance Vision |  | Distant | 1-6점(1점 Best) |  |  | 12개월 | 명 | 29 | 2.67 | 1.03
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | Satisfaction | Intermediate and near VA |  | Intermediate & near | 1-6점(1점 Best) |  |  | 12개월 | 명 | 29 | 1.94 | 0.78
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | Satisfaction | Daily activities at home |  |  | 1-6점(1점 Best) |  |  | 12개월 | 명 | 14 | 1.82 | 0.84 | 15 | 2.7 | 0.86 |  |  |  |  | 0.015
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) | 시술방법 | Satisfaction | House work |  |  | 1-6점(1점 Best) |  |  | 12개월 | 명 | 14 | 1.67 | 0.61 | 15 | 2.18 | 0.6 |  |  |  |  | 0.41
457 | Xu (2020) | Prospective, non-randomized, non-comparative study | Emmetropic presbyopia | LaserACE |  |  | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Binocular |  | 수술전 | 명 | 4 | 0.41 | 0.1
457 | Xu (2020) | Prospective, non-randomized, non-comparative study | Emmetropic presbyopia | LaserACE |  |  | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Binocular |  | 6개월 | 명 | 4 | 0.24 | 0.12 | [7 empty] | <0.05
457 | Xu (2020) | Prospective, non-randomized, non-comparative study | Emmetropic presbyopia | LaserACE |  |  | VA | DCNVA | Distance corrected | Near(40cm) | LogMAR | Binocular |  | 12개월 | 명 | 4 | 0.26 | 0.09 | [7 empty] | <0.05
457 | Xu (2020) | Prospective, non-randomized, non-comparative study | Emmetropic presbyopia | LaserACE |  |  | VA | DCIVA | Distance corrected | Intermediate(60cm) | LogMAR | Binocular |  | 수술전 | 명 | 4 | 0.33 | 0.12
457 | Xu (2020) | Prospective, non-randomized, non-comparative study | Emmetropic presbyopia | LaserACE |  |  | VA | DCIVA | Distance corrected | Intermediate(60cm) | LogMAR | Binocular |  | 6개월 | 명 | 4 | 0.12 | 0.12 | [7 empty] | <0.05
457 | Xu (2020) | Prospective, non-randomized, non-comparative study | Emmetropic presbyopia | LaserACE |  |  | VA | DCIVA | Distance corrected | Intermediate(60cm) | LogMAR | Binocular |  | 12개월 | 명 | 4 | 0.13 | 0.12 | [7 empty] | <0.05
457 | Xu (2020) | Prospective, non-randomized, non-comparative study | Emmetropic presbyopia | LaserACE |  |  | VA | UDVA | Uncorrected | Distant(4m) | LogMAR | Binocular |  | 수술전 | 명 | 4 | -0.1 | 0.08
457 | Xu (2020) | Prospective, non-randomized, non-comparative study | Emmetropic presbyopia | LaserACE |  |  | VA | UDVA | Uncorrected | Distant(4m) | LogMAR | Binocular |  | 6개월 | 명 | 4 | -0.16 | 0.07 | [7 empty] | >0.05
457 | Xu (2020) | Prospective, non-randomized, non-comparative study | Emmetropic presbyopia | LaserACE |  |  | VA | UDVA | Uncorrected | Distant(4m) | LogMAR | Binocular |  | 12개월 | 명 | 4 | -0.14 | 0.05 | [7 empty] | >0.05
457 | Xu (2020) | Prospective, non-randomized, non-comparative study | Emmetropic presbyopia | LaserACE |  |  | Satisfaction | General satisfaction |  |  | 0-3점(0점 Best) |  |  | 수술전 | 명 |  | 24.5 | 3.1
457 | Xu (2020) | Prospective, non-randomized, non-comparative study | Emmetropic presbyopia | LaserACE |  |  | Satisfaction | General satisfaction |  |  | 0-3점(0점 Best) |  |  | 6개월 | 명 |  | 15.3 | 3.2
457 | Xu (2020) | Prospective, non-randomized, non-comparative study | Emmetropic presbyopia | LaserACE |  |  | Satisfaction | General satisfaction |  |  | 0-3점(0점 Best) |  |  | 12개월 | 명 |  | 17.2 | 2.9 | [7 empty] | <0.05 | vs 수술전
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 74 | 1.01 | 0.67
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | CDVA | Corrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 74 | -0.1 | 0.06
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | UNVA | Uncorrected | Near | LogRAD | Monocular |  | 수술전 | 안 | 74 | 0.65 | 0.32
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | CNVA | Corrected | Near | LogRAD | Monocular |  | 수술전 | 안 | 74 | 0 | 0.08
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 수술전 | 명 | 37 | 1.03 | 0.76
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | CDVA | Corrected | Distant | LogMAR | Binocular |  | 수술전 | 명 | 37 | -0.18 | 0.04
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | UNVA | Uncorrected | Near | LogRAD | Binocular |  | 수술전 | 명 | 37 | 0.62 | 0.43
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | CNVA | Corrected | Near | LogRAD | Binocular |  | 수술전 | 명 | 37 | 0 | 0.05
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 12개월 | 안 | 68 | 0.01 | 0.15
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | CDVA | Corrected | Distant | LogMAR | Monocular |  | 12개월 | 안 | 68 | 0 | 0.04
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | UNVA | Uncorrected | Near | LogRAD | Monocular |  | 12개월 | 안 | 68 | 0.1 | 0.13
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | CNVA | Corrected | Near | LogRAD | Monocular |  | 12개월 | 안 | 68 | 0.07 | 0.09
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 12개월 | 명 | 34 | 0 | 0.18
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | CDVA | Corrected | Distant | LogMAR | Binocular |  | 12개월 | 명 | 34 | -0.05 | 0.05
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | UNVA | Uncorrected | Near | LogRAD | Binocular |  | 12개월 | 명 | 34 | 0 | 0.16
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | CNVA | Corrected | Near | LogRAD | Binocular |  | 12개월 | 명 | 34 | -0.03 | 0.05
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | Satisfaction | Near vision |  |  | 0-100점(100점 Best) | Binocular |  | 수술전 | 명 | 37 | 30 | 11
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | Satisfaction | Distance vision |  |  | 0-100점(100점 Best) | Binocular |  | 수술전 | 명 | 37 | 37 | 12
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | Spectacle Independence | Dependence on correction |  |  | 0-100점(100점 Best) | Binocular |  | 수술전 | 명 | 37 | 30 | 15
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | Satisfaction | Satisfaction with correction |  |  | 0-100점(100점 Best) | Binocular |  | 수술전 | 명 | 37 | 56 | 14
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | Satisfaction | Near vision |  |  | 0-100점(100점 Best) | Binocular |  | 12개월 | 명 | 34 | 94 | 6 | [7 empty] | <0.05 | vs 수술전
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | Satisfaction | Distance vision |  |  | 0-100점(100점 Best) | Binocular |  | 12개월 | 명 | 34 | 92 | 7 | [7 empty] | <0.05 | vs 수술전
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | Spectacle Independence | Dependence on correction |  |  | 0-100점(100점 Best) | Binocular |  | 12개월 | 명 | 34 | 95 | 2 | [7 empty] | <0.05 | vs 수술전
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | Satisfaction | Satisfaction with correction |  |  | 0-100점(100점 Best) | Binocular |  | 12개월 | 명 | 34 | 93 | 8 | [7 empty] | <0.05 | vs 수술전
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | CS | 3cpd |  |  | Logarithmic scale |  |  | 수술전 | 명 | 19 | 1.54 | NA
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | CS | 5cpd |  |  | Logarithmic scale |  |  | 수술전 | 명 | 19 | 1.62 | NA
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | CS | 12cpd |  |  | Logarithmic scale |  |  | 수술전 | 명 | 19 | 1.29 | NA
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | CS | 18cpd |  |  | Logarithmic scale |  |  | 수술전 | 명 | 19 | 0.83 | NA
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | CS | 3cpd |  |  | Logarithmic scale |  |  | 6개월? | 명 | 19 | 1.49 | NA | [7 empty] | 0.221 | p값으로 SD 산출
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | CS | 5cpd |  |  | Logarithmic scale |  |  | 6개월? | 명 | 19 | 1.56 | NA | [7 empty] | 0.164 | p값으로 SD 산출
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | CS | 12cpd |  |  | Logarithmic scale |  |  | 6개월? | 명 | 19 | 1.31 | NA | [7 empty] | 0.535 | p값으로 SD 산출
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | CS | 18cpd |  |  | Logarithmic scale |  |  | 6개월? | 명 | 19 | 0.84 | NA | [7 empty] | 0.367 | p값으로 SD 산출
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | CS | 3cpd |  |  | Logarithmic scale |  |  | 수술전 | 명 | 14 | 1.47 | NA
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | CS | 5cpd |  |  | Logarithmic scale |  |  | 수술전 | 명 | 14 | 1.56 | NA
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | CS | 12cpd |  |  | Logarithmic scale |  |  | 수술전 | 명 | 14 | 1.34 | NA
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | CS | 18cpd |  |  | Logarithmic scale |  |  | 수술전 | 명 | 14 | 1.05 | NA
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | CS | 3cpd |  |  | Logarithmic scale |  |  | 6개월? | 명 | 14 | 1.38 | NA | [7 empty] | 0.276 | p값으로 SD 산출
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | CS | 5cpd |  |  | Logarithmic scale |  |  | 6개월? | 명 | 14 | 1.52 | NA | [7 empty] | 0.412 | p값으로 SD 산출
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | CS | 12cpd |  |  | Logarithmic scale |  |  | 6개월? | 명 | 14 | 1.27 | NA | [7 empty] | 0.364 | p값으로 SD 산출
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | CS | 18cpd |  |  | Logarithmic scale |  |  | 6개월? | 명 | 14 | 0.91 | NA | [7 empty] | 0.021 | p값으로 SD 산출
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.00D) | CS | 3cpd |  |  | Logarithmic scale |  |  | 수술전 | 명 | 17 | 1.42 | NA
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.01D) | CS | 5cpd |  |  | Logarithmic scale |  |  | 수술전 | 명 | 17 | 1.6 | NA
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.02D) | CS | 12cpd |  |  | Logarithmic scale |  |  | 수술전 | 명 | 17 | 1.25 | NA
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.03D) | CS | 18cpd |  |  | Logarithmic scale |  |  | 수술전 | 명 | 17 | 1.1 | NA
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.04D) | CS | 3cpd |  |  | Logarithmic scale |  |  | 6개월? | 명 | 17 | 1.49 | NA | [7 empty] | 0.34 | p값으로 SD 산출
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.05D) | CS | 5cpd |  |  | Logarithmic scale |  |  | 6개월? | 명 | 17 | 1.51 | NA | [7 empty] | 0.262 | p값으로 SD 산출
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.06D) | CS | 12cpd |  |  | Logarithmic scale |  |  | 6개월? | 명 | 17 | 1.18 | NA | [7 empty] | 0.42 | p값으로 SD 산출
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.07D) | CS | 18cpd |  |  | Logarithmic scale |  |  | 6개월? | 명 | 17 | 0.92 | NA | [7 empty] | 0.371 | p값으로 SD 산출
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UNVA | Uncorrected | Near | LogMAR | Monocular |  | 수술전 | 안 | 72 | 0.29 | 0.3
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UIVA | Uncorrected | Intermediate | LogMAR | Monocular |  | 수술전 | 안 | 72 | 0.38 | 0.36
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 72 | 0.75 | 0.45
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UNVA | Uncorrected | Near | LogMAR | Monocular |  | 6개월 | 안 | 72 | -0.01 | 0.11
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UIVA | Uncorrected | Intermediate | LogMAR | Monocular |  | 6개월 | 안 | 72 | -0.01 | 0.07
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 6개월 | 안 | 72 | -0.09 | 0.09
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | dominant eye | VA | UNVA | Uncorrected | Near | LogMAR | Monocular |  | 6개월 | 안 | 36 | 0.09 | 0.11
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | dominant eye | VA | UIVA | Uncorrected | Intermediate | LogMAR | Monocular |  | 6개월 | 안 | 36 | 0.02 | 0.04
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 6개월 | 안 | 36 | 0.02 | 0.07
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | dominant eye | VA | UNVA | Uncorrected | Near | LogMAR | Monocular |  | 3개월 | 안 | 36 | 0.11 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | dominant eye | VA | UIVA | Uncorrected | Intermediate | LogMAR | Monocular |  | 3개월 | 안 | 36 | 0.02 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 3개월 | 안 | 36 | 0.03 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | dominant eye | VA | UNVA | Uncorrected | Near | LogMAR | Monocular |  | 1개월 | 안 | 36 | 0.1 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | dominant eye | VA | UIVA | Uncorrected | Intermediate | LogMAR | Monocular |  | 1개월 | 안 | 36 | 0.04 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 1개월 | 안 | 36 | 0.07 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | dominant eye | VA | UNVA | Uncorrected | Near | LogMAR | Monocular |  | 수술전 | 안 | 36 | 0.33 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | dominant eye | VA | UIVA | Uncorrected | Intermediate | LogMAR | Monocular |  | 수술전 | 안 | 36 | 0.39 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 36 | 0.79 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | non-dominant eye | VA | UNVA | Uncorrected | Near | LogMAR | Monocular |  | 6개월 | 안 | 36 | 0.04 | 0.1
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | non-dominant eye | VA | UIVA | Uncorrected | Intermediate | LogMAR | Monocular |  | 6개월 | 안 | 36 | 0.01 | 0.03
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | non-dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 6개월 | 안 | 36 | 0.08 | 0.08
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | non-dominant eye | VA | UNVA | Uncorrected | Near | LogMAR | Monocular |  | 3개월 | 안 | 36 | 0.05 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | non-dominant eye | VA | UIVA | Uncorrected | Intermediate | LogMAR | Monocular |  | 3개월 | 안 | 36 | 0.02 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | non-dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 3개월 | 안 | 36 | 0.1 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | non-dominant eye | VA | UNVA | Uncorrected | Near | LogMAR | Monocular |  | 1개월 | 안 | 36 | 0.06 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | non-dominant eye | VA | UIVA | Uncorrected | Intermediate | LogMAR | Monocular |  | 1개월 | 안 | 36 | 0.03 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | non-dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 1개월 | 안 | 36 | 0.17 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | non-dominant eye | VA | UNVA | Uncorrected | Near | LogMAR | Monocular |  | 수술전 | 안 | 36 | 0.32 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | non-dominant eye | VA | UIVA | Uncorrected | Intermediate | LogMAR | Monocular |  | 수술전 | 안 | 36 | 0.36 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | non-dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 36 | 0.76 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UNVA | Uncorrected | Near | LogMAR | Binocular |  | 6개월 | 명 | 36 | 0.03 | 0.1
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UIVA | Uncorrected | Intermediate | LogMAR | Binocular |  | 6개월 | 명 | 36 | 0.01 | 0.02
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 6개월 | 명 | 36 | 0 | 0.05
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UNVA | Uncorrected | Near | LogMAR | Binocular |  | 3개월 | 명 | 36 | 0.06 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UIVA | Uncorrected | Intermediate | LogMAR | Binocular |  | 3개월 | 명 | 36 | 0.01 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 3개월 | 명 | 36 | 0.01 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UNVA | Uncorrected | Near | LogMAR | Binocular |  | 1개월 | 명 | 36 | 0.05 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UIVA | Uncorrected | Intermediate | LogMAR | Binocular |  | 1개월 | 명 | 36 | 0.02 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 1개월 | 명 | 36 | 0.05 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UNVA | Uncorrected | Near | LogMAR | Binocular |  | 수술전 | 명 | 36 | 0.29 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UIVA | Uncorrected | Intermediate | LogMAR | Binocular |  | 수술전 | 명 | 36 | 0.38 | NA
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 수술전 | 명 | 36 | 0.75 | NA
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | VA | CDVA | Corrected | Distant | Radner reading charts | Monocular |  | 수술전 | 안 | 39 | 1.07 | 0.09
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | VA | UDVA | Uncorrected |  | Radner reading charts | Monocular |  | 수술전 | 안 | 39 | 0.56 | 0.31
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | VA | UDVA | Uncorrected |  | Radner reading charts | Monocular |  | 12개월 | 안 | 39 | 0.9 | 0.15
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | VA | UNVA | Uncorrected | Near(40cm) | Radner reading charts | Monocular |  | 수술전 | 안 | 39 | 0.15 | 0.07
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | VA | UNVA | Uncorrected | Near(40cm) | Radner reading charts | Monocular |  | 12개월 | 안 | 39 | 0.62 | 0.24
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | dominant eye | VA | UDVA | Uncorrected | Distant | Radner reading charts | Monocular |  | 수술전 | 안 | 19 | 0.56 | 0.25
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | dominant eye | VA | UDVA | Uncorrected | Distant | Radner reading charts | Monocular |  | 12개월 | 안 | 19 | 0.93 | 0.13
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | non-dominant eye | VA | UDVA | Uncorrected | Distant | Radner reading charts | Monocular |  | 수술전 | 안 | 20 | 0.51 | 0.32
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | non-dominant eye | VA | UDVA | Uncorrected | Distant | Radner reading charts | Monocular |  | 12개월 | 안 | 20 | 0.82 | 0.15
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | CS | CS |  |  | Logarithmic scale |  |  | 12개월 | 안 | 39 | 1.83 | 0.18
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Spectacle Independence | Near vision improved without glasses |  |  | 0-10점(10점 best) |  |  | 12개월 | 명 | 20 | 8.2 | 1.9
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Spectacle Independence | Distance vision improved without glasses |  |  | 0-10점(10점 best) |  |  | 12개월 | 명 | 20 | 6.4 | 3.1
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Spectacle Independence | SMS reading without glasses |  |  | 0-10점(10점 best) |  |  | 12개월 | 명 | 20 | 7.35 | 3.4
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Spectacle Independence | Book reading without glasses |  |  | 0-10점(10점 best) |  |  | 12개월 | 명 | 20 | 7.1 | 3.5
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Spectacle Independence | Shopping without glasses |  |  | 0-10점(10점 best) |  |  | 12개월 | 명 | 20 | 7.5 | 3
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Spectacle Independence | Work at PC screen without glasses |  |  | 0-10점(10점 best) |  |  | 12개월 | 명 | 20 | 7.6 | 2.95
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Complication | Glare |  |  | 1-5점(1점 low) |  |  | 12개월 | 명 | 20 | 1.95 | 1.4
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Complication | Halos |  |  | 1-5점(1점 low) |  |  | 12개월 | 명 | 20 | 1.3 | 1.6
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Complication | Double vision |  |  | 1-5점(1점 low) |  |  | 12개월 | 명 | 20 | 0.35 | 0.75
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Complication | Fluctuations in vision |  |  | 1-5점(1점 low) |  |  | 12개월 | 명 | 20 | 1.5 | 1.2
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Complication | Headaches |  |  | 1-5점(1점 low) |  |  | 12개월 | 명 | 20 | 0.4 | 1.2
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Complication | Reduced distance vision |  |  | 1-5점(1점 low) |  |  | 12개월 | 명 | 20 | 1.8 | 1.7
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 YES | VA | UDVA | Uncorrected | Distant | Decimal | Binocular |  | 12개월 | 명 | 15 | 0.82 | 0.15
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 No | VA | UDVA | Uncorrected | Distant |  | Binocular |  | 12개월 | 명 | 5 | 0.93 | 0.2
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 YES | VA | UNVA | Uncorrected | Near(40cm) | Decimal | Binocular |  | 12개월 | 명 | 15 | 0.65 | 0.2
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 No | VA | UNVA | Uncorrected | Near(40cm) |  | Binocular |  | 12개월 | 명 | 5 | 0.54 | 0.2
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 YES | Complication | Glare |  |  | 1-5점(1점 low) |  |  | 12개월 | 명 | 20 | 1.7 | 1.5
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 YES | Complication | Halos |  |  | 1-5점(1점 low) |  |  | 12개월 | 명 | 20 | 1 | 1.25
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 YES | Complication | Double vision |  |  | 1-5점(1점 low) |  |  | 12개월 | 명 | 20 | 0.2 | 0.4
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 YES | Complication | Fluctuations in vision |  |  | 1-5점(1점 low) |  |  | 12개월 | 명 | 20 | 1.33 | 1.2
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 YES | Complication | Headaches |  |  | 1-5점(1점 low) |  |  | 12개월 | 명 | 20 | 0.1 | 0.3
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 YES | Complication | Reduced distance vision |  |  | 1-5점(1점 low) |  |  | 12개월 | 명 | 20 | 1.3 | 1.2
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 No | Complication | Glare |  |  | 1-5점(1점 low) |  |  | 12개월 | 명 | 20 | 2.6 | 1.1
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 No | Complication | Halos |  |  | 1-5점(1점 low) |  |  | 12개월 | 명 | 20 | 2.2 | 2.3
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 No | Complication | Double vision |  |  | 1-5점(1점 low) |  |  | 12개월 | 명 | 20 | 0.8 | 1.3
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 No | Complication | Fluctuations in vision |  |  | 1-5점(1점 low) |  |  | 12개월 | 명 | 20 | 2 | 1.2
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 No | Complication | Headaches |  |  | 1-5점(1점 low) |  |  | 12개월 | 명 | 20 | 1.4 | 2.2
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 No | Complication | Reduced distance vision |  |  | 1-5점(1점 low) |  |  | 12개월 | 명 | 20 | 3.4 | 2.1
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 1개월 | 명 | 69 | 0.07 | 0.11
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 3개월 | 명 | 69 | 0.06 | 0.09
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 6개월 | 명 | 69 | 0.03 | 0.07
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 9개월 | 명 | 69 | -0.01 | 0.05
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 12개월 | 명 | 69 | -0.04 | 0.05
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | Uncorrected | Intermediate(60cm) | LogMAR | Binocular |  | 1개월 | 명 | 69 | -0.01 | 0.16
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | Uncorrected | Intermediate(61cm) | LogMAR | Binocular |  | 3개월 | 명 | 69 | -0.04 | 0.13
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | Uncorrected | Intermediate(62cm) | LogMAR | Binocular |  | 6개월 | 명 | 69 | -0.02 | 0.15
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | Uncorrected | Intermediate(63cm) | LogMAR | Binocular |  | 9개월 | 명 | 69 | -0.07 | 0.11
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | Uncorrected | Intermediate(64cm) | LogMAR | Binocular |  | 12개월 | 명 | 69 | -0.12 | 0.13
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Binocular |  | 1개월 | 명 | 69 | 0.13 | 0.12
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Binocular |  | 3개월 | 명 | 69 | 0.11 | 0.15
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Binocular |  | 6개월 | 명 | 69 | 0.13 | 0.22
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Binocular |  | 9개월 | 명 | 69 | 0.15 | 0.15
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Binocular |  | 12개월 | 명 | 69 | 0.1 | 0.08
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 수술전 | 명 | 69 | 0.42 | 0.55
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Binocular |  | 수술전 | 명 | 69 | 0.78 | 0.2
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q | VA | UIVA | Uncorrected | Intermediate(60cm) | LogMAR | Binocular |  | 수술전 | 명 | 69 | NA
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q, non-dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 69 | 0.43 | 0.4
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q, non-dominant eye | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Monocular |  | 수술전 | 안 | 69 | NA
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 1개월 | 명 | 69 | 0.08 | 0.12
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Binocular |  | 1개월 | 명 | 69 | 0.16 | 0.19
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q | VA | UIVA | Uncorrected | Intermediate(60cm) | LogMAR | Binocular |  | 1개월 | 명 | 69 | -0.053 | 0.15
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q, non-dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 1개월 | 안 | 69 | 0.39 | 0.25
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q, non-dominant eye | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Monocular |  | 1개월 | 안 | 69 | 0.17 | 0.24
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 3개월 | 명 | 69 | 0.068 | 0.14
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Binocular |  | 3개월 | 명 | 69 | 0.094 | 0.17
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q | VA | UIVA | Uncorrected | Intermediate(60cm) | LogMAR | Binocular |  | 3개월 | 명 | 69 | -0.023 | 0.14
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q, non-dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 3개월 | 안 | 69 | 0.34 | 0.27
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q, non-dominant eye | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Monocular |  | 3개월 | 안 | 69 | 0.12 | 0.16
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 6개월 | 명 | 69 | 0.078 | 0.1
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Binocular |  | 6개월 | 명 | 69 | 0.089 | 0.18
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q | VA | UIVA | Uncorrected | Intermediate(60cm) | LogMAR | Binocular |  | 6개월 | 명 | 69 | -0.033 | 0.225
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q, non-dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 6개월 | 안 | 69 | 0.39 | 0.25
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q, non-dominant eye | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Monocular |  | 6개월 | 안 | 69 | 0.1 | 0.15
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 9개월 | 명 | 69 | 0.007 | 0.1
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Binocular |  | 9개월 | 명 | 69 | 0.16 | 0.15
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q | VA | UIVA | Uncorrected | Intermediate(60cm) | LogMAR | Binocular |  | 9개월 | 명 | 69 | -0.033 | 0.06
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q, non-dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 9개월 | 안 | 69 | 0.31 | 0.26
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q, non-dominant eye | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Monocular |  | 9개월 | 안 | 69 | 0.1 | 0.1
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 12개월 | 명 | 69 | -0.05 | 0.05
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Binocular |  | 12개월 | 명 | 69 | 0.08 | 0.07
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q | VA | UIVA | Uncorrected | Intermediate(60cm) | LogMAR | Binocular |  | 12개월 | 명 | 69 | -0.13 | 0.15
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q, non-dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 12개월 | 안 | 69 | 0.35 | 0.24
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.8△Q, non-dominant eye | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Monocular |  | 12개월 | 안 | 69 | 0.1 | 0.1
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 수술전 | 명 | 69 | 0.36 | 0.46
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Binocular |  | 수술전 | 명 | 69 | 0.78 | 0.43
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q | VA | UIVA | Uncorrected | Intermediate(60cm) | LogMAR | Binocular |  | 수술전 | 명 | 69 | NA
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q, non-dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 69 | 0.45 | 0.47
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q, non-dominant eye | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Monocular |  | 수술전 | 안 | 69 | NA
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 1개월 | 명 | 69 | 0.039 | 0.06
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Binocular |  | 1개월 | 명 | 69 | 0.11 | 0.1
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q | VA | UIVA | Uncorrected | Intermediate(60cm) | LogMAR | Binocular |  | 1개월 | 명 | 69 | -0.075 | 0.096
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q, non-dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 1개월 | 안 | 69 | 0.35 | 0.28
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q, non-dominant eye | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Monocular |  | 1개월 | 안 | 69 | 0.16 | 0.11
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 3개월 | 명 | 69 | 0.051 | 0.07
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Binocular |  | 3개월 | 명 | 69 | 0.15 | 0.14
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q | VA | UIVA | Uncorrected | Intermediate(60cm) | LogMAR | Binocular |  | 3개월 | 명 | 69 | 0 | 0.21
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q, non-dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 3개월 | 안 | 69 | 0.39 | 0.36
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q, non-dominant eye | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Monocular |  | 3개월 | 안 | 69 | 0.18 | 0.18
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 6개월 | 명 | 69 | 0.009 | 0.03
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Binocular |  | 6개월 | 명 | 69 | 0.12 | 0.11
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q | VA | UIVA | Uncorrected | Intermediate(60cm) | LogMAR | Binocular |  | 6개월 | 명 | 69 | -0.038 | 0.1
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q, non-dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 6개월 | 안 | 69 | 0.15 | 0.14
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q, non-dominant eye | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Monocular |  | 6개월 | 안 | 69 | 0.15 | 0.15
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 9개월 | 명 | 69 | 0 | 0.001
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Binocular |  | 9개월 | 명 | 69 | 0.19 | 0.01
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q | VA | UIVA | Uncorrected | Intermediate(60cm) | LogMAR | Binocular |  | 9개월 | 명 | 69 | -0.05 | 0.06
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q, non-dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 9개월 | 안 | 69 | 0.26 | 0.21
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q, non-dominant eye | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Monocular |  | 9개월 | 안 | 69 | 0.25 | 0.2
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 12개월 | 명 | 69 | -0.025 | 0.005
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Binocular |  | 12개월 | 명 | 69 | 0.13 | 0.08
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q | VA | UIVA | Uncorrected | Intermediate(60cm) | LogMAR | Binocular |  | 12개월 | 명 | 69 | -0.13 | 0.15
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q, non-dominant eye | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 12개월 | 안 | 69 | 0.28 | 0.34
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  | -0.6△Q, non-dominant eye | VA | UNVA | Uncorrected | Near(33cm) | LogMAR | Monocular |  | 12개월 | 안 | 69 | 0.19 | 0.01
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | distance eyes | VA | UDVA | Uncorrected | Distant(3m) | LogMAR | Monocular |  | 3개월 | 안 | 40 | -0.1 | 0.06
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | near eyes | VA | UDVA | Uncorrected | Distant(3m) | LogMAR | Monocular |  | 3개월 | 안 | 40 | 0.22 | 0.12
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | VA | UDVA | Uncorrected | Distant(3m) | LogMAR | Binocular |  | 3개월 | 명 | 80 | -0.11 | 0.06
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | distance eyes | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 40 | 0.19 | 0.21
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | near eyes | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Monocular |  | 3개월 | 안 | 40 | 0.07 | 0.07
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | VA | UNVA | Uncorrected | Near(40cm) | LogMAR | Binocular |  | 3개월 | 명 | 80 | 0.04 | 0.06
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | mesopic | CS | AULCSF |  |  | Logarithmic scale |  |  | 수술전 | 안 | 80 | 1.38 | NA
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | mesopic | CS | AULCSF |  |  | Logarithmic scale |  |  | 3개월 | 안 | 80 | 1.41 | NA
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | photopic | CS | AULCSF |  |  | Logarithmic scale |  |  | 수술전 | 안 | 80 | 1.42 | NA
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | photopic | CS | AULCSF |  |  | Logarithmic scale |  |  | 3개월 | 안 | 80 | 1.43 | NA
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | Satisfaction | Near vision |  |  | 0-100점(100점 best) |  |  | 수술전 | 명 | 40 | 75 | 21
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | Satisfaction | Near vision |  |  | 0-100점(100점 best) |  |  | 3개월 | 명 | 40 | 90 | 6 | [7 empty] | <.05
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | Satisfaction | distance vision |  |  | 0-100점(100점 best) |  |  | 수술전 | 명 | 40 | 39 | 17
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | Satisfaction | distance vision |  |  | 0-100점(100점 best) |  |  | 3개월 | 명 | 40 | 89 | 8 | [7 empty] | <.05
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | Spectacle Independence | Dependence on correction |  |  | 0-100점(100점 best) |  |  | 수술전 | 명 | 40 | 28 | 24
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | Spectacle Independence | Dependence on correction |  |  | 0-100점(100점 best) |  |  | 3개월 | 명 | 40 | 94 | 5 | [7 empty] | <.05
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | Satisfaction | Satisfaction with correction |  |  | 0-100점(100점 best) |  |  | 수술전 | 명 | 40 | 76 | 15
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | Satisfaction | Satisfaction with correction |  |  | 0-100점(100점 best) |  |  | 3개월 | 명 | 40 | 92 | 6 | [7 empty] | <.05
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | asymmetrical presbyLASIK | symmetrical presbyLASIK |  | VA | UDVA | Uncorrected | Distant | Decimal(snellen) | Monocular |  | 3개월 | 안 | 21 | 0.8 | 0.2 | 19 | 0.9 | 0.2 |  |  |  |  | <0.01
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | asymmetrical presbyLASIK | symmetrical presbyLASIK |  | VA | UNVA | Uncorrected | Near | Decimal(snellen) | Monocular |  | 3개월 | 안 | 21 | 0.8 | 0.3 | 19 | 0.8 | 0.2 |  |  |  |  | <0.01
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | asymmetrical presbyLASIK | symmetrical presbyLASIK |  | VA | UDVA | Uncorrected | Distant | Decimal(snellen) | Binocular |  | 3개월 | 명 | 14 | 1 | 0.007 | 16 | 1 | NA |  |  |  |  | <0.01
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | asymmetrical presbyLASIK | symmetrical presbyLASIK |  | VA | UNVA | Uncorrected | Near | Decimal(snellen) | Binocular |  | 3개월 | 명 | 14 | 0.8 | 0.3 | 16 | 0.8 | 0.2 |  |  |  |  | <0.01
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | asymmetrical presbyLASIK | symmetrical presbyLASIK |  | VA | UDVA | Uncorrected | Distant | Decimal(snellen) | Monocular |  | 18개월 | 안 | 21 | 0.9 | 0.2 | 19 | 0.8 | 0.2
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | PresLASIK | symmetrical presbyLASIK |  | VA | UDVA | Uncorrected | Distant | Decimal(snellen) | Monocular | All | 18개월 | 안 | 40 | 0.9 | 0.2
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | asymmetrical presbyLASIK | symmetrical presbyLASIK |  | VA | UNVA | Uncorrected | Near | Decimal(snellen) | Monocular |  | 18개월 | 안 | 21 | 0.8 | 0.2 | 19 | 0.9 | 0.2
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | PresLASIK | symmetrical presbyLASIK |  | VA | UNVA | Uncorrected | Near | Decimal(snellen) | Monocular | All | 18개월 | 안 | 40 | 0.8 | 0.2
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | asymmetrical presbyLASIK | symmetrical presbyLASIK | Posterior enhancement | VA | UDVA | Uncorrected | Distant | Decimal(snellen) | Monocular |  | 18개월 | 안 | 7 | 0.9 | 0.1 | 13 | 0.9 | 0.1
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | PresLASIK | symmetrical presbyLASIK | Posterior enhancement | VA | UDVA | Uncorrected | Distant | Decimal(snellen) | Monocular | All | 18개월 | 안 | 20 | 0.9 | 0.1
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | asymmetrical presbyLASIK | symmetrical presbyLASIK | Posterior enhancement | VA | UNVA | Uncorrected | Near | Decimal(snellen) | Monocular |  | 18개월 | 안 | 7 | 0.8 | 0.2 | 13 | 0.8 | 0.2
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | PresLASIK | symmetrical presbyLASIK | Posterior enhancement | VA | UNVA | Uncorrected | Near | Decimal(snellen) | Monocular | All | 18개월 | 안 | 20 | 0.8 | 0.1
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | asymmetrical presbyLASIK | symmetrical presbyLASIK |  | Satisfaction | General satisfaction |  |  | 0-10점(10점 best) |  |  | 18개월 | 명 | 14 | 7.3 | 2.8 | 16 | 7 | 2.6
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | asymmetrical presbyLASIK | symmetrical presbyLASIK |  | Satisfaction | Reading satisfaction |  |  | 0-10점(10점 best) |  |  | 18개월 | 명 | 14 | 8.4 | 2.1 | 16 | 7.3 | 2
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | asymmetrical presbyLASIK | symmetrical presbyLASIK |  | Satisfaction | Far vision difficulty |  |  | 0-4점(0점 best) |  |  | 18개월 | 명 | 14 | 1.5 | 1.5 | 16 | 0.7 | 1.1
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | asymmetrical presbyLASIK | symmetrical presbyLASIK |  | Satisfaction | Near vision difficulty |  |  | 0-4점(0점 best) |  |  | 18개월 | 명 | 14 | 1.3 | 1.5 | 16 | 1 | 1
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | asymmetrical presbyLASIK | symmetrical presbyLASIK |  | Satisfaction | Intermediate difficulty |  |  | 0-4점(0점 best) |  |  | 18개월 | 명 | 14 | 1 | 0 | 16 | 0.1 | 0.4
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | asymmetrical presbyLASIK | symmetrical presbyLASIK |  | Satisfaction | Halos difficulty |  |  | 0-4점(0점 best) |  |  | 18개월 | 명 | 14 | 0.6 | 1.5 | 16 | 1 | 1.4
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | SUPRACOR |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 3개월 | 명 | 23 | 0.2 | 0.15
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | SUPRACOR |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 3개월 | 명 | 23 | 0.1 | 0.1
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | SUPRACOR |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 6개월 | 명 | 23 | 0.21 | 0.17
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | SUPRACOR |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 6개월 | 명 | 23 | 0.07 | 0.12
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | SUPRACOR |  | Overcorrected by more than 0.50D | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 6개월? | 명 | 23 | 0.33 | 0.19
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | SUPRACOR |  | Undercorrected by more than 0.50D | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 6개월? | 명 | 23 | 0.14 | 0.12
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | SUPRACOR |  |  | VA | CDVA | Corrected | Distant | LogMAR | Binocular |  | 수술전 | 명 | 23 | -0.03 | 0.07
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | SUPRACOR |  |  | VA | CDVA | Corrected | Distant | LogMAR | Binocular |  | 3개월 | 명 | 23 | -0.02 | 0.6
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | SUPRACOR |  |  | VA | CDVA | Corrected | Distant | LogMAR | Binocular |  | 6개월 | 명 | 23 | -0.01 | 0.1
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | SUPRACOR |  |  | VA | UDVA | Uncorrected | Distant | logMAR | Monocular |  | 수술전 | 명 | 23 | 0.32 | 0.09
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | SUPRACOR |  |  | VA | UDVA | Uncorrected | Distant | logMAR | Binocular |  | 수술전 | 명 | 23 | 0.23 | 0.07
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | SUPRACOR |  |  | VA | UDVA | Uncorrected | Distant | logMAR | Monocular |  | 6개월? | 명 | 23 | 0
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | SUPRACOR |  |  | VA | UDVA | Uncorrected | Distant | logMAR | Binocular |  | 6개월? | 명 | 23 | 0
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 66 | 1.07 | 0.8
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | CDVA | Corrected | Distant | LogMAR | Monocular |  | 수술전 | 안 | 66 | -0.1 | 0.06
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | UNVA | Uncorrected | Near | LogRAD | Monocular |  | 수술전 | 안 | 66 | 0.67 | 0.4
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | CNVA | corrected | Near | LogRAD | Monocular |  | 수술전 | 안 | 66 | 0 | 0.09
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 수술전 | 안 | 33 | 1.02 | 0.75
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | CDVA | Corrected | Distant | LogMAR | Binocular |  | 수술전 | 안 | 33 | -0.1 | 0.05
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | UNVA | Uncorrected | Near | LogRAD | Binocular |  | 수술전 | 안 | 33 | 0.62 | 0.35
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | CNVA | corrected | Near | LogRAD | Binocular |  | 수술전 | 안 | 33 | -0.05 | 0.05
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Monocular |  | 1년 | 안 | 62 | 0.19 | 0.18
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | CDVA | Corrected | Distant | LogMAR | Monocular |  | 1년 | 안 | 62 | 0 | 0.05
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | UNVA | Uncorrected | Near | LogRAD | Monocular |  | 1년 | 안 | 62 | 0.15 | 0.15
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | CNVA | corrected | Near | LogRAD | Monocular |  | 1년 | 안 | 62 | 0.07 | 0.09
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 1년 | 안 | 31 | 0.07 | 0.12
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | CDVA | Corrected | Distant | LogMAR | Binocular |  | 1년 | 안 | 31 | -0.05 | 0.04
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | UNVA | Uncorrected | Near | LogRAD | Binocular |  | 1년 | 안 | 31 | 0.05 | 0.1
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | CNVA | corrected | Near | LogRAD | Binocular |  | 1년 | 안 | 31 | -0.03 | 0.05
1310 | Oh (2013) | Prospective comparative study | Presbyopia | Aspheric corneal ablation with the CST | conventional presbyopic excimer laser (NCST) | CST | VA | UDVA | Uncorrected | Distant | Lines | Binocular |  | baseline | 명 | 12 | 20/20 | 0.8
1310 | Oh (2013) | Prospective comparative study | Presbyopia | Aspheric corneal ablation with the CST | conventional presbyopic excimer laser (NCST) | CST | VA | UNVA | Uncorrected | Near | Lines | Binocular |  | baseline | 명 | 12 | 20/60 | 0.4
1310 | Oh (2013) | Prospective comparative study | Presbyopia | Aspheric corneal ablation with the CST | conventional presbyopic excimer laser (NCST) | NCST | VA | UDVA | Uncorrected | Distant | Lines | Binocular |  | baseline | 명 | 11 | 20/20 | 1.2
1310 | Oh (2013) | Prospective comparative study | Presbyopia | Aspheric corneal ablation with the CST | conventional presbyopic excimer laser (NCST) | NCST | VA | UNVA | Uncorrected | Near | Lines | Binocular |  | baseline | 명 | 11 | 20/50 | 0.8
1310 | Oh (2013) | Prospective comparative study | Presbyopia | Aspheric corneal ablation with the CST | conventional presbyopic excimer laser (NCST) | CST | VA | UDVA | Uncorrected | Distant | Lines | Binocular |  | 3개월 | 명 | 12 | 20/16 | 0.4
1310 | Oh (2013) | Prospective comparative study | Presbyopia | Aspheric corneal ablation with the CST | conventional presbyopic excimer laser (NCST) | CST | VA | UNVA | Uncorrected | Near | Lines | Binocular |  | 3개월 | 명 | 12 | 20/20 | 0.6
1310 | Oh (2013) | Prospective comparative study | Presbyopia | Aspheric corneal ablation with the CST | conventional presbyopic excimer laser (NCST) | NCST | VA | UDVA | Uncorrected | Distant | Lines | Binocular |  | 3개월 | 명 | 11 | 20/20 | 0.2
1310 | Oh (2013) | Prospective comparative study | Presbyopia | Aspheric corneal ablation with the CST | conventional presbyopic excimer laser (NCST) | NCST | VA | UNVA | Uncorrected | Near | Lines | Binocular |  | 3개월 | 명 | 11 | 20/20 | 0.6
1310 | Oh (2013) | Prospective comparative study | Presbyopia | Aspheric corneal ablation with the CST | conventional presbyopic excimer laser (NCST) | CST | VA | UDVA | Uncorrected | Distant | Lines | Binocular |  | 2년 | 명 | 12 | 20/20 | 0.4
1310 | Oh (2013) | Prospective comparative study | Presbyopia | Aspheric corneal ablation with the CST | conventional presbyopic excimer laser (NCST) | CST | VA | UNVA | Uncorrected | Near | Lines | Binocular |  | 2년 | 명 | 12 | 12/15 | 0.6
1310 | Oh (2013) | Prospective comparative study | Presbyopia | Aspheric corneal ablation with the CST | conventional presbyopic excimer laser (NCST) | NCST | VA | UDVA | Uncorrected | Distant | Lines | Binocular |  | 2년 | 명 | 11 | 20/20 | 2.1
1310 | Oh (2013) | Prospective comparative study | Presbyopia | Aspheric corneal ablation with the CST | conventional presbyopic excimer laser (NCST) | NCST | VA | UNVA | Uncorrected | Near | Lines | Binocular |  | 2년 | 명 | 11 | 10/20 | 1.5
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | hyperopic | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 수술전 | 명 | 10 | 0.28 | 0.25
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | hyperopic | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 6개월 | 명 | 10 | -0.04 | 0.07
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | emmetropic | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 수술전 | 명 | 10 | -0.05 | 0.07
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | emmetropic | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 6개월 | 명 | 10 | 0.03 | 0.11
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | myopic | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 수술전 | 명 | 10 | 0.78 | 0.27
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | myopic | VA | UDVA | Uncorrected | Distant | LogMAR | Binocular |  | 6개월 | 명 | 10 | 0.09 | 0.08
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | hyperopic | VA | UNVA | Uncorrected | Near | LogRAD | Binocular |  | 수술전 | 명 | 10 | 0.86 | 0.62
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | hyperopic | VA | UNVA | Uncorrected | Near | LogRAD | Binocular |  | 6개월 | 명 | 10 | 0.24 | 0.23
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | emmetropic | VA | UNVA | Uncorrected | Near | LogRAD | Binocular |  | 수술전 | 명 | 10 | 0.48 | 0.14
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | emmetropic | VA | UNVA | Uncorrected | Near | LogRAD | Binocular |  | 6개월 | 명 | 10 | 0.18 | 0.11
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | myopic | VA | UNVA | Uncorrected | Near | LogRAD | Binocular |  | 수술전 | 명 | 10 | 0.04 | 0.19
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | myopic | VA | UNVA | Uncorrected | Near | LogRAD | Binocular |  | 6개월 | 명 | 10 | 0.12 | 0.18
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | hyperopic | CS |  |  |  | Log | Monocular |  | 수술전 | 안 | 20 | 1.82 | 0.16
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | hyperopic | CS |  |  |  | Log | Monocular |  | 6개월 | 안 | 20 | 1.69 | 0.23
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | hyperopic | CS |  |  |  | Log | Binocular |  | 수술전 | 명 | 10 | 1.94 | 0.13
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | hyperopic | CS |  |  |  | Log | Binocular |  | 6개월 | 명 | 10 | 1.89 | 0.12
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | emmetropic | CS |  |  |  | Log | Monocular |  | 수술전 | 안 | 20 | 1.85 | 0.16
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | emmetropic | CS |  |  |  | Log | Monocular |  | 6개월 | 안 | 20 | 1.73 | 0.16
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | emmetropic | CS |  |  |  | Log | Binocular |  | 수술전 | 명 | 10 | 1.92 | 0.09
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | emmetropic | CS |  |  |  | Log | Binocular |  | 6개월 | 명 | 10 | 1.92 | 0.09
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | myopic | CS |  |  |  | Log | Monocular |  | 수술전 | 안 | 20 | 1.86 | 0.16
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | myopic | CS |  |  |  | Log | Monocular |  | 6개월 | 안 | 20 | 1.74 | 0.22
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | myopic | CS |  |  |  | Log | Binocular |  | 수술전 | 명 | 10 | 1.95 | 0
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | myopic | CS |  |  |  | Log | Binocular |  | 6개월 | 명 | 10 | 1.89 | 0.12
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | all | CS |  |  |  | Log | Monocular |  | 수술전 | 안 | 20 | 1.85 | 0.16
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | all | CS |  |  |  | Log | Binocular |  | 수술전 | 명 | 30 | 1.94 | 0.09
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | all | CS |  |  |  | Log | Monocular | postoperative monocular performance was significantly reduced in all groups | 6개월 | 안 | 20 | 1.72 | 0.2
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | PresbyMAX |  | all | CS |  |  |  | Log | Binocular | binocular values were comparable to preoperative ones. | 6개월 | 명 | 30 | 1.9 | 0.11

## 결과지표_범주형_RS
3. 범주형 결과변수
no. | 1저자 (연도) | 연구설계 | 대상자 | 중재군명 | 대조군명 | 하위그룹 | 결과그룹 | 결과지표명 | 구분(나안, 교정( | 구분(거리) | 단위 | 안구기준 (Monocular, Binocular) | 정의 | 측정시점 (개월) | 분석단위(명/안) | 중재군 |  | 대조군 |  | 두 군간 차이 | OR, RR, HR |  | 95% CI | 비고
[16 empty] | Total N | event | Total N | event | P value | 단위 | risk rate
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | CDVA | corrected | Distant | Snellen | Binocular | lost 2 line | 6개월 | 명 | 30 | 6
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | CDVA | corrected | Distant | Snellen | Binocular | lost 3 line | 6개월 | 명 | 30 | 0
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | CDVA | corrected | Distant | Snellen | Binocular | lost 1 line | 6개월 | 명 | 30 | 16.98
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | CDVA | corrected | Distant | Snellen | Binocular | 변화없음 | 6개월 | 명 | 30 | 18
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | CDVA | corrected | Distant | Snellen | Binocular | Gain 1 line | 6개월 | 명 | 30 | 1.02
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | CDVA | corrected | Distant | Snellen | Binocular | Gain 2 line | 6개월 | 명 | 30 | 0
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | CDVA | corrected | Distant | Snellen | Binocular | Gain 3 line | 6개월 | 명 | 30 | 0
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | CNVA | corrected | Near | Snellen | Binocular | lost 2 line | 6개월 | 명 | 30 | 3
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | CNVA | corrected | Near | Snellen | Binocular | lost 3 line | 6개월 | 명 | 30 | 1.98
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | CNVA | corrected | Near | Snellen | Binocular | lost 1 line | 6개월 | 명 | 30 | 16.98
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | CNVA | corrected | Near | Snellen | Binocular | 변화없음 | 6개월 | 명 | 30 | 37.98
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | CNVA | corrected | Near | Snellen | Binocular | Gain 1 line | 6개월 | 명 | 30 | 0
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | CNVA | corrected | Near | Snellen | Binocular | Gain 2 line | 6개월 | 명 | 30 | 0
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | CNVA | corrected | Near | Snellen | Binocular | Gain 3 line | 6개월 | 명 | 30 | 0
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | DCNVA | corrected | Near | Snellen | Binocular | lost 2 line | 6개월 | 명 | 30 | 0
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | DCNVA | corrected | Near | Snellen | Binocular | lost 3 line | 6개월 | 명 | 30 | 0
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | DCNVA | corrected | Near | Snellen | Binocular | lost 1 line | 6개월 | 명 | 30 | 1.02
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | DCNVA | corrected | Near | Snellen | Binocular | 변화없음 | 6개월 | 명 | 30 | 9
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | DCNVA | corrected | Near | Snellen | Binocular | Gain 1 line | 6개월 | 명 | 30 | 7.98
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | DCNVA | corrected | Near | Snellen | Binocular | Gain 2 line | 6개월 | 명 | 30 | 13.98
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA | DCNVA | corrected | Near | Snellen | Binocular | Gain 3 line | 6개월 | 명 | 30 | 28.02
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA |  | uncorrected | Distant/Near | Snellen+Jarger | monocular | 20/25 or better and J3 | 수술전 | 안 | 60 | 0
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA |  | uncorrected | Distant/Near | Snellen+Jarger | monocular | 20/25 or better and J3 | 6개월 | 안 | 60 | 23.1
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA |  | uncorrected | Distant/Near | Snellen+Jarger | Binocular | 20/25 or better and J3 | 6개월 | 명 | 30 | 27.9
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA |  | uncorrected | Distant/Near | Snellen+Jarger | Binocular | 20/25 or better and J3 | 12개월 | 명 | 25 | 25
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | VA |  | uncorrected | Distant/Near | Snellen+Jarger | Binocular | 20/25 or better and J1 | 12개월 | 명 | 25 | 18
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Satisfaction | Satisfied or very satified before surgery | uncorrected |  |  | monocular |  | 수술전 | 안 | 66 | 48
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Satisfaction | Reported sharpness and clarity did not diminish after aspheric treatment |  |  |  | monocular |  | 12개월 | 안 | 48 | 39.84
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Satisfaction | Somewhat or very dissatisfied or not sure before surgery | uncorrected |  |  | monocular |  | 수술전 | 안 | 66 | 18
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Satisfaction | Reported very satisfied or satisfied with the sharpness and clarity of vision after aspheric treatment |  |  |  | monocular |  | 12개월 | 안 | 18 | 14.4
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Satisfaction | “satisfied” or “very satisfied” with night vision under glare conditions | corrected |  |  | monocular |  | 수술전 | 안 | 66 | 36.96
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Satisfaction | “satisfied” or “very satisfied” with night vision under glare conditions | uncorrected |  |  | monocular |  | 12개월 | 안 | 66 | 39
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Reading |  |  |  | monocular | Y/N | 수술전 | 안 | 66 | 66
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Reading |  |  |  | monocular | Y/N | 6개월 | 안 | 60 | 36
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Reading |  |  |  | monocular | Y/N | 12개월 | 안 | 50 | 34
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Computer |  |  |  | monocular | Y/N | 수술전 | 안 | 66 | 66
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Computer |  |  |  | monocular | Y/N | 6개월 | 안 | 60 | 12.6
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Computer |  |  |  | monocular | Y/N | 12개월 | 안 | 50 | 24
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Writing Checks |  |  |  | monocular | Y/N | 수술전 | 안 | 66 | 66
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Writing Checks |  |  |  | monocular | Y/N | 6개월 | 안 | 60 | 10.2
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Writing Checks |  |  |  | monocular | Y/N | 12개월 | 안 | 50 | 14
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Driving Day |  |  |  | monocular | Y/N | 수술전 | 안 | 66 | 50.16
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Driving Day |  |  |  | monocular | Y/N | 6개월 | 안 | 60 | 0
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Driving Day |  |  |  | monocular | Y/N | 12개월 | 안 | 50 | 0
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Driving Night |  |  |  | monocular | Y/N | 수술전 | 안 | 66 | 54.12
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Driving Night |  |  |  | monocular | Y/N | 6개월 | 안 | 60 | 1.8
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Driving Night |  |  |  | monocular | Y/N | 12개월 | 안 | 50 | 0
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Leaving Home |  |  |  | monocular | Y/N | 수술전 | 안 | 66 | 46.2
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Leaving Home |  |  |  | monocular | Y/N | 6개월 | 안 | 60 | 0
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Leaving Home |  |  |  | monocular | Y/N | 12개월 | 안 | 50 | 0
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Recreation |  |  |  | monocular | Y/N | 수술전 | 안 | 66 | 50.16
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Recreation |  |  |  | monocular | Y/N | 6개월 | 안 | 60 | 0
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | Aspheric Wavefront-guided LASIK |  |  | Spectacle dependence | Recreation |  |  |  | monocular | Y/N | 12개월 | 안 | 50 | 0
329 | Khalifa (2011) | Prospective consecutive comparative randomized double-masked study | Myopic presbyopia | all high-order aberrations were treated | Vertical coma aberrations were selectively left untreated, while all other HOAs were treated |  | VA | UCVA | uncorrected | Distant | decimal scale | monocular | 20/30 or better | 수술후 | 안 | 30 | 30 | 22 | 22
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | distance corrected | Near | LogMAR | Binocular | lost ≥ 3 line | 12개월 | 명 | 29 | 0
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | distance corrected | Near | LogMAR | Binocular | lost 2 line | 12개월 | 명 | 29 | 2.03
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | distance corrected | Near | LogMAR | Binocular | lost 1 line | 12개월 | 명 | 29 | 0.87
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | distance corrected | Near | LogMAR | Binocular | 변화없음 | 12개월 | 명 | 29 | 2.03
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | distance corrected | Near | LogMAR | Binocular | Gain 1 line | 12개월 | 명 | 29 | 2.03
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | distance corrected | Near | LogMAR | Binocular | Gain 2 line | 12개월 | 명 | 29 | 4.06
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | DCNVA | distance corrected | Near | LogMAR | Binocular | Gain ≥ 3 line | 12개월 | 명 | 29 | 17.98
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | corrected | Distant | LogMAR | Binocular | lost ≥ 3 line | 12개월 | 명 | 29 | 6.09
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | corrected | Distant | LogMAR | Binocular | lost 2 line | 12개월 | 명 | 29 | 2.9
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | corrected | Distant | LogMAR | Binocular | lost 1 line | 12개월 | 명 | 29 | 8.12
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | corrected | Distant | LogMAR | Binocular | 변화없음 | 12개월 | 명 | 29 | 8.99
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | corrected | Distant | LogMAR | Binocular | Gain 1 line | 12개월 | 명 | 29 | 0.87
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | corrected | Distant | LogMAR | Binocular | Gain 2 line | 12개월 | 명 | 29 | 2.03
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | VA | CDVA | corrected | Distant | LogMAR | Binocular | Gain ≥ 3 line | 12개월 | 명 | 29 | 0
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | Spectacle dependence | Distance vision |  |  |  |  | Y/N | 12개월 | 명 | 29 | 7
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | Spectacle dependence | Near vision |  |  |  |  | Y/N | 12개월 | 명 | 29 | 3
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | PresbyMAX (Hybrid micro-Monovision) | PresbyMAX (micro-Monovision) |  | Satisfaction | Would undergo the same procedure again |  |  |  |  | 재시술의향 |  | 명 | 29 | 24
457 | Xu (2020) | Prospective, non-randomized, non-comparative study | Emmetropic presbyopia | LaserACE |  |  | Spectacle dependence | Near and intermediate task | [5 empty] | 6개월 | 명 | 4 | 0
457 | Xu (2020) | Prospective, non-randomized, non-comparative study | Emmetropic presbyopia | LaserACE |  |  | Spectacle dependence | Near tasks (reading books for a long time) | [5 empty] | 12개월 | 명 | 4 | 1
457 | Xu (2020) | Prospective, non-randomized, non-comparative study | Emmetropic presbyopia | LaserACE |  |  | Complication | Intraocular pressure | [6 empty] | 안 | 8 | 0
457 | Xu (2020) | Prospective, non-randomized, non-comparative study | Emmetropic presbyopia | LaserACE |  |  | Complication | Subconjunctival hemorrhage | [6 empty] | 안 | 8 | 1
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 0.0 logMAR (20/20) or better |  | 명 | 34 | 27
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | UNVA | uncorrected | Near | logRAD | Binocular | 0.1 logRAD or better |  | 명 | 34 | 27
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA |  | uncorrected | Distant/Near | LogMAR+LogRAD | Binocular | at least 0.0 logMAR (20/20) of binocular UDVA and 0.1 logRAD of UNVA |  | 명 | 34 | 21.76
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | CDVA | corrected | Distant |  | Binocular | 0.1 logMAR (20/25) or better |  | 명 | 34 | 34
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  | distance eye | VA | CDVA | corrected | Distant |  | Binocular | lost 1 line |  | 명 | 34 | 2
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  | near eye | VA | CDVA | corrected | Distant |  | monocular | lost 1 line |  | 안 | 34 | 4
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  | near eye | VA | CDVA | corrected | Distant |  | monocular | lost 2 line |  | 안 | 34 | 3
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | CDVA | corrected | Distant |  | Binocular | lost 1 line |  | 명 | 34 | 1
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | VA | CDVA | corrected | Distant |  | Binocular | lost 2 line |  | 명 | 34 | 0
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | PresbyMAX with micro-Monovision |  |  | Complication | Intraoperative complication | [6 empty] | 안 | 34 | 0
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | VA | CDVA | corrected | Distant | decimal scale | Binocular | >20/20 | 6개월 | 명 | 19 | 19
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | VA | CDVA | corrected | Distant | decimal scale | Binocular | >20/25 | 6개월 | 명 | 19 | 19
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | VA | CDVA | corrected | Distant | decimal scale | Binocular | >20/30 | 6개월 | 명 | 19 | 19
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | VA | CDVA | corrected | Distant | decimal scale | Binocular | >20/40 | 6개월 | 명 | 19 | 19
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | VA | UDVA | uncorrected | Distant | decimal scale | Binocular | >20/20 | 6개월 | 명 | 19 | 5.3599
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | VA | UDVA | uncorrected | Distant | decimal scale | Binocular | >20/25 | 6개월 | 명 | 19 | 13.1898
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | VA | UDVA | uncorrected | Distant | decimal scale | Binocular | >20/30 | 6개월 | 명 | 19 | 19
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | VA | UDVA | uncorrected | Distant | decimal scale | Binocular | >20/40 | 6개월 | 명 | 19 | 19
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | VA | CDVA | corrected | Distant | decimal scale | Binocular | >20/20 | 6개월 | 명 | 14 | 14
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | VA | CDVA | corrected | Distant | decimal scale | Binocular | >20/25 | 6개월 | 명 | 14 | 14
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | VA | CDVA | corrected | Distant | decimal scale | Binocular | >20/30 | 6개월 | 명 | 14 | 14
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | VA | CDVA | corrected | Distant | decimal scale | Binocular | >20/40 | 6개월 | 명 | 14 | 14
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | VA | UDVA | uncorrected | Distant | decimal scale | Binocular | >20/20 | 6개월 | 명 | 14 | 7.1736
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | VA | UDVA | uncorrected | Distant | decimal scale | Binocular | >20/25 | 6개월 | 명 | 14 | 10.9998
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | VA | UDVA | uncorrected | Distant | decimal scale | Binocular | >20/30 | 6개월 | 명 | 14 | 14
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | VA | UDVA | uncorrected | Distant | decimal scale | Binocular | >20/40 | 6개월 | 명 | 14 | 14
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.00D) | VA | CDVA | corrected | Distant | decimal scale | Binocular | >20/20 | 6개월 | 명 | 17 | 17
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.00D) | VA | CDVA | corrected | Distant | decimal scale | Binocular | >20/25 | 6개월 | 명 | 17 | 17
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.00D) | VA | CDVA | corrected | Distant | decimal scale | Binocular | >20/30 | 6개월 | 명 | 17 | 17
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.00D) | VA | CDVA | corrected | Distant | decimal scale | Binocular | >20/40 | 6개월 | 명 | 17 | 17
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.00D) | VA | UDVA | uncorrected | Distant | decimal scale | Binocular | >20/20 | 6개월 | 명 | 17 | 5.9857
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.00D) | VA | UDVA | uncorrected | Distant | decimal scale | Binocular | >20/25 | 6개월 | 명 | 17 | 10.999
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.00D) | VA | UDVA | uncorrected | Distant | decimal scale | Binocular | >20/30 | 6개월 | 명 | 17 | 17
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.00D) | VA | UDVA | uncorrected | Distant | decimal scale | Binocular | >20/40 | 6개월 | 명 | 17 | 17
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | VA | CNVA | corrected | Near(33cm) | Jaeger scale | Binocular | J1 or better | 6개월 | 명 | 19 | 19
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | VA | CNVA | corrected | Near(33cm) | Jaeger scale | Binocular | J2 or better | 6개월 | 명 | 19 | 19
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | VA | CNVA | corrected | Near(33cm) | Jaeger scale | Binocular | J4 or better | 6개월 | 명 | 19 | 19
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | VA | UNVA | uncorrected | Near(33cm) | Jaeger scale | Binocular | J1 or better | 6개월 | 명 | 19 | 3.0837
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | VA | UNVA | uncorrected | Near(33cm) | Jaeger scale | Binocular | J2 or better | 6개월 | 명 | 19 | 16.9993
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | VA | UNVA | uncorrected | Near(33cm) | Jaeger scale | Binocular | J4 or better | 6개월 | 명 | 19 | 19
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | VA | CNVA | corrected | Near(33cm) | Jaeger scale | Binocular | J1 or better | 6개월 | 명 | 14 | 14
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | VA | CNVA | corrected | Near(33cm) | Jaeger scale | Binocular | J2 or better | 6개월 | 명 | 14 | 14
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | VA | CNVA | corrected | Near(33cm) | Jaeger scale | Binocular | J4 or better | 6개월 | 명 | 14 | 14
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | VA | UNVA | uncorrected | Near(33cm) | Jaeger scale | Binocular | J1 or better | 6개월 | 명 | 14 | 1.1522
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | VA | UNVA | uncorrected | Near(33cm) | Jaeger scale | Binocular | J2 or better | 6개월 | 명 | 14 | 12.999
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group2(spherical equivalent less than -3.00D) | VA | UNVA | uncorrected | Near(33cm) | Jaeger scale | Binocular | J4 or better | 6개월 | 명 | 14 | 14
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.00D) | VA | CNVA | corrected | Near(33cm) | Jaeger scale | Binocular | J1 or better | 6개월 | 명 | 17 | 17
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.00D) | VA | CNVA | corrected | Near(33cm) | Jaeger scale | Binocular | J2 or better | 6개월 | 명 | 17 | 17
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.00D) | VA | CNVA | corrected | Near(33cm) | Jaeger scale | Binocular | J4 or better | 6개월 | 명 | 17 | 17
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.00D) | VA | UNVA | uncorrected | Near(33cm) | Jaeger scale | Binocular | J1 or better | 6개월 | 명 | 17 | 0.9996
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.00D) | VA | UNVA | uncorrected | Near(33cm) | Jaeger scale | Binocular | J2 or better | 6개월 | 명 | 17 | 15.9987
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.00D) | VA | UNVA | uncorrected | Near(33cm) | Jaeger scale | Binocular | J4 or better | 6개월 | 명 | 17 | 17
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.00D) | VA | CDVA | corrected | Distant |  |  | lost 2 line |  | 명 | 17 | 1
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.00D) | VA | CNVA | corrected | Near(33cm) |  |  | lost lines |  | 명 | 17 | 0
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  |  | Complication | Retreatment for residual hyperopia | [6 empty] | 명 | 50 | 4
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group3(spherical equivalent greater than -3.00D) | Complication | Retreatment for residual hyperopia | [6 empty] | 명 | 17 | 1
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  | Group1(refraction +0.50 D and +3.00D) | Complication | Retreatment for residual hyperopia | [6 empty] | 명 | 19 | 3
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  |  | Complication | Repositioning of corneal flap | [6 empty] | 명 | 50 | 0
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  |  | Complication | Case of diffuse lamellar kertitis, bacterial keratitis, or haze | [6 empty] | 명 | 50 | 0
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | Presbyopia correction using micro-Monovision and aspheric aberration patterns |  |  | Complication | Dry eye | [6 empty] | 명 | 50 | 2
734 | Taneri (2019) | Prospective, non-randomized, comparative case series | Hyperopic presbyopia | Varifocal LASIK |  |  | VA | UDVA | uncorrected | Distant(3m) | snellen line | monocular | 20/20 or better | 3개월 | 안 | 16 | 7 | 14 | 10
734 | Taneri (2019) | Prospective, non-randomized, comparative case series | Hyperopic presbyopia | Varifocal LASIK |  |  | VA | UDVA | uncorrected | Distant(3m) | snellen line | monocular | lost 2 or more line | 3개월 | 안 | 16 | 3 | 14 | 0
734 | Taneri (2019) | Prospective, non-randomized, comparative case series | Hyperopic presbyopia | Varifocal LASIK |  |  | VA | CDVA | corrected | Distant(3m) | snellen line |  | lost one or more linse | 3개월 | 안구 | 16 | 0 | 14 | 0
734 | Taneri (2019) | Prospective, non-randomized, comparative case series | Hyperopic presbyopia | Varifocal LASIK |  |  | Complication | Retreatment | [5 empty] | 3개월 | 안구 | 16 | 0 | 14 | 0
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | Complication | Side effect, a single complication | [5 empty] | 6개월 | 안 | 72 | 0
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | dominant eye | VA | UNVA | uncorrected | Near |  | monocular | 20/20 or better | 6개월 | 안 | 36 | 12.96
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | non-dominant eye | VA | UNVA | uncorrected | Near |  | monocular | 20/20 or better | 6개월 | 안 | 36 | 23.04
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | dominant eye | VA | UNVA | uncorrected | Near |  | monocular | 20/40 or better | 6개월 | 안 | 36 | 33.12
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | non-dominant eye | VA | UNVA | uncorrected | Near |  | monocular | 20/40 or better | 6개월 | 안 | 36 | 34.92
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | dominant eye | VA | UIVA | uncorrected | Intermediate |  | monocular | 20/40 or better | 6개월 | 안 | 36 | 36
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | non-dominant eye | VA | UIVA | uncorrected | Intermediate |  | monocular | 20/40 or better | 6개월 | 안 | 36 | 36
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | dominant eye | VA | UDVA | uncorrected | Distant |  | monocular | 20/40 or better | 6개월 | 안 | 36 | 36
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  | non-dominant eye | VA | UDVA | uncorrected | Distant |  | monocular | 20/40 or better | 6개월 | 안 | 36 | 36
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UNVA | uncorrected | Near |  | Binocular | 20/20 | 6개월 | 명 | 36 | 28.08
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UIVA | uncorrected | Intermediate |  | Binocular | 20/20 | 6개월 | 명 | 36 | 33.12
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | multifocal presbyLASIK |  |  | VA | UDVA | uncorrected | Distant |  | Binocular | 20/20 | 6개월 | 명 | 36 | 30.96
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | VA | UDVA | uncorrected | Distant | LogMAR |  | 0.10 logMAR | 12개월 | 안 | 39 | 30
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | dominant eye with emmetropic correction | VA | UDVA | uncorrected | Distant | LogMAR |  | 0.10 logMAR | 12개월 | 안 | 19 | 18
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | VA | CDVA | corrected | Distant |  |  | lost 2 lines | 12개월 | 안 | 39 | 4
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | VA | UNVA | uncorrected | Near(40cm) |  |  | ≥0.4 | 12개월 | 안 | 39 | 0
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | VA | UNVA | uncorrected | Near(40cm) |  |  | ≥0.4 | 12개월 | 안 | 39 | 34
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Complication | Glare |  |  |  |  | no side effect | 12개월 | 명 | 20 | 4
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Complication | Halos |  |  |  |  | no side effect | 12개월 | 명 | 20 | 9
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Complication | Double vision |  |  |  |  | no side effect | 12개월 | 명 | 20 | 15
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Complication | Fluctuations in vision |  |  |  |  | no side effect | 12개월 | 명 | 20 | 5
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Complication | Headaches |  |  |  |  | no side effect | 12개월 | 명 | 20 | 17
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Complication | Artificial tears every day | [5 empty] | 12개월 | 명 | 20 | 17
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Complication | Artificial tears more than before treatment | [5 empty] | 12개월 | 명 | 20 | 12
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Complication | Reduced distance vision |  |  |  |  | no side effect | 12개월 | 명 | 20 | 6
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Satisfaction | Would undergo SUPRACOR again |  |  | Yes/No |  |  |  | 명 | 20 | 15
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 YES | VA | CDVA | corrected | Distant | line | Binocular | loss of 2 line | 12개월 | 명 | 15 | 2
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 No | VA | CDVA | corrected | Distant | line | binocular | loss of 2 line | 12개월 | 명 | 5 | 1
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 YES | Spectacle dependence | Daily use of distance glasses | [5 empty] | 12개월 | 명 | 15 | 1
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 No | Spectacle dependence | Daily use of distance glasses | [5 empty] | 12개월 | 명 | 5 | 0
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 YES | Spectacle dependence | Daily use of reading glasses | [5 empty] | 12개월 | 명 | 15 | 6
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  | 재시술 의향 No | Spectacle dependence | Daily use of reading glasses | [5 empty] | 12개월 | 명 | 5 | 2
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Spectacle dependence | Daily use of reading glasses | [5 empty] | 12개월 | 명 | 20 | 8
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | SUPRACOR |  |  | Spectacle dependence | Daily use of distance glasses | [5 empty] | 12개월 | 명 | 20 | 1
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/12.5 | 1개월 | 명 | 69 | 0
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/12.5 | 3개월 | 명 | 69 | 0
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/12.5 | 6개월 | 명 | 69 | 0
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/12.5 | 9개월 | 명 | 69 | 0
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/12.5 | 12개월 | 명 | 69 | 0
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/16 | 1개월 | 명 | 69 | 4.2435
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/16 | 3개월 | 명 | 69 | 2.8152
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/16 | 6개월 | 명 | 69 | 0
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/16 | 9개월 | 명 | 69 | 12.1785
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/16 | 12개월 | 명 | 69 | 25.0884
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/20 | 1개월 | 명 | 69 | 31.8435
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/20 | 3개월 | 명 | 69 | 38.019
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/20 | 6개월 | 명 | 69 | 43.6977
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/20 | 9개월 | 명 | 69 | 56.8215
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/20 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/25 | 1개월 | 명 | 69 | 50.9565
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/25 | 3개월 | 명 | 69 | 56.3247
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/25 | 6개월 | 명 | 69 | 57.4977
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/25 | 9개월 | 명 | 69 | 64.9428
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/25 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/32 | 1개월 | 명 | 69 | 63.6939
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/32 | 3개월 | 명 | 69 | 64.7772
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/32 | 6개월 | 명 | 69 | 62.1
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/32 | 9개월 | 명 | 69 | 64.9428
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/32 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/40 | 1개월 | 명 | 69 | 66.8748
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/40 | 3개월 | 명 | 69 | 66.1848
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/40 | 6개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/40 | 9개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/40 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/50 | 1개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/50 | 3개월 | 명 | 69 | 67.5924
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/50 | 6개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/50 | 9개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/50 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/63 | 1개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/63 | 3개월 | 명 | 69 | 67.5924
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/63 | 6개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/63 | 9개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | 20/63 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | >20/63 | 1개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | >20/63 | 3개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | >20/63 | 6개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | >20/63 | 9개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UDVA | uncorrected | Distant | LogMAR | Binocular | >20/63 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | ≤J1 | 1개월 | 명 | 69 | 0
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | ≤J1 | 3개월 | 명 | 69 | 2.9394
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | ≤J1 | 6개월 | 명 | 69 | 2.553
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | ≤J1 | 9개월 | 명 | 69 | 0
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | ≤J1 | 12개월 | 명 | 69 | 0
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | <J1 | 1개월 | 명 | 69 | 4.4988
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | <J1 | 3개월 | 명 | 69 | 7.3416
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | <J1 | 6개월 | 명 | 69 | 7.6659
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | <J1 | 9개월 | 명 | 69 | 0
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | <J1 | 12개월 | 명 | 69 | 0
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J1 | 1개월 | 명 | 69 | 33.0027
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J1 | 3개월 | 명 | 69 | 17.4777
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J1 | 6개월 | 명 | 69 | 25.5576
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J1 | 9개월 | 명 | 69 | 24.3501
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J1 | 12개월 | 명 | 69 | 20.7
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J2 | 1개월 | 명 | 69 | 33.0027
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J2 | 3개월 | 명 | 69 | 42.573
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J2 | 6개월 | 명 | 69 | 46.0023
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J2 | 9개월 | 명 | 69 | 36.5286
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J2 | 12개월 | 명 | 69 | 48.3
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J3 | 1개월 | 명 | 69 | 50.9979
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J3 | 3개월 | 명 | 69 | 58.7259
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J3 | 6개월 | 명 | 69 | 51.1083
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J3 | 9개월 | 명 | 69 | 44.6499
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J3 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J4 | 1개월 | 명 | 69 | 64.5012
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J4 | 3개월 | 명 | 69 | 66.0606
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J4 | 6개월 | 명 | 69 | 63.8871
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J4 | 9개월 | 명 | 69 | 60.8856
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J4 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J5 | 1개월 | 명 | 69 | 64.5012
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J5 | 3개월 | 명 | 69 | 67.5303
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J5 | 6개월 | 명 | 69 | 66.447
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J5 | 9개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J5 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J6 | 1개월 | 명 | 69 | 65.9985
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J6 | 3개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J6 | 6개월 | 명 | 69 | 66.447
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J6 | 9개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | J6 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | >J8 | 1개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | >J8 | 3개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | >J8 | 6개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | >J8 | 9개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UNVA | uncorrected | Near(33cm) | LogMAR | Binocular | >J8 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/12.5 | 1개월 | 명 | 69 | 20.9967
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/12.5 | 3개월 | 명 | 69 | 0
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/12.5 | 6개월 | 명 | 69 | 22.9977
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/12.5 | 9개월 | 명 | 69 | 0
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/12.5 | 12개월 | 명 | 69 | 38.3364
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/16 | 1개월 | 명 | 69 | 30.0012
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/16 | 3개월 | 명 | 69 | 2.8152
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/16 | 6개월 | 명 | 69 | 30.6636
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/16 | 9개월 | 명 | 69 | 12.1785
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/16 | 12개월 | 명 | 69 | 46.0023
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/20 | 1개월 | 명 | 69 | 45.0018
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/20 | 3개월 | 명 | 69 | 38.019
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/20 | 6개월 | 명 | 69 | 61.3341
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/20 | 9개월 | 명 | 69 | 56.8215
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/20 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/25 | 1개월 | 명 | 69 | 65.9985
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/25 | 3개월 | 명 | 69 | 56.3247
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/25 | 6개월 | 명 | 69 | 65.1636
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/25 | 9개월 | 명 | 69 | 64.9428
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/25 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/32 | 1개월 | 명 | 69 | 65.9985
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/32 | 3개월 | 명 | 69 | 64.7772
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/32 | 6개월 | 명 | 69 | 65.1636
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/32 | 9개월 | 명 | 69 | 64.9428
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/32 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/40 | 1개월 | 명 | 69 | 65.9985
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/40 | 3개월 | 명 | 69 | 66.1848
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/40 | 6개월 | 명 | 69 | 65.1636
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/40 | 9개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/40 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/50 | 1개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/50 | 3개월 | 명 | 69 | 67.5924
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/50 | 6개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/50 | 9개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/50 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/63 | 1개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/63 | 3개월 | 명 | 69 | 67.5924
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/63 | 6개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/63 | 9개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | 20/63 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | >20/63 | 1개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | >20/63 | 3개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | >20/63 | 6개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | >20/63 | 9개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | UIVA | uncorrected | Intermediate(60cm) | LogMAR | Binocular | >20/63 | 12개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | CDVA | corrected | Distant |  |  | lost 2 lines | 1개월 | 안 | 84 | 1.0248
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | CDVA | corrected | Distant |  |  | lost 1 line | 1개월 | 안 | 84 | 5.124
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | CDVA | corrected | Distant |  |  | 변화없음 | 1개월 | 안 | 84 | 61.4628
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | CDVA | corrected | Distant |  |  | gain 1 line | 1개월 | 안 | 84 | 13.314
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | CDVA | corrected | Distant |  |  | Gain 2 line | 1개월 | 안 | 84 | 3.0744
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | VA | CDVA | corrected | Distant |  |  | gain 1 line or more | 12개월 | 안 | 82 | 16
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | Complication | Retreatment | [5 empty] | 12개월 | 명 | 69 | 16
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | Satisfaction | Everyday activities | [5 empty] | 1개월 | 명 | 69 | 63
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | Satisfaction | Everyday activities | [5 empty] | 3개월 | 명 | 69 | 66
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | Satisfaction | Everyday activities | [5 empty] | 9개월 | 명 | 69 | 69
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | Satisfaction | Recommend the surgery | [5 empty] | 1개월 | 명 | 69 | 62
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | central presbyopic LASIK with Custom Q algorithm |  |  | Satisfaction | Recommend the surgery | [5 empty] | 6개월 | 명 | 69 | 69
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | preop-CDVA | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/12.5 or better | 수술전 | 안 | 80 | 14.4
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | preop-CDVA | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/16 or better | 수술전 | 안 | 80 | 68
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | preop-CDVA | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/20 or better | 수술전 | 안 | 80 | 75.2
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | preop-CDVA | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/25 or better | 수술전 | 안 | 80 | 80
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | preop-CDVA | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/32 or better | 수술전 | 안 | 80 | 80
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | preop-CDVA | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/40 or better | 수술전 | 안 | 80 | 80
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | preop-CDVA | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/63 or better | 수술전 | 안 | 80 | 80
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | VA | UDVA | uncorrected | Distant(3m) |  | Binocular | 20/12.5 or better | 3개월 | 명 | 40 | 16
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | VA | UDVA | uncorrected | Distant(3m) |  | Binocular | 20/16 or better | 3개월 | 명 | 40 | 38
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | VA | UDVA | uncorrected | Distant(3m) |  | Binocular | 20/20 or better | 3개월 | 명 | 40 | 39.2
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | VA | UDVA | uncorrected | Distant(3m) |  | Binocular | 20/25 or better | 3개월 | 명 | 40 | 40
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | VA | UDVA | uncorrected | Distant(3m) |  | Binocular | 20/32 or better | 3개월 | 명 | 40 | 40
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | VA | UDVA | uncorrected | Distant(3m) |  | Binocular | 20/40 or better | 3개월 | 명 | 40 | 40
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | VA | UDVA | uncorrected | Distant(3m) |  | Binocular | 20/63 or better | 3개월 | 명 | 40 | 40
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | distance eye | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/12.5 or better | 3개월 | 안 | 40 | 13.2
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | distance eye | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/16 or better | 3개월 | 안 | 40 | 35.2
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | distance eye | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/20 or better | 3개월 | 안 | 40 | 39.2
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | distance eye | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/25 or better | 3개월 | 안 | 40 | 40
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | distance eye | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/32 or better | 3개월 | 안 | 40 | 40
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | distance eye | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/40 or better | 3개월 | 안 | 40 | 40
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | distance eye | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/63 or better | 3개월 | 안 | 40 | 40
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | near eye | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/12.5 or better | 3개월 | 안 | 40 | 0
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | near eye | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/16 or better | 3개월 | 안 | 40 | 0
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | near eye | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/20 or better | 3개월 | 안 | 40 | 4
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | near eye | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/25 or better | 3개월 | 안 | 40 | 10
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | near eye | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/32 or better | 3개월 | 안 | 40 | 23.2
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | near eye | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/40 or better | 3개월 | 안 | 40 | 35.2
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | near eye | VA | UDVA | uncorrected | Distant(3m) |  | monocular | 20/63 or better | 3개월 | 안 | 40 | 40
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | distance eye | VA | Difference between UDVA and CDVA | NA | Distant(3m) |  | monocular | 3 or more worse | 3개월 | 안 | 40 | 0
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | distance eye | VA | Difference between UDVA and CDVA | NA | Distant(3m) |  | monocular | 2 worse | 3개월 | 안 | 40 | 1
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | distance eye | VA | Difference between UDVA and CDVA | NA | Distant(3m) |  | monocular | 1 worse | 3개월 | 안 | 40 | 16
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | distance eye | VA | Difference between UDVA and CDVA | NA | Distant(3m) |  | monocular | same | 3개월 | 안 | 40 | 21
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | distance eye | VA | Difference between UDVA and CDVA | NA | Distant(3m) |  | monocular | 1 better | 3개월 | 안 | 40 | 1
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | distance eye | VA | Difference between UDVA and CDVA | NA | Distant(3m) |  | monocular | 2 better | 3개월 | 안 | 40 | 1
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  | distance eye | VA | Difference between UDVA and CDVA | NA | Distant(3m) |  | monocular | 3 or more better | 3개월 | 안 | 40 | 0
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | VA | change in CDVA | corrected | Distant(3m) | snellen line | monocular | loss 3 line or more | 3개월 | 안 | 80 | 0
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | VA | change in CDVA | corrected | Distant(3m) | snellen line | monocular | loss 2 line | 3개월 | 안 | 80 | 0
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | VA | change in CDVA | corrected | Distant(3m) | snellen line | monocular | loss 1 line | 3개월 | 안 | 80 | 6
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | VA | change in CDVA | corrected | Distant(3m) | snellen line | monocular | 변화 | 3개월 | 안 | 80 | 49.8
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | VA | change in CDVA | corrected | Distant(3m) | snellen line | monocular | gain 1 line | 3개월 | 안 | 80 | 24
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | VA | change in CDVA | corrected | Distant(3m) | snellen line | monocular | Gain 2 line | 3개월 | 안 | 80 | 1
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | VA | change in CDVA | corrected | Distant(3m) | snellen line | monocular | gain 3 line or more | 3개월 | 안 | 80 | 0
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | VA | CDVA | corrected | Distant(3m) |  |  | 20/25 or better | 3개월 | 안 | 80 | 80
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | Complication | Diffuse lamellar keratitis | [5 empty] | 1주 | 안 | 80 | 1
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | Complication | Severe glare or double vision | [5 empty] | 3개월 | 안 | 80 | 0
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | Spectacle dependence | Near spectacles | [5 empty] | 3개월 | 명 | 40 | 1
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | Aspheric Micro-monovision LASIK (Laser Blended Vision) |  |  | Spectacle dependence | Distance vision | [5 empty] | 3개월 | 명 | 40 | 0
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | asymmetrical presbyLASIK | symmetrical presbyLASIK |  | Satisfaction | Would you repeat the procecure |  |  |  |  | Y/N | 18개월 | 명 | 14 | 9.94 | 16 | 9.6
1225 | Cosar (2014) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | monocular | 20/20 or better | 수술전 | 안 | 123 | 7
1225 | Cosar (2014) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | monocular | 20/25 or better | 수술전 | 안 | 123 | 12
1225 | Cosar (2014) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | monocular | 20/20 or better | 1개월 | 안 | 123 | 19
1225 | Cosar (2014) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | monocular | 20/25 or better | 1개월 | 안 | 123 | 38
1225 | Cosar (2014) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | monocular | 20/20 or better | 6개월 | 안 | 123 | 27
1225 | Cosar (2014) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | monocular | 20/25 or better | 6개월 | 안 | 123 | 45
1225 | Cosar (2014) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | 20/20 or better | 수술전 | 안 | 123 | 123
1225 | Cosar (2014) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | 20/20 or better | 1개월 | 안 | 123 | 90
1225 | Cosar (2014) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | 20/25 or better | 1개월 | 안 | 123 | 116
1225 | Cosar (2014) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | 20/20 or better | 6개월 | 안 | 123 | 76
1225 | Cosar (2014) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | 20/25 or better | 6개월 | 안 | 123 | 110
1225 | Cosar (2014) | [5 empty] | VA | BSCVA | best spectacle-corrected | Distant | Snellen | monocular | lost 1 line | 1개월 | 안 | 123 | 26
1225 | Cosar (2014) | [5 empty] | VA | BSCVA | best spectacle-corrected | Distant | Snellen | monocular | lost 2 line | 1개월 | 안 | 123 | 7
1225 | Cosar (2014) | [5 empty] | VA | BSCVA | best spectacle-corrected | Distant | Snellen | monocular | lost 1 line | 6개월 | 안 | 123 | 35
1225 | Cosar (2014) | [5 empty] | VA | BSCVA | best spectacle-corrected | Distant | Snellen | monocular | lost 2 line | 6개월 | 안 | 123 | 13
1225 | Cosar (2014) | [5 empty] | VA | UNVA | uncorrected | Near | Snellen | monocular | 20/20 or better | 수술전 | 안 | 123 | 0
1225 | Cosar (2014) | [5 empty] | VA | UNVA | uncorrected | Near | Snellen | monocular | 20/25 or better | 수술전 | 안 | 123 | 5
1225 | Cosar (2014) | [5 empty] | VA | UNVA | uncorrected | Near | Snellen | monocular | 20/20 or better | 1개월 | 안 | 123 | 97
1225 | Cosar (2014) | [5 empty] | VA | UNVA | uncorrected | Near | Snellen | monocular | 20/25 or better | 1개월 | 안 | 123 | 115
1225 | Cosar (2014) | [5 empty] | VA | UNVA | uncorrected | Near | Snellen | monocular | 20/20 or better | 6개월 | 안 | 123 | 95
1225 | Cosar (2014) | [5 empty] | VA | UNVA | uncorrected | Near | Snellen | monocular | 20/25 or better | 6개월 | 안 | 123 | 110
1225 | Cosar (2014) | [5 empty] | VA | CNVA | corrected | near | Snellen | monocular | 20/20 or better | 수술전 | 안 | 123 | 123
1225 | Cosar (2014) | [5 empty] | VA | CNVA | corrected | near | Snellen | monocular | 20/20 or better | 1개월 | 안 | 123 | 117
1225 | Cosar (2014) | [5 empty] | VA | CNVA | corrected | near | Snellen | monocular | 20/25 or better | 1개월 | 안 | 123 | 123
1225 | Cosar (2014) | [5 empty] | VA | CNVA | corrected | near | Snellen | monocular | 20/20 or better | 6개월 | 안 | 123 | 117
1225 | Cosar (2014) | [5 empty] | VA | CNVA | corrected | near | Snellen | monocular | 20/25 or better | 6개월 | 안 | 123 | 123
1225 | Cosar (2014) | [5 empty] | VA | CNVA | corrected | near | Snellen | monocular | lost 1 line | 1개월 | 안 | 123 | 6
1225 | Cosar (2014) | [5 empty] | VA | CNVA | corrected | near | Snellen | monocular | lost 1 line | 6개월 | 안 | 123 | 6
1225 | Cosar (2014) | [5 empty] | Complication | Mild diffuse lamellar keratitis | [6 empty] | 안 | 123 | 5
1225 | Cosar (2014) | [5 empty] | Complication | Mild dryness | [6 empty] | 안 | 123 | 8
1225 | Cosar (2014) | [5 empty] | Complication | Moderate dryness | [6 empty] | 안 | 123 | 1
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | 6/4.8 or better | 수술전 | 안 | 46 | 17.94
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | monocular | 6/4.8 or better | 6개월 | 안 | 46 | 1.84
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | 6/6 or better | 수술전 | 안 | 46 | 43.24
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | monocular | 6/6 or better | 6개월 | 안 | 46 | 5.98
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | 6/7.5 or better | 수술전 | 안 | 46 | 44.16
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | monocular | 6/7.5 or better | 6개월 | 안 | 46 | 13.8
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | 6/9.5 or better | 수술전 | 안 | 46 | 46
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | monocular | 6/9.5 or better | 6개월 | 안 | 46 | 30.82
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | 6/12 or better | 수술전 | 안 | 46 | 46
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | monocular | 6/12 or better | 6개월 | 안 | 46 | 38.18
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | 6/15 or better | 수술전 | 안 | 46 | 46
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | monocular | 6/15 or better | 6개월 | 안 | 46 | 40.94
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | 6/19 or better | 수술전 | 안 | 46 | 46
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | monocular | 6/19 or better | 6개월 | 안 | 46 | 45.08
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | 6/24 or better | 수술전 | 안 | 46 | 46
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | monocular | 6/24 or better | 6개월 | 안 | 46 | 45.08
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | 6/30 or better | 수술전 | 안 | 46 | 46
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | monocular | 6/30 or better | 6개월 | 안 | 46 | 45.08
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | 6/38 | 수술전 | 안 | 46 | 46
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | monocular | 6/38 | 6개월 | 안 | 46 | 46
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | binocular | 6/4.8 or better | 수술전 | 명 | 23 | 11.96
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | binocular | 6/4.8 or better | 6개월 | 명 | 23 | 2.99
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | binocular | 6/6 or better | 수술전 | 명 | 23 | 22.08
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | binocular | 6/6 or better | 6개월 | 명 | 23 | 11.04
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | binocular | 6/7.5 or better | 수술전 | 명 | 23 | 23
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | binocular | 6/7.5 or better | 6개월 | 명 | 23 | 17.94
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | binocular | 6/9.5 or better | 수술전 | 명 | 23 | 23
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | binocular | 6/9.5 or better | 6개월 | 명 | 23 | 20.93
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | binocular | 6/12 or better | 수술전 | 명 | 23 | 23
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | binocular | 6/12 or better | 6개월 | 명 | 23 | 22.08
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | binocular | 6/15 or better | 수술전 | 명 | 23 | 23
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | binocular | 6/15 or better | 6개월 | 명 | 23 | 23
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | binocular | 6/19 or better | 수술전 | 명 | 23 | 23
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | binocular | 6/19 or better | 6개월 | 명 | 23 | 23
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | binocular | 6/24 or better | 수술전 | 명 | 23 | 23
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | binocular | 6/24 or better | 6개월 | 명 | 23 | 23
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | binocular | 6/30 or better | 수술전 | 명 | 23 | 23
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | binocular | 6/30 or better | 6개월 | 명 | 23 | 23
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | binocular | 6/38 | 수술전 | 명 | 23 | 23
1268 | Ryan (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | Snellen | binocular | 6/38 | 6개월 | 명 | 23 | 23
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | loss 3 or more | 3개월 | 안 | 46 | 1.012
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | loss 3 or more | 6개월 | 안 | 46 | 1.012
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | loss 2 | 3개월 | 안 | 46 | 1.012
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | loss 2 | 6개월 | 안 | 46 | 1.978
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | loss 1 | 3개월 | 안 | 46 | 13.018
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | loss 1 | 6개월 | 안 | 46 | 6.992
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | no change | 3개월 | 안 | 46 | 20.01
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | no change | 6개월 | 안 | 46 | 24.978
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | gain 1 | 3개월 | 안 | 46 | 9.982
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | gain 1 | 6개월 | 안 | 46 | 9.982
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | gain 2 | 3개월 | 안 | 46 | 1.012
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | gain 2 | 6개월 | 안 | 46 | 1.012
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | gain 3 or more | 3개월 | 안 | 46 | 0
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | monocular | gain 3 or more | 6개월 | 안 | 46 | 0
1268 | Ryan (2013) | [5 empty] | VA | CDVA | corrected | Distant | Snellen | binocular | losing 2 or more line | 6개월 | 명 | 23 | 1
1268 | Ryan (2013) |  |  |  |  | all of the time | Spectacle dependence | Reading glasses | [6 empty] | 명 | 23 | 2 | [6 empty] | Independence 보고를 dependence로 전환
1268 | Ryan (2013) |  |  |  |  | most of the time | Spectacle dependence | Reading glasses | [6 empty] | 명 | 23 | 1 | [6 empty] | Independence 보고를 dependence로 전환
1268 | Ryan (2013) |  |  |  |  | still dependent on reading glasses all the time | Spectacle dependence | Reading glasses | [6 empty] | 명 | 23 | 1 | [6 empty] | Independence 보고를 dependence로 전환
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | SUPRACOR |  | nondominant eye | Complication | Retreatment to enhance reading ability and was subsequently able to read N5 print unaided and free of reading glasses | [5 empty] | 6개월 이후 | 명 | 23 | 1
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | SUPRACOR |  | dominant eye | Complication | Retreatment to enhance the UDVA | [5 empty] | 6개월 이후 | 명 | 23 | 5
1268 | Ryan (2013) | [5 empty] | Complication | Dry eye | [5 empty] | 3개월 | 명 | 23 | 5
1268 | Ryan (2013) | [5 empty] | Complication | Dry eye | [5 empty] | 6개월 | 명 | 23 | 3
1268 | Ryan (2013) | [5 empty] | Complication | Sensitivity to light | [6 empty] | 명 | 23 | 4
1268 | Ryan (2013) | [5 empty] | Complication | Intraoperative complication | [6 empty] | 명 | 23 | 0
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | SUPRACOR |  |  | Satisfaction | Happy they had the procedure | [6 empty] | 명 | 23 | 21
1268 | Ryan (2013) | [5 empty] | Spectacle dependence | newpaper |  |  |  |  | self-reported ablility to read without glasses | 6개월 | 명 | 23 | 0.989 | [6 empty] | Independence 보고를 dependence로 전환
1268 | Ryan (2013) | [5 empty] | Spectacle dependence | moblie phone |  |  |  |  | self-reported ablility to read without glasses | 6개월 | 명 | 23 | 0.989 | [6 empty] | Independence 보고를 dependence로 전환
1268 | Ryan (2013) | [5 empty] | Spectacle dependence | menu |  |  |  |  | self-reported ablility to read without glasses | 6개월 | 명 | 23 | 2.001 | [6 empty] | Independence 보고를 dependence로 전환
1268 | Ryan (2013) | [5 empty] | Spectacle dependence | computer |  |  |  |  | self-reported ablility to read without glasses | 6개월 | 명 | 23 | 0.989 | [6 empty] | Independence 보고를 dependence로 전환
1268 | Ryan (2013) | [5 empty] | Spectacle dependence | small label |  |  |  |  | self-reported ablility to read without glasses | 6개월 | 명 | 23 | 0.989 | [6 empty] | Independence 보고를 dependence로 전환
1268 | Ryan (2013) | [5 empty] | Spectacle dependence | dim light |  |  |  |  | self-reported ablility to read without glasses | 6개월 | 명 | 23 | 4.991 | [6 empty] | Independence 보고를 dependence로 전환
1268 | Ryan (2013) | [5 empty] | Spectacle dependence | prolonged reading |  |  |  |  | self-reported ablility to read without glasses | 6개월 | 명 | 23 | 0.989 | [6 empty] | Independence 보고를 dependence로 전환
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | SUPRACOR |  |  | satisfaction | Satisfied or very satisfied with their UDVA | [6 empty] | 명 | 23 | 18
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | SUPRACOR |  |  | satisfaction | Not fully satisfied with their UDVA and had a retreatment in their dominant eye to improve UDVA without loss of reading ability. | [6 empty] | 명 | 23 | 5
1309 | Luger (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | LogMAR | binocular | -0.2 LogMAR | 1년 | 명 | 31 | 0
1309 | Luger (2013) | [5 empty] | VA | UNVA | uncorrected | Near | LogRAD | binocular | -0.2 LogRAD | 1년 | 명 | 31 | 0
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | DCNVA | distance corrected | Near | LogRAD | binocular | -0.2 LogRAD | 1년 | 명 | 31 | 0
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | U(D+N)VA | uncorrected | Distant+Near | LogMAR+LogRAD | binocular | -0.2 | 1년 | 명 | 31 | 0
1309 | Luger (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | LogMAR | binocular | -0.1 LogMAR | 1년 | 명 | 31 | 7.13
1309 | Luger (2013) | [5 empty] | VA | UNVA | uncorrected | Near | LogRAD | binocular | -0.1 LogRAD | 1년 | 명 | 31 | 3.1
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | DCNVA | distance corrected | Near | LogRAD | binocular | -0.1 LogRAD | 1년 | 명 | 31 | 1.24
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | U(D+N)VA | uncorrected | Distant+Near | LogMAR+LogRAD | binocular | -0.1 | 1년 | 명 | 31 | 0
1309 | Luger (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | LogMAR | binocular | 0.0 LogMAR | 1년 | 명 | 31 | 14.88
1309 | Luger (2013) | [5 empty] | VA | UNVA | uncorrected | Near | LogRAD | binocular | 0.0 LogRAD | 1년 | 명 | 31 | 17.05
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | DCNVA | distance corrected | Near | LogRAD | binocular | 0.0 LogRAD | 1년 | 명 | 31 | 4.65
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | U(D+N)VA | uncorrected | Distant+Near | LogMAR+LogRAD | binocular | 0.0 | 1년 | 명 | 31 | 8.37
1309 | Luger (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | LogMAR | binocular | 0.1 LogMAR | 1년 | 명 | 31 | 21.7
1309 | Luger (2013) | [5 empty] | VA | UNVA | uncorrected | Near | LogRAD | binocular | 0.1 LogRAD | 1년 | 명 | 31 | 26.04
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | DCNVA | distance corrected | Near | LogRAD | binocular | 0.1 LogRAD | 1년 | 명 | 31 | 13.64
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | U(D+N)VA | uncorrected | Distant+Near | LogMAR+LogRAD | binocular | 0.1 | 1년 | 명 | 31 | 18.91
1309 | Luger (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | LogMAR | binocular | 0.2 LogMAR | 1년 | 명 | 31 | 27.9
1309 | Luger (2013) | [5 empty] | VA | UNVA | uncorrected | Near | LogRAD | binocular | 0.2 LogRAD | 1년 | 명 | 31 | 29.14
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | DCNVA | distance corrected | Near | LogRAD | binocular | 0.2 LogRAD | 1년 | 명 | 31 | 26.35
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | U(D+N)VA | uncorrected | Distant+Near | LogMAR+LogRAD | binocular | 0.2 | 1년 | 명 | 31 | 26.35
1309 | Luger (2013) | [5 empty] | VA | UDVA | uncorrected | Distant | LogMAR | binocular | 0.3 LogMAR | 1년 | 명 | 31 | 31
1309 | Luger (2013) | [5 empty] | VA | UNVA | uncorrected | Near | LogRAD | binocular | 0.3 LogRAD | 1년 | 명 | 31 | 31
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | DCNVA | distance corrected | Near | LogRAD | binocular | 0.3 LogRAD | 1년 | 명 | 31 | 31
1309 | Luger (2013) | Prospective study | Presbyopia | Schwind PresbyMAX |  |  | VA | U(D+N)VA | uncorrected | Distant+Near | LogMAR+LogRAD | binocular | 0.3 | 1년 | 명 | 31 | 31
1309 | Luger (2013) | [5 empty] | VA | CDVA | corrected | Distant | snellen | monocular | loss 3 or more | 1년 | 안 | 62 | 0
1309 | Luger (2013) | [5 empty] | VA | CNVA | corrected | near | snellen | monocular | loss 3 or more | 1년 | 안 | 62 | 0
1309 | Luger (2013) | [5 empty] | VA | CDVA | corrected | Distant | snellen | monocular | loss 2 | 1년 | 안 | 62 | 1.86
1309 | Luger (2013) | [5 empty] | VA | CNVA | corrected | near | snellen | monocular | loss 2 | 1년 | 안 | 62 | 4.96
1309 | Luger (2013) | [5 empty] | VA | CDVA | corrected | Distant | snellen | monocular | loss 1 | 1년 | 안 | 62 | 20.46
1309 | Luger (2013) | [5 empty] | VA | CNVA | corrected | near | snellen | monocular | loss 1 | 1년 | 안 | 62 | 14.26
1309 | Luger (2013) | [5 empty] | VA | CDVA | corrected | Distant | snellen | monocular | no change | 1년 | 안 | 62 | 37.82
1309 | Luger (2013) | [5 empty] | VA | CNVA | corrected | near | snellen | monocular | no change | 1년 | 안 | 62 | 36.58
1309 | Luger (2013) | [5 empty] | VA | CDVA | corrected | Distant | snellen | monocular | gain 1 | 1년 | 안 | 62 | 1.86
1309 | Luger (2013) | [5 empty] | VA | CNVA | corrected | near | snellen | monocular | gain 1 | 1년 | 안 | 62 | 6.82
1309 | Luger (2013) | [5 empty] | VA | CDVA | corrected | Distant | snellen | monocular | gain 2 | 1년 | 안 | 62 | 0
1309 | Luger (2013) | [5 empty] | VA | CNVA | corrected | near | snellen | monocular | gain 2 | 1년 | 안 | 62 | 0
1309 | Luger (2013) | [5 empty] | VA | CDVA | corrected | Distant | snellen | monocular | gain 3 or more | 1년 | 안 | 62 | 0
1309 | Luger (2013) | [5 empty] | VA | CNVA | corrected | near | snellen | monocular | gain 3 or more | 1년 | 안 | 62 | 0
1310 | Oh (2013) |  |  |  |  | CST | VA | UNVA | uncorrected | Near |  |  | J1(≥20/25) or better | 2년 | 명 | 12 | 8.4
1310 | Oh (2013) |  |  |  |  | NCST | VA | UNVA | uncorrected | Near |  |  | J1(≥20/25) or better | 2년 | 명 | 11 | 3.3
1310 | Oh (2013) |  |  |  |  | CST | VA | UNVA | uncorrected | Near |  |  | J2-J3 (20/40≤ <20/25) | 2년 | 명 | 12 | 3.6
1310 | Oh (2013) |  |  |  |  | NCST | VA | UNVA | uncorrected | Near |  |  | J2-J3 (20/40≤ <20/25) | 2년 | 명 | 11 | 4.62
1310 | Oh (2013) |  |  |  |  | CST | VA | UNVA | uncorrected | Near |  |  | J4 or worse (≤20/50) | 2년 | 명 | 12 | 0
1310 | Oh (2013) |  |  |  |  | NCST | VA | UNVA | uncorrected | Near |  |  | J4 or worse (≤20/50) | 2년 | 명 | 11 | 3.08
1392 | Uthoff (2012) |  |  |  |  | hyperopic | VA | UDVA | uncorrected | Distant | logMAR |  | 0.1 logMAR or better | 6개월 | 명 | 10 | 10
1392 | Uthoff (2012) |  |  |  |  | emmetropic | VA | UDVA | uncorrected | Distant | logMAR |  | 0.1 logMAR or better | 6개월 | 명 | 10 | 8
1392 | Uthoff (2012) |  |  |  |  | myopic | VA | UDVA | uncorrected | Distant | logMAR |  | 0.1 logMAR or better | 6개월 | 명 | 10 | 7
1392 | Uthoff (2012) |  |  |  |  | all | VA | UDVA | uncorrected | Distant | logMAR |  | 0.1 logMAR or better | 6개월 | 명 | 30 | 24.9
1392 | Uthoff (2012) |  |  |  |  | hyperopic | VA | UDVA | uncorrected | Distant | logMAR |  | 0.3 logMAR or better | 6개월 | 명 | 10 | 10
1392 | Uthoff (2012) |  |  |  |  | emmetropic | VA | UDVA | uncorrected | Distant | logMAR |  | 0.3 logMAR or better | 6개월 | 명 | 10 | 10
1392 | Uthoff (2012) |  |  |  |  | myopic | VA | UDVA | uncorrected | Distant | logMAR |  | 0.3 logMAR or better | 6개월 | 명 | 10 | 10
1392 | Uthoff (2012) |  |  |  |  | hyperopic | VA | UNVA | uncorrected | Near | logRAD |  | 0.1 logRAD or better | 6개월 | 명 | 10 | 3
1392 | Uthoff (2012) |  |  |  |  | emmetropic | VA | UNVA | uncorrected | Near | logRAD |  | 0.1 logRAD or better | 6개월 | 명 | 10 | 4
1392 | Uthoff (2012) |  |  |  |  | myopic | VA | UNVA | uncorrected | Near | logRAD |  | 0.1 logRAD or better | 6개월 | 명 | 10 | 6
1392 | Uthoff (2012) |  |  |  |  | hyperopic | VA | UNVA | uncorrected | Near | logRAD |  | 0.3 logRAD or better | 6개월 | 명 | 10 | 8
1392 | Uthoff (2012) |  |  |  |  | emmetropic | VA | UNVA | uncorrected | Near | logRAD |  | 0.3 logRAD or better | 6개월 | 명 | 10 | 9
1392 | Uthoff (2012) |  |  |  |  | myopic | VA | UNVA | uncorrected | Near | logRAD |  | 0.3 logRAD or better | 6개월 | 명 | 10 | 8
1392 | Uthoff (2012) |  |  |  |  | all | VA | UNVA | uncorrected | Near | logRAD |  | 0.3 logRAD or better | 6개월 | 명 | 30 | 24.9
1392 | Uthoff (2012) |  |  |  |  | hyperopic | VA | DBCVA | distance best-corrected | Distant |  |  | lost 2 line | 6개월 | 명 | 10 | 1
1392 | Uthoff (2012) |  |  |  |  | hyperopic | VA | DBCVA | distance best-corrected | Distant |  |  | lost 1 line | 6개월 | 명 | 10 | 4
1392 | Uthoff (2012) |  |  |  |  | hyperopic | VA | DBCVA | distance best-corrected | Distant |  |  | unchanged to the preoperative level | 6개월 | 명 | 10 | 5
1392 | Uthoff (2012) |  |  |  |  | hyperopic | VA | DBCVA | distance best-corrected | Distant |  |  | gained any line | 6개월 | 명 | 10 | 0
1392 | Uthoff (2012) |  |  |  |  | emmetropic | VA | DBCVA | distance best-corrected | Distant |  |  | lost 2 line | 6개월 | 명 | 10 | 1
1392 | Uthoff (2012) |  |  |  |  | emmetropic | VA | DBCVA | distance best-corrected | Distant |  |  | lost 1 line | 6개월 | 명 | 10 | 4
1392 | Uthoff (2012) |  |  |  |  | emmetropic | VA | DBCVA | distance best-corrected | Distant |  |  | unchanged to the preoperative level | 6개월 | 명 | 10 | 5
1392 | Uthoff (2012) |  |  |  |  | emmetropic | VA | DBCVA | distance best-corrected | Distant |  |  | gained any line | 6개월 | 명 | 10 | 0
1392 | Uthoff (2012) |  |  |  |  | myopic | VA | DBCVA | distance best-corrected | Distant |  |  | lost 3 line | 6개월 | 명 | 10 | 1
1392 | Uthoff (2012) |  |  |  |  | myopic | VA | DBCVA | distance best-corrected | Distant |  |  | lost 2 line | 6개월 | 명 | 10 | 1
1392 | Uthoff (2012) |  |  |  |  | myopic | VA | DBCVA | distance best-corrected | Distant |  |  | lost 1 line | 6개월 | 명 | 10 | 1
1392 | Uthoff (2012) |  |  |  |  | myopic | VA | DBCVA | distance best-corrected | Distant |  |  | unchanged | 6개월 | 명 | 10 | 6
1392 | Uthoff (2012) |  |  |  |  | myopic | VA | DBCVA | distance best-corrected | Distant |  |  | gained a line | 6개월 | 명 | 10 | 1
1392 | Uthoff (2012) |  |  |  |  | all | VA | DBCVA | distance best-corrected | Distant |  | monocular | lost 3 or more | 6개월 | 명 | 30 | 0.9
1392 | Uthoff (2012) |  |  |  |  | all | VA | DBCVA | distance best-corrected | Distant |  | binocular | lost 3 or more | 6개월 | 명 | 30 | 0.9
1392 | Uthoff (2012) |  |  |  |  | all | VA | DBCVA | distance best-corrected | Distant |  | monocular | loss 2 | 6개월 | 명 | 30 | 3.6
1392 | Uthoff (2012) |  |  |  |  | all | VA | DBCVA | distance best-corrected | Distant |  | binocular | loss 2 | 6개월 | 명 | 30 | 3
1392 | Uthoff (2012) |  |  |  |  | all | VA | DBCVA | distance best-corrected | Distant |  | monocular | loss 1 | 6개월 | 명 | 30 | 12.9
1392 | Uthoff (2012) |  |  |  |  | all | VA | DBCVA | distance best-corrected | Distant |  | binocular | loss 1 | 6개월 | 명 | 30 | 9
1392 | Uthoff (2012) |  |  |  |  | all | VA | DBCVA | distance best-corrected | Distant |  | monocular | unchanged to the preoperative level | 6개월 | 명 | 30 | 9.6
1392 | Uthoff (2012) |  |  |  |  | all | VA | DBCVA | distance best-corrected | Distant |  | binocular | unchanged to the preoperative level | 6개월 | 명 | 30 | 15.9
1392 | Uthoff (2012) |  |  |  |  | all | VA | DBCVA | distance best-corrected | Distant |  | monocular | gain 1 | 6개월 | 명 | 30 | 2.4
1392 | Uthoff (2012) |  |  |  |  | all | VA | DBCVA | distance best-corrected | Distant |  | binocular | gain 1 | 6개월 | 명 | 30 | 0.9
1392 | Uthoff (2012) |  |  |  |  | all | VA | DBCVA | distance best-corrected | Distant |  | monocular | gain 2 or more | 6개월 | 명 | 30 | 0.6
1392 | Uthoff (2012) |  |  |  |  | all | VA | DBCVA | distance best-corrected | Distant |  | binocular | gain 2 or more | 6개월 | 명 | 30 | 0

## 비뚤림 위험평가(RoB)_RS
4. 비뚤림위험평가
L: 낮음, H: 높음, U: 불확실, N: 해당없음
no. | 1저자 (연도) | 연구설계 | 질환 | 1. 무작위 배정순서 생성 | 1. 판단근거 | 2. 배정순서 은폐 | 2. 판단근거 | 3. 연구참여자, 연구자에 대한 눈가림 | 3. 판단근거 | 4. 결과평가에 대한 눈가림 | 4. 판단근거 | 5. 불충분한 결과자료 | 5. 판단근거 | 6. 선택적 보고 | 6. 판단근거 | 7. 민간연구지원 비뚤림 | 7. 판단근거
329 | Khalifa (2011) | Prospective consecutive comparative randomized double-masked study | Myopic presbyopia | U | They were randomly divided into two groups, A and B. | U | 언급없음 | L | 눈가림이 결과에 영향을 미치지 않을 것으로 판단됨 | L | Double-masked postoperative follow-up visits were ­scheduled at 1 day, 1 week, 1 month, and 3 months | L | 결측치 없음 | L | 연구방법에 기술된 결과 모두 보고 | L | The authors have no conflicts of interest to declare in relation to this paper.

## 비뚤림 위험평가 (RoBANS)_RS
4. 비뚤림위험평가
L: 낮음, H: 높음, U: 불확실, N: 해당없음
no. | 1저자 (연도) | 연구설계 | 대상자 | 1. 대상군 비교가능성 | 1. 판단근거 | 2. 대상군 선정 | 2. 판단근거 | 3. 교란변수 | 3. 판단근거 | 4. 노출측정 | 4. 판단근거 | 5. 평가자의 눈가림 | 5. 판단근거 | 6. 결과평가 | 6. 판단근거 | 7. 불완전한 결과자료 | 7. 판단근거 | 8. 선택적 결과보고 | 8. 판단근거 | 9. 민간연구지원 비뚤림 | 9. 판단근거
225 | Jackson (2011) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | L | 중재에 대한 노출전후의 인구집단이 동일하다 | U | 자료를 전향적으로 수집하였으나 연속적 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있다 | L | 수술기록이 명확함 | L | 평가자 눈가림은 없지만 눈가림 여부가 노출측정에 영향을 미치지 않음 | H | Questionnaires were created for the study and were not validated independently | U | 탈락자와 완료자의 기저상태차이 확인할 수 없음 | L | 방법에 기술된 결과변수를 모두 보고함 | H | This study was supported by a VISX/Abbott Medical Optics (AMO) research trial grant (Jackson).
430 | Kohnen (2020) | Prospective, nonrandomized trial | Presbyopia | L | 중재에 대한 노출전후의 인구집단이 동일하다 | U | 자료를 전향적으로 수집하였으나 연속적 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있다 | L | 수술기록이 명확함 | L | 평가자 눈가림은 없지만 눈가림 여부가 노출측정에 영향을 미치지 않음 | L | The questionnaire is a modified version of the one already used by the authors in other studies. Questions regarding spectacle independence were similar to those on the modified Cataract Type Specification QoV instrument. | U | 탈락률 3%, 탈락자와 완료자의 기저상태차이 확인할 수 없음 | L | 방법에 기술된 결과변수를 모두 보고함 | H | The surgical procedure costs were paid by SCHWIND eye-tech-solutions GmbH, Kleinostheim, Germany.
457 | Xu (2020) | Prospective, non-randomized, non-comparative study | Emmetropic presbyopia | L | 중재에 대한 노출전후의 인구집단이 동일하다 | U | 자료를 전향적으로 수집하였으나 연속적 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있다 | L | 수술기록이 명확함 | L | 평가자 눈가림은 없지만 눈가림 여부가 노출측정에 영향을 미치지 않음 | U | 결과평가에 대한 세부기술이 없음 | L | 탈락자 없음 | U | 방법에 결과변수에 대한 기술이 없음 | L | The study was supported in part by the National Natural Science Foundation of China (Grant No.81770955), and Shanghai Municipal Commission of Health and Family Planning (Grant No. 20154Y0150).
497 | Liu (2020) | Prospective, non-comparative case series | Myopic presbyopia | L | 중재에 대한 노출전후의 인구집단이 동일하다 | U | 자료를 전향적으로 수집하였으나 연속적 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있다 | L | 수술기록이 명확함 | L | 평가자 눈가림은 없지만 눈가림 여부가 노출측정에 영향을 미치지 않음 | L | 선행연구에서 활용된 도구를 이용함 | U | 탈락률 8%, 탈락자와 완료자의 기저상태차이 확인할 수 없음 | L | 방법에 기술된 결과변수를 모두 보고함 | L | None of Confilicts of interest
695 | Romero (2019) | Prospective case series | Myopic presbyopia, hyperopic presbyopia | L | 중재에 대한 노출전후의 인구집단이 동일하다 | U | 자료를 전향적으로 수집하였으나 연속적 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있다 | L | 수술기록이 명확함 | L | 평가자 눈가림은 없지만 눈가림 여부가 노출측정에 영향을 미치지 않음 | L | 타당도 입증된 방법으로 결과측정 | L | 탈락자 없음 | L | 방법에 기술된 결과변수를 모두 보고함 | L | None of the authors has a financial or proprietary in- terest in any material or method mentioned.
734 | Taneri (2019) | Prospective, non-randomized, comparative case series | Hyperopic presbyopia | L | 중재에 대한 노출과 비교군이 질병특성에 차이가 없어 비교가능함 | L | 참여자 모집전략이 두 군에서 동일하다 | U | 교란변수에 대한 확인할 근거가 부족함 | L | 수술기록이 명확함 | L | 평가자 눈가림은 없지만 눈가림 여부가 노출측정에 영향을 미치지 않음 | L | 타당도 입증된 방법으로 결과측정 | L | 탈락률 20%, 결측치가 생긴이유가 결과에 주는 영향이 없을 것으로 판단됨 | U | 방법에 결과변수에 대한 기술이 없음 | H | 저자 중 1명이 기기회사소속임
897 | Pajic (2017) | Prospective, single-surgeon study | Myopic presbyopia | L | 중재에 대한 노출전후의 인구집단이 동일하다 | U | 참여자 모집전략에 대한 기술없음 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있다 | L | 수술기록이 명확함 | L | 평가자 눈가림은 없지만 눈가림 여부가 노출측정에 영향을 미치지 않음 | L | 타당도 입증된 방법으로 결과측정 | L | 탈락자 없음 | L | 방법에 기술된 결과변수를 모두 보고함 | L | The study was supported by the Swiss Commission for Technology and Innovation (CTI), through CTI project 13404.1.
929 | Schlote (2017) | Prospective consecutive one-center cross sectional study | Hyperopic presbyopia | L | 중재에 대한 노출전후의 인구집단이 동일하다 | L | prospective consecutive cross sectional study | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있다 | L | 수술기록이 명확함 | L | 평가자 눈가림은 없지만 눈가림 여부가 노출측정에 영향을 미치지 않음 | U | questionnaire의 타당도 입증여부 확인할 수 있는 문장이 없음 | L | 탈락자 없음 | L | 방법에 기술된 결과변수를 모두 보고함 | L | No financial support was received for this submission
958 | Wang Yin (2016) | Prospective cohort study | Hyperopic presbyopia | L | 중재에 대한 노출전후의 인구집단이 동일하다 | U | 자료를 전향적으로 수집하였으나 연속적 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있다 | L | 수술기록이 명확함 | L | 평가자 눈가림은 없지만 눈가림 여부가 노출측정에 영향을 미치지 않음 | L | 타당도 입증된 방법으로 결과측정 | L | 탈락자 없음 | L | 방법에 기술된 결과변수를 모두 보고함 | L | None of the authors has a financial or proprietary interest in any material or method mentioned.
963 | Zhang (2016) | Prospective, non-comparative case series | Myopic presbyopia | L | 중재에 대한 노출전후의 인구집단이 동일하다 | U | 자료를 전향적으로 수집하였으나 연속적 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있다 | L | 수술기록이 명확함 | L | 평가자 눈가림은 없지만 눈가림 여부가 노출측정에 영향을 미치지 않음 | U | 만족도 조사도구의 타당도 입증여부 확인불가 | L | 탈락자 없음 | L | 방법에 기술된 결과변수를 모두 보고함 | L | Supported by grants from the Natural Science Foundation of China (81261130021); The Science and Technology Program of Guangzhou (2013B 021800126); and The Science and Technology Program of Guangzhou (12S 202060079).
1123 | Soler Tomas (2015) | Prospective, longitudinal case series | Hyperopic presbyopia | L | 중재에 대한 노출군과 비교군이 적응증과 질병의 중증도 등에 차이가 없어 비교할만한 인구집단임 | U | 참여자 모집전략이 모두 동일함 | U | 교란변수에 대한 확인할 근거가 부족함 | L | 수술기록이 명확함 | L | 평가자 눈가림은 없지만 눈가림 여부가 노출측정에 영향을 미치지 않음 | L | 타당도 입증된 방법으로 결과측정 | L | 탈락자 없음 | L | 방법에 기술된 결과변수를 모두 보고함 | L | The authors have no funding or conflicts of interest to disclose.
1225 | Cosar (2014) | Prospective, nonrandomized, clinical trial | Hyperopic presbyopia | L | 중재에 대한 노출전후의 인구집단이 동일하다 | U | A total of 123 eyes from 68 patients that underwent Supracor treatment from July 1, 2011, to September 30, 2012, were enrolled in this prospective, nonrandomized, clinical trial | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있다 | L | 수술기록이 명확함 | L | 평가자 눈가림은 없지만 눈가림 여부가 노출측정에 영향을 미치지 않음 | L | 타당도 입증된 방법으로 결과측정 | L | 탈락자 없음 | L | 방법에 기술된 결과변수를 모두 보고함 | L | No financial support was received for this submission.
1268 | Ryan (2013) | Prospective case series | Hyperopic presbyopia | L | 중재에 대한 노출전후의 인구집단이 동일하다 | L | 자료를 전향적, 연속적으로 수집(This was a prospective evaluation of the first consecutive series of hyperopic presbyopic patients treated with the Supracor multifocal excimer LASIK procedure over a 3-month period ) | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있다 | L | 수술기록이 명확함 | L | 평가자 눈가림은 없지만 눈가림 여부가 노출측정에 영향을 미치지 않음 | U | questionnaire의 타당도 입증여부 확인할 수 있는 문장이 없음 | L | 탈락자 없음 | L | 방법에 기술된 결과변수를 모두 보고함 | L | Neither author has a financial or proprietary interest in any material or method mentioned.
1309 | Luger (2013) | Prospective study | Presbyopia | L | 중재에 대한 노출전후의 인구집단이 동일하다 | U | 자료를 전향적으로 수집하였으나 연속적 모집여부 확인불가 Sixty-six eyes of 33 patients undergoing bilateral laser in situ keratomileusis (LASIK) for refractive presbyopic corrections were enrolled | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있다 | L | 수술기록이 명확함 | L | 평가자 눈가림은 없지만 눈가림 여부가 노출측정에 영향을 미치지 않음 | L | 타당도 입증된 방법으로 결과측정 | U | 탈락률 6%, 탈락자와 완료자의 기저상태차이 확인할 수 없음 | L | 방법에 기술된 결과변수를 모두 보고함 | H | The authors state that they have no proprietary interest in the products named in this article. T. Ewering and S. Arba-Mosquera are employees of Schwind eye-techsolutions.
1310 | Oh (2013) | Prospective comparative study | Presbyopia | L | 중재에 대한 노출군과 비교군이 적응증과 질병의 중증도 등에 차이가 없어 비교할만한 인구집단임 | L | 참여자 모집전략이 두 군에서 동일하다 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있다 | L | 수술기록이 명확함 | L | 평가자 눈가림은 없지만 눈가림 여부가 노출측정에 영향을 미치지 않음 | L | 타당도 입증된 방법으로 결과측정 | L | 탈락자 없음 | L | 방법에 기술된 결과변수를 모두 보고함 | L | Supported by the National Research Foundation of Korea, 2011.
1392 | Uthoff (2012) | Prospective non-randomised clinical evaluation | Hyperopic presbyopia, emmetropic presbyopia, myopic presbyopia | L | 중재에 대한 노출전후의 인구집단이 동일하다 | U | 자료를 전향적으로 수집하였으나 연속적 모집여부 확인불가 | L | 질병특성상 시간경과에 따른 전후차이를 배제할 수 있다 | L | 수술기록이 명확함 | L | 평가자 눈가림은 없지만 눈가림 여부가 노출측정에 영향을 미치지 않음 | L | 타당도 입증된 방법으로 결과측정 | L | 탈락자 없음 | L | 방법에 기술된 결과변수를 모두 보고함 | L | The authors have no financial interests in the presented products, and no financial relationship with the Schwind Company.